## MET
Date | Julian Day | Time (UTC) | Observer's Initials | 2m T | 2m TD | 9m SONIC | 9m RM Young | Check Enclosure Temp (~12 C) | Check date and time | Comments (Climbing, removing instruments, etc)
2024-01-01 | 1 | 1550 | GM | rm lf | rm lf | clr | hf | * | * | Enclosure hovering around 1C
2024-01-02 | 2 | 1240 | GM | rm lf | rm lf | clr | hf | * | *
2024-01-03 | 3 | 1600 | GM | rm lf | rm lf | clr | hf | * | * | Enclosure has been above 12C since 0515Z
2024-01-04 | 4 | 1200 | GM | rm lf | rm lf | clr | hf | * | *
2024-01-05 | 5 | [8 empty] | Did not make it out to AWO after ICESat
2024-01-06 | 6 | 1325 | GM | rm lf | rm lf | clr | hf | * | *
2024-01-07 | 7 | 1420 | GM | clr | clr | clr | clr | * | *
2024-01-08 | 8 | 1615 | GM | clr | clr | clr | clr | * | *
2024-01-09 | 9 | 1445 | GM | clr | clr | clr | clr | * | *
2024-01-10 | 10 | 1155 | GM | clr | clr | clr | clr | * | *
2024-01-11 | 11 | 1400 | GM | clr | clr | clr | clr | * | *
2024-01-12 | 12 | 1555 | GM | rm lf | rm lf | clr | lf | * | * | Enclosure back below 12C again
2024-01-13 | 13 | 1250 | GM | rm lf | rm lf | clr | lf | * | *
2024-01-14 | 14 | 1505 | GM | rm lf | rm lf | clr | hf | * | * | RM Young has too much frost to move.
2024-01-15 | 15 | 1440 | GM | rm lf | rm lf | clr | lf | * | * | RM Young has been revived. The Snow Beast was ran for about 15min on station at the beginning of North winds 1600-1615Z
2024-01-16 | 16 | 1310 | GM | rm lf | rm lf | clr | lf | * | * | Encosure below 12C. Climbed Tower from 1600-1735Z for general cleaning and cable checking
2024-01-17 | 17 | 1550 | GM | clr | clr | clr | clr | * | * | Enclosure still below 12C
2024-01-18 | 18 | 1620 | GM | rm lf | rm lf | clr | clr | * | *
2024-01-19 | 19 | 1240 | GM | rm lf | rm lf | clr | clr | * | *
2024-01-20 | 20 | 1215 | GM | rm lf | rm lf | clr | clr | * | *
2024-01-21 | 21 | 1730 | GM | rm lf | rm lf | clr | lf | * | *
2024-01-22 | 22 | 1400 | GM | clr | clr | clr | lf | * | *
2024-01-23 | 23 | 1305 | GM | clr | rm lf | clr | lf | * | *
2024-01-24 | 24 | 1225 | GM | rm lf | rm lf | clr | lf | * | *
2024-01-25 | 25 | 1700 | GM | clr | clr | clr | lf | * | * | Enclosure below freezing
2024-01-26 | 26 | 1250 | GM | rm lf | rm lf | clr | lf | * | * | Enclosure below freezing
2024-01-27 | 27 | 1315 | GM | rm lf | rm lf | clr | lf | * | *
2024-01-28 | 28 | 1405 | GM | clr | clr | clr | lf | * | *
2024-01-29 | 29 | 1335 | GM | rm lf | rm lf | clr | lf | * | *
2024-01-30 | 30 | 1335 | GM | rm lf | rm lf | clr | lf | * | *
2024-01-31 | 31 | 1705 | GM | rm lf | rm lf | clr | hf | * | *
2024-02-01 | 32 | 1315 | GM | rm lf | rm lf | clr | hf | * | *
2024-02-02 | 33 | [6 empty] | * | * | Did not make it out to AWO after ICESat
2024-02-03 | 34 | 1310 | GM | rm lf | rm lf | clr | hf | * | * | RM Young too frosted to move again
2024-02-04 | 35 | 1520 | GM | rm lf | rm lf | clr | hf | * | * | RM Young too frosted to move again
2024-02-05 | 36 | 1405 | GM | rm lf | rm lf | clr | hf | * | * | RM Young too frosted to move again
2024-02-06 | 37 | 1650 | GM | rm lf | rm lf | clr | hf | * | * | RM Young too frosted to move again
2024-02-07 | 38 | 1340 | GM | rm lf | rm lf | clr | hf | * | * | RM Young too frosted to move again
2024-02-08 | 39 | 1945 | GM | clr | clr | clr | hf | * | * | RM Young too frosted to move again
2024-02-09 | 40 | 1515 | GM DR | clr | rm lf | clr | hf | * | * | RM Young too frosted to move again
2024-02-10 | 41 | 1355 | DR GM | rm lf | rm lf | clr | hf | * | * | RM Young too frosted to move again
2024-02-11 | 42 | 1350 | DR GM | rm lf | rm lf | clr | hf | * | * | RM Young too frosted to move again
2024-02-12 | 43 | 1245 | DR GM | clr | clr | clr | hf | * | * | RM Young too frosted to move again
2024-02-13 | 44 | [8 empty] | Did not make it out to AWO after ICESat
2024-02-14 | 45 | [8 empty] | Did not make it out to AWO due to Con 1 conditions
2024-02-15 | 46 | 1420 | GM IG | rm lf | rm lf | clr | hf | * | * | RM Young too frosted to move again
2024-02-16 | 47 | 1350 | DR | rm lf | rm lf | clr | hf | * | * | RM Young too frosted to move again
2024-02-17 | 48 | 1835 | DR | rm lf | rm lf | clr | hf | * | * | RM Young too frosted to move again
2024-02-18 | 49 | 1405 | DR | rm hf | rm hf | clr | hf | * | * | RM Young too frosted to move again
2024-02-19 | 50 | 1345 | DR | rm hf | rm hf | clr | hf | * | * | RM Young too frosted to move again
2024-02-20 | 51 | 1605 | DR | rm hf | rm hf | clr | hf | * | * | RM Young too frosted to move again. Climbed tower 1430-1520UTC - only cleaned half, will plan another climb next week.
2024-02-21 | 52 | 1545 | DR | rm lf | rm lf | clr | hf | * | * | RM Young too frosted to move again
2024-02-22 | 53 | 1620 | DR | rm lf | rm lf | clr | hf | * | * | RM Young too frosted to move again
2024-02-23 | 54 | 1335 | DR | rm lf | rm lf | clr | hf | * | * | RM Young too frosted to move again
2024-02-24 | 55 | 1410 | DR | rm lf | rm lf | clr | hf | * | * | RM Young too frosted to move again
2024-02-25 | 56 | 1645 | DR | rm lf | rm hf | clr | hf | * | * | RM Young too frosted to move again. RM Young stopped reporting data (0) from 0850-1026.
2024-02-26 | 57 | 1405 | DR | rm lf | rm lf | clr | hf | * | * | RM Young too frosted to move again
2024-02-27 | 58 | 1445 | DR | rm lf | rm lf | clr | hf | * | * | RM Young too frosted to move again
2024-02-28 | 59 | 1345 | DR | rm lf | rm lf | clr | hf | * | * | RM Young too frosted to fully move, but is moving slightly back and forth. RM Young stopped reporting data (0) at 2100.
2024-02-29 | 60 | 1440 | DR | rm lf | rm lf | clr | lf | * | * | RM Young is freely spinning. RM Young started reporting data again at 2309. Wind shear from inversion experienced.
2024-03-01 | 61 | 1410 | DR | clr | rm lf | clr | hf | * | * | RM Young too frosted to fully move, but is moving slightly back and forth.
2024-03-02 | 62 | 1055 | DR | rm lf | rm lf | clr | lf | * | * | RM Young too frosted to fully move, but is moving slightly back and forth.
2024-03-03 | 63 | 1835 | DR | clr | clr | clr | lf | * | *
2024-03-04 | 64 | 1805 | DR | rm lf | rm lf | clr | rm lf | * | * | Tower climbed at 1353-1427UTC. All instruments cleared of rime buildup.
2024-03-05 | 65 | 1505 | DR | clr | clr | clr | clr | * | *
2024-03-06 | 66 | 1305 | DR | clr | clr | clr | clr | * | *
2024-03-07 | 67 | 1610 | DR | clr | clr | clr | clr | * | *
2024-03-08 | 68 | 1145 | DR | rm lf | rm lf | clr | clr | * | *
2024-03-09 | 69 | 1250 | DR | clr | clr | clr | clr | * | *
2024-03-10 | 70 | 1545 | DR | clr | clr | clr | clr | * | *
2024-03-11 | 71 | 1535 | DR | rm lf | rm lf | clr | clr | * | *
2024-03-12 | 72 |  | DR | [6 empty] | Did not make it out to AWO after ICESat
2024-03-13 | 73 | 1330 | DR | rm lf | rm lf | clr | lf | * | *
2024-03-14 | 74 | 1645 | DR | clr | clr | clr | lf | * | * | NOAA weather websites not updated since 3/13 1239
2024-03-15 | 75 | 1425 | DR | rm lf | rm lf | clr | lf | * | * | NOAA weather websites not updated since 3/13 1239
2024-03-16 | 76 | 1330 | DR | rm lf | rm lf | clr | lf | * | *
2024-03-17 | 77 | 1915 | DR | rm lf | rm lf | clr | lf | * | *
2024-03-18 | 78 | 1605 | DR | rm lf | rm lf | clr | lf | * | *
2024-03-19 | 79 | [8 empty] | Did not make it out to AWO due to CON1
2024-03-20 | 80 | 1135 | DR | rm lf | rm lf | clr | clr | * | *
2024-03-21 | 81 | 1635 | DR | rm lf | rm lf | clr | clr | * | *
2024-03-22 | 82 | 1250 | DR | rm lf | rm lf | clr | clr | * | *
2024-03-23 | 83 | 1735 | DR | rm lf | rm lf | clr | clr | * | *
2024-03-24 | 84 | 1515 | DR | rm lf | rm lf | clr | clr | * | *
2024-03-25 | 85 | 1650 | DR | rm hf | rm hf | clr | clr | * | * | A lot of buildup snow from inside the enclosures came out.
2024-03-26 | 86 | 1500 | DR | clr | clr | clr | clr | * | *
2024-03-27 | 87 | 1820 | DR | rm lf | rm lf | clr | clr | * | *
2024-03-28 | 88 | 1755 | DR | rm lf | rm lf | clr | clr | * | *
2024-03-29 | 89 | 1310 | DR | rm lf | rm lf | clr | clr | * | *
2024-03-30 | 90 | 1315 | DR | rm lf | rm lf | clr | clr | * | * | NOAA weather websites not updated since 3/29 1448
2024-03-31 | 91 | 1540 | DR | clr | clr | clr | clr | * | *
2024-04-01 | 92 | 1540 | DR | rm lf | rm lf | clr | lf | * | *
2024-04-02 | 93 | 1635 | DR | clr | clr | clr | lf | * | *
2024-04-03 | 94 |  | DR | [6 empty] | Did not make it out to AWO after ICESat
2024-04-04 | 95 | 1200 | DR | rm lf | rm lf | clr | clr | * | *
2024-04-05 | 96 | 1215 | DR | rm lf | rm lf | clr | lf | * | *
2024-04-06 | 97 | 1235 | DR | rm lf | rm lf | clr | clr | * | *
2024-04-07 | 98 | 1545 | DR | clr | clr | clr | clr | * | *
2024-04-08 | 99 | 1125 | DR | rm lf | rm lf |  |  | * | *
2024-04-09 | 100 | 1710 | DR | rm lf | rm lf | clr | clr | * | *
2024-04-10 | 101 | 1440 | DR | rm lf | rm lf | clr | clr | * | *
2024-04-11 | 102 | 1230 | DR | rm lf | rm lf | clr | clr | * | *
2024-04-12 | 103 | 1200 | DR | rm lf | rm lf | clr | clr | * | *
2024-04-13 | 104 | 1135 | DR | rm lf | rm lf | clr | clr | * | *
2024-04-14 | 105 | 1905 | DR | clr | clr | clr | clr | * | *
2024-04-15 | 106 | 1625 | DR | clr | clr | clr | clr | * | *
2024-04-16 | 107 | 1300 | DR | rm lf | rm lf | clr | clr | * | *
2024-04-17 | 108 | 2045 | DR | clr | clr | clr | clr | * | *
2024-04-18 | 109 | 1735 | DR | rm lf | rm hf | clr | clr | * | * | Climbed entire tower 1622-1704. Unclasped bottom portion og 2m aspirated humidity enclosure - cleared out snow buildup.
2024-04-19 | 110 | 1745 | DR | clr | clr | clr | clr | * | *
2024-04-20 | 111 | 1615 | DR | clr | clr | clr | clr | * | *
2024-04-21 | 112 | 1340 | DR | clr | clr | clr | clr | * | *
2024-04-22 | 113 | 1225 | DR | clr | clr | clr | clr | * | *
2024-04-23 | 114 | 1035 | DR | clr | clr | clr | clr | * | *
2024-04-24 | 115 | 1550 | DR | clr | clr | clr | clr | * | *
2024-04-25 | 116 |  | DR | [6 empty] | HM - did not make it out to AWO.
2024-04-26 | 117 | 1120 | DR | rm lf | clr | clr | clr | * | *
2024-04-27 | 118 | 1220 | DR | clr | clr | clr | clr | * | *
2024-04-28 | 119 | 1515 | DR | clr | clr | clr | clr | * | *
2024-04-29 | 120 | 1800 | DR | clr | clr | clr | clr | * | *
2024-04-30 | 121 | 1230 | DR | clr | clr | clr | clr | * | *
2024-05-01 | 122 | 1520 | DR | clr | clr | clr | clr | * | *
2024-05-02 | 123 | 1220 | DR | clr | clr | clr | clr | * | *
2024-05-03 | 124 | 1700 | DR | clr | clr | clr | clr | * | *
2024-05-04 | 125 | 1640 | DR | clr | clr | clr | clr | * | *
2024-05-05 | 126 |  | DR | [6 empty] | HW - did not make it out to AWO.
2024-05-06 | 127 | 1115 | DR | clr | clr | clr | clr | * | *
2024-05-07 | 128 | 1715 | DR | rm lf | rm lf | clr | clr | * | *
2024-05-08 | 129 | 940 | DR | rm lf | rm lf | clr | clr | * | *
2024-05-09 | 130 | 2010 | DR | clr | clr | clr | clr | * | *
2024-05-10 | 131 |  | DR | [6 empty] | Did not make it out to AWO after ICESat
2024-05-11 | 132 | 1650 | DR | clr | clr | clr | clr | * | *
2024-05-12 | 133 | 1350 | DR | rm lf | rm lf | clr | clr | * | *
2024-05-13 | 134 | 1735 | DR | rm lf | clr | clr | clr | * | *
2024-05-14 | 135 | 1600 | DR | clr | clr | clr | clr | * | *
2024-05-15 | 136 |  | DR | [6 empty] | Did not make it out to AWO after ICESat
2024-05-16 | 137 | 1515 | DR | clr | clr | clr | clr | * | *
2024-05-17 | 138 | 1220 | DR | clr | clr | clr | clr | * | *
2024-05-18 | 139 | 1600 | DR | rm lf | rm lf | clr | clr | * | * | Unclasped longer enclosure and cleared light rime.
2024-05-19 | 140 | 1835 | DR | rm lf | rm lf | clr | clr | * | *
2024-05-20 | 141 | 1210 | DR | rm lf | rm lf | clr | clr | * | *
2024-05-21 | 142 |  | DR | [6 empty] | HM - did not make it out to AWO.
2024-05-22 | 143 | 1045 | DR | rm lf | rm lf | clr | clr | * | *
2024-05-23 | 144 | 1150 | DR | rm lf | rm lf | clr | clr | * | *
2024-05-24 | 145 | 1700 | DR | clr | clr | clr | clr | * | *
2024-05-25 | 146 | 1605 | DR | clr | clr | clr | clr | * | *
2024-05-26 | 147 |  | DR | [6 empty] | Unplanned power outage - storm conditions - did not check tower.
2024-05-27 | 148 |  | DR | [6 empty] | Day off.
2024-05-28 | 149 |  | DR | [6 empty] | Con 1 - didn't go out to AWO.
2024-05-29 | 150 |  | DR
2024-05-30 | 151 | 1525 | DR | clr | clr | clr | clr | * | *
2024-05-31 | 152
2024-06-01 | 153 | 1130 | DR | clr | clr | clr | clr | * | *
2024-06-02 | 154 | 1115 | DR | clr | clr | clr | clr | * | *
2024-06-03 | 155 | 1650 | DR | clr | clr | clr | clr | * | *
2024-06-04 | 156 |  | DR | [6 empty] | Did not make it out to AWO after ICESat
2024-06-05 | 157 | 1040 | DR | rm lf | rm lf | clr | lf | * | *
2024-06-06 | 158 | 1610 | CW | clr | clr | clr | clr | * | *
2024-06-07 | 159 | 1530 | CW | clr | clr | clr | clr | * | *
2024-06-08 | 160 | 845 | CW | clr | clr | clr | clr | * | *
2024-06-09 | 161 | 1030 | CW | clr | clr | clr | clr |  |  | NOAA site unresponsive
2024-06-10 | 162 | 935 | CW | clr | clr | clr | clr |  |  | NOAA site unresponsive
2024-06-11 | 163 | 1115 | CW | clr | clr | clr | clr | * | *
2024-06-12 | 164 | 1040 | CW | clr | clr | clr | clr | * | *
2024-06-13 | 165 | 1045 | CW | clr | clr | clr | clr | * | *
2024-06-14 | 166 | 955 | CW | clr | clr | clr | clr | * | *
2024-06-15 | 167 | 1510 | CW | clr | clr | clr | clr | * | *
2024-06-16 | 168 | 1440 | CW | clr | clr | clr | clr | * | *
2024-06-17 | 169 | 1230 | CW | clr | clr | clr | clr | * | * | NOAA data gapped out between 0327-0731Z.
2024-06-18 | 170 | 1020 | CW | lf | lf | clr | clr | * | *
2024-06-19 | 171 | 1503 | CW | clr | rm lf | clr | clr | * | *
2024-06-20 | 172 | 1402 | CW | clr | lf | clr | clr | * | *
2024-06-21 | 173 | 1235 | CW | clr | clr | clr | clr | * | *
2024-06-22 | 174 | 1507 | CW | clr | clr | clr | clr | * | *
2024-06-23 | 175 | 1412 | CW | rm lf | rm lf | clr | clr | * | *
2024-06-24 | 176 | 1413 | CW | clr | clr | clr | clr | * | *
2024-06-25 | 177 | 1450 | CW | clr | clr | clr | clr | * | *
2024-06-26 | 178 | 1113 | CW | clr | clr | clr | clr |  |  | NOAA site unresponsive
2024-06-27 | 179 | 1204 | CW | clr | rm lf | clr | clr | * | * | Raised bamboo to Met tower and Albedo stand.
2024-06-28 | 180 | 1017 | CW | clr | clr | clr | clr | * | *
2024-06-29 | 181 | 1020 | CW | clr | clr | clr | clr | * | * | Power outage during UPS maintenance, UPS instruments affected from 1128-1134Z.
2024-06-30 | 182 | 1240 | CW | clr | clr | clr | clr | * | *
2024-07-01 | 183 | 1557 | CW | clr | clr | clr | clr | * | *
2024-07-02 | 184 | 1450 | CW | clr | clr | clr | clr | * | *
2024-07-03 | 185 | 923 | CW | clr | clr | clr | clr | * | *
2024-07-04 | 186 |  | CW |  |  |  |  | * | * | Did not make it to AWO after ICESAT.
2024-07-05 | 187 | 1435 | CW | clr | clr | clr | clr | * | *
2024-07-06 | 188 | 1455 | CW | clr | clr | clr | clr | * | *
2024-07-07 | 189 | 1225 | CW | rm lf | rm lf | clr | clr | * | *
2024-07-08 | 190 | 1010 | CW | rm lf | rm lf | clr | clr | * | *
2024-07-09 | 191 | 1040 | CW | clr | clr | clr | clr |  |  | NOAA page timed out error
2024-07-10 | 192 | 1436 | CW | clr | clr | clr | clr | * | *
2024-07-11 | 193 | 1155 | CW | clr | clr | clr | clr |  |  | NOAA page stuck on 2024-07-10 at 2354Z
2024-07-12 | 194 | 1500 | CW | clr | clr | clr | clr | * | *
2024-07-13 | 195 | 1710 | CW | clr | clr | clr | clr | * | *
2024-07-14 | 196 | 1655 | CW | clr | clr | clr | clr | * | *
2024-07-15 | 197 | 1158 | CW | clr | clr | clr | clr | * | *
2024-07-16 | 198 | 1248 | CW | clr | clr | clr | clr | * | * | Riggers working around Met tower until 7/21.
2024-07-17 | 199 | 1401 | CW | clr | clr | clr | clr | * | * | Riggers on tower in morning, did not log times.
2024-07-18 | 200 | 1737 | CW | clr | rm lf | clr | clr | * | *
2024-07-19 | 201 | 1254 | CW | clr | clr | clr | clr | * | * | Riggers on tower 1456-1509Z
2024-07-20 | 202 | 1145 | CW | clr | clr | clr | clr | * | *
2024-07-21 | 203 | 1025 | CW | clr | clr | clr | clr | * | * | Riggers on tower 0951-1133Z & 1336-1554Z. Rigging work complete.
2024-07-22 | 204 | 1311 | CW | clr | clr | clr | clr | * | *
2024-07-23 | 205 | 1340 | CW | clr | clr | clr | clr | * | *
2024-07-24 | 206 | 1745 | CW | clr | clr | clr | clr | * | *
2024-07-25 | 207 | 1212 | CW | clr | clr | clr | clr | * | *
2024-07-26 | 208 | 1440 | CW | clr | clr | clr | clr | * | *
2024-07-27 | 209 | 1045 | CW | clr | clr | clr | clr | * | *
2024-07-28 | 210 | 1150 | CW | clr | clr | clr | clr | * | *
2024-07-29 | 211 | 1230 | CW | clr | clr | clr | clr | * | *
2024-07-30 | 212 | 1241 | CW | clr | clr | clr | clr | * | *
2024-07-31 | 213 | 1709 | CW | clr | clr | clr | clr | * | *
2024-08-01 | 214 | 1448 | CW | clr | clr | clr | clr | * | * | From 1033-1332Z Erin opened aspirated humidity and temp enclosures for calibration.
2024-08-02 | 215 | 1530 | CW | rm lf | rm lf | clr | clr | * | * | Tower climb, adjusted RM Young orientation ~7degrees east.
2024-08-03 | 216 | 1445 | CW | rm lf | lf | clr | clr | * | *
2024-08-04 | 217 | 1755 | CW | clr | clr | clr | clr | * | *
2024-08-05 | 218 | [6 empty] | * | * | Did not make it to AWO after ICESAT.
2024-08-06 | 219 | 1616 | CW | clr | clr | clr | clr | * | *
2024-08-07 | 220 | 1106 | CW | lf | lf | clr | lf | * | *
2024-08-08 | 221 | 1006 | CW | clr | clr | clr | clr | * | *
2024-08-09 | 222 | 1259 | CW | clr | clr | clr | clr | * | *
2024-08-10 | 223 | 1044 | CW | rm lf | rm lf | clr | lf | * | * | lf frost on prop of RM Young. Two data outages from 2055-2211 and 2234-2337Z.
2024-08-11 | 224 | 1838 | CW | clr | clr | clr | clr | * | *
2024-08-12 | 225 | 1200 | CW | clr | clr | clr | clr | * | *
2024-08-13 | 226 | 1012 | CW | clr | clr | clr | clr | * | *
2024-08-14 | 227 | [6 empty] | * | * | ICESat
2024-08-15 | 228 | 1449 | CW | clr | clr | clr | clr | * | *
2024-08-16 | 229 | 950 | CW | clr | clr | clr | clr | * | *
2024-08-17 | 230 | 1522 | CW | clr | clr | clr | clr |  |  | NOAA SUM Met page down, blip in data from yesterday at 2039Z- today at 1338Z.
2024-08-18 | 231 | [6 empty] | * | * | Bear protocols, did not go to AWO.
2024-08-19 | 232 | 1530 | CW | rm lf | rm lf | clr | clr | * | *
2024-08-20 | 233 | 1339 | CW | rm lf | rm lf | clr | clr |  |  | NOAA SUM met page(s) down
2024-08-21 | 234 | 1158 | CW | rm lf | rm lf | lf | clr | * | *
2024-08-22 | 235 | 1849 | CW | rm lf | rm lf | clr | lf | * | *
2024-08-23 | 236 | 935 | CW | rm lf | rm lf | clr | lf | * | * | Internet outage from 2024-08-22 at 2224Z- 2024-08-23 at 0907Z.
2024-08-24 | 237 | 1016 | CW | rm lf | rm lf | clr | lf | * | *
2024-08-25 | 238 | 1440 | CW | rm lf | rm lf | clr | lf | * | *
2024-08-26 | 239 | 1200 | CW | rm lf | rm lf | clr | lf | * | *
2024-08-27 | 240 | 1318 | CW | rm lf | rm lf |  |  | * | *
2024-08-28 | 241 | 1526 | CW | rm lf | rm lf | clr | lf | * | *
2024-08-29 | 242 | 1345 | CW | clr | clr | clr | clr | * | *
2024-08-30 | 243 | 943 | CW | clr | clr | clr | clr | * | *
2024-08-31 | 244 | 1215 | CW | clr | clr | clr | clr | * | *
2024-09-01 | 245 | 1238 | CW | clr | clr | clr | clr | * | *
2024-09-02 | 246 | 1517 | CW | clr | clr | clr | clr | * | *
2024-09-03 | 247 |  | CW |  |  |  |  | * | * | Did not make it to AWO after ICESat.
2024-09-04 | 248 | 1331 | CW | rm lf | rm lf | clr | clr | * | *
2024-09-05 | 249 | 1245 | CW | lf | lf | clr | lf | * | *
2024-09-06 | 250 | 1406 | CW | rm lf | rm lf | clr | lf | * | *
2024-09-07 | 251 |  | CW |  |  |  |  | * | * | Did not go to AWO due to poor visibility
2024-09-08 | 252 | 1857 | CW | clr | clr | clr | clr | * | *
2024-09-09 | 253 | 1311 | CW | clr | clr | clr | clr | * | *
2024-09-10 | 254 |  | CW |  |  |  |  | * | * | Did not go to AWO due to poor visibility
2024-09-11 | 255 | 1333 | CW | rm lf | rm lf | clr | lf | * | *
2024-09-12 | 256 | 824 | CW | rm lf | rm lf | clr | lf | * | *
2024-09-13 | 257 | 949 | CW | rm lf | rm lf | clr | lf | * | *
2024-09-14 | 258 | 1123 | CW | rm lf | rm lf | clr | lf | * | *
2024-09-15 | 259 | 1824 | CW | rm lf | rm lf | clr | lf | * | *
2024-09-16 | 260 | 1211 | CW | rm lf | rm lf | clr | lf | * | *
2024-09-17 | 261 | 1820 | CW | rm lf | rm lf | clr | lf | * | *
2024-09-18 | 262 | 1254 | CW | rm lf | rm lf | clr | lf | * | *
2024-09-19 | 263 | 1208 | CW | rm lf | rm lf | clr | lf | * | *
2024-09-20 | 264 | 1623 | CW | rm lf | rm lf | clr | lf | * | *
2024-09-21 | 265 | 954 | CW | clr | clr | clr | lf | * | *
2024-09-22 | 266 | 1248 | CW | clr | clr | clr | lf | * | *
2024-09-23 | 267 | 1330 | CW | clr | clr | clr | clr | * | * | Frost on instruments appears sun effected
2024-09-24 | 268 | 1050 | CW | lf | lf | lf | lf | * | *
2024-09-25 | 269 | 1355 | CW | rm lf | rm hf | clr | hf | * | *
2024-09-26 | 270 | 1511 | CW | rm lf | rm hf | clr | hf | * | *
2024-09-27 | 271 | 1016 | CW | rm lf | rm lf | clr | lf | * | *
2024-09-28 | 272 |  | CW |  |  |  |  | * | * | Did not make it to AWO after raising sector lines
2024-09-29 | 273 | 1234 | CW | rm lf | rm lf | clr | lf | * | *
2024-09-30 | 274 | 1039 | CW | clr | clr | clr | clr | * | *
2024-10-01 | 275 | 1356 | CW | rm lf | rm lf | clr | lf | * | *
2024-10-02 | 276 | 1522 | CW | rm lf | rm lf | clr | hf | * | * | Wingbird not moving
2024-10-03 | 277 | 1448 | CW | rm lf | clr | clr | lf | * | *
2024-10-04 | 278 | 1634 | CW | rm lf | rm lf | rm lf | rm lf | * | * | Tower climb 1422-1545Z for cleaning.
2024-10-05 | 279 | [6 empty] | * | * | Did not go to AWO after ICESat
2024-10-06 | 280 | 1710 | CW | clr | clr | clr | lf | - | - | Network outage - NOAA page not updating
2024-10-07 | 281 | 1300 | CW | rm lf | rm lf | clr | clr | - | - | Network outage - NOAA page not updating
2024-10-08 | 282 | 1025 | CW | rm lf | rm lf | clr | lf | * | *
2024-10-09 | 283 | 1117 | CW | clr | clr | clr | clr | * | *
2024-10-10 | 284 | [6 empty] | * | *
2024-10-11 | 285 | 1345 | ALW | rm lf | rm lf | clr | ice | * | * | Wingbird anemometer prop is iced, whole body of wingbird is spinning on bearing
2024-10-12 | 286 | 1319 | CW | rm lf | rm lf | clr | lf | * | *
2024-10-13 | 287 | 1422 | ALW | rm lf | rm lf | clr | lf | * | *
2024-10-14 | 288 | 1708 | ALW | rm lf | rm lf | clr | lf | * | *
2024-10-15 | 289 | 1642 | ALW | rm lf | rm lf | clr | lf | * | * | Wingbird anemometer prop is iced, whole body of wingbird is spinning on bearing
2024-10-16 | 290 | 1820 | ALW | clr | rm lf | clr | lf | * | *
2024-10-17 | 291 | 1256 | ALW | rm lf | rm lf | clr | lf | * | *
2024-10-18 | 292 | 1323 | ALW | rm lf | rm lf | clr | lf | * | * | Wingbird anemometer prop is iced and whole body of wingbird is spinning on bearing.
2024-10-19 | 293 | 1159 | ALW | rm lf | rm lf | clr | lf | * | *
2024-10-20 | 294 | 1450 | ALW | rm lf | rm lf | clr | lf | * | *
2024-10-21 | 295 | 1530 | ALW | rm lf | rm lf | clr | lf | * | *
2024-10-22 | 296 | 1309 | ALW | clr | clr | clr | lf | * | *
2024-10-23 | 297 | 1514 | ALW | rm lf | rm lf | clr | lf | * | *
2024-10-24 | 298 | 1542 | ALW | rm lf | rm lf | clr | lf | * | *
2024-10-25 | 299 | 1153 | ALW | rm lf | rm lf | clr | clr | * | * | Heavy winds last night
2024-10-26 | 300 | 1217 | ALW | rm lf | rm lf | clr | lf | * | * | Wingbird anemometer prop is iced and whole body of wingbird is spinning on bearing.
2024-10-27 | 301 | 1538 | ALW | rm lf | rm lf | clr | lf | * | *
2024-10-28 | 302 | 1442 | ALW | rm lf | rm lf | clr | lf | * | *
2024-10-29 | 303 | 1109 | ALW | - | - | - | - | * | * | Con 1 - didn't go out to AWO.
2024-10-30 | 304 | 1250 | ALW | rm lf | rm lf | rm lf | lf | * | *
2024-10-31 | 305 | 1630 | ALW | rm lf | rm lf | clr | lf | * | *
2024-11-01 | 306 | 1330 | ALW | rm lf | clr | clr | lf | * | *
2024-11-02 | 307 | - | ALW | - | - | - | - | * | * | ICESat
2024-11-03 | 308 | 1410 | ALW | rm lf | rm lf | clr | lf | * | *
2024-11-04 | 309 | 1436 | ALW | clr | rm lf | clr | lf | * | *
2024-11-05 | 310 | 1330 | ALW | rm lf | rm lf | clr | lf | * | *
2024-11-06 | 311 | 1400 | ALW | clr | clr | clr | clr | * | *
2024-11-07 | 312 | 1912 | ALW | rm lf | rm lf | rm lf | rm lf | * | * | Conducted tower climb today 1600-1912
2024-11-08 | 313 | 1600 | ALW | rm lf | rm lf | clr | clr | * | *
2024-11-09 | 314 | 1230 | ALW | clr | rm lf | clr | clr | * | *
2024-11-10 | 315 | 1350 | ALW | rm lf | rm lf | clr | clr | * | *
2024-11-11 | 316 | 1230 | ALW | clr | clr | clr | clr | * | *
2024-11-12 | 317 | 1652 | ALW | clr | clr | clr | clr | * | *
2024-11-13 | 318 | - | ALW | - | - | - | - | * | *
2024-11-14 | 319 | 1618 | ALW | rm lf | rm lf | clr | clr | * | *
2024-11-15 | 320 | 1410 | ALW | rm lf | rm lf | clr | lf | * | *
2024-11-16 | 321 | 1335 | ALW | rm lf | rm lf | lf | lf | * | *
2024-11-17 | 322 | 1540 | ALW | rm lf | rm lf | lf | lf | * | * | 2mTD came unclasped. Was reclasped
2024-11-18 | 323 | 1618 | ALW | rm lf | rm lf | lf | lf | * | *
2024-11-19 | 324 | 1618 | ALW | clr | rm lf | lf | lf | * | *
2024-11-20 | 325 | 1541 | ALW | rm lf | rm lf | lf | lf | * | *
2024-11-21 | 326 | 1438 | ALW | rm lf | rm lf | lf | lf | * | *
2024-11-22 | 327 | 1415 | ALW | rm lf | rm lf | lf | lf | * | * | Wingbird anemometer prop is iced and whole body of wingbird is spinning on bearing.
2024-11-23 | 328 | 1330 | ALW | rm lf | rm lf | lf | lf | * | *
2024-11-24 | 329 | 1521 | ALW | rm lf | rm lf | lf | lf | * | *
2024-11-25 | 330 | - | ALW | - | - | - | - | - | - | GML page won't load
2024-11-26 | 331 | 1739 | ALW | rm lf | rm lf | lf | lf | * | *
2024-11-27 | 332 | - | ALW | - | - | - | - | - | - | Con 1 - didn't go out to AWO. GML won't load.
2024-11-28 | 333 | - | ALW | - | - | - | - | * | * | Holiday- did not go to AWO.
2024-11-29 | 334 | 1549 | ALW | rm lf | rm lf | lf | lf | * | * | T logger did go below freezing around midnight.
2024-11-30 | 335 | 1320 | ALW | rm lf | rm lf | lf | lf | * | *
2024-12-01 | 336 | 1612 | ALW | - | - | - | - | * | * | T has dropped below 0
2024-12-02 | 337 | 1320 | ALW | rm lf | rm lf | lf | lf | * | * | Couldn't get enclosure open to plug in heater
2024-12-03 | 338 | 1605 | ALW | rm lf | rm lf | lf | lf |  |  | GML won't update past 2024-12-03 1933Z
2024-12-04 | 339 | - | ALW | - | - | - | - | * | * | ICESAT- did not go to AWO
2024-12-05 | 340 | - | ALW | - | - | - | - | * | * | Sick- did not go to AWO
2024-12-06 | 341 | 1642 | ALW | rm lf | rm lf | lf | lf | * | *
2024-12-07 | 342 | 1308 | ALW | rm lf | rm lf | lf | lf |  |  | GML won't load
2024-12-08 | 343 | 1503 | ALW | clr | clr | clr | clr | * | *
2024-12-09 | 344 | 1309 | ALW | rm lf | rm lf | lf | lf | * | *
2024-12-10 | 345 | 1703 | ALW | rm lf | rm lf | lf | lf | * | *
2024-12-11 | 346 | 1533 | ALW | rm lf | rm lf | lf | lf | * | *
2024-12-12 | 347 | 1647 | ALW | rm lf | rm lf | lf | lf | * | *
2024-12-13 | 348 | 1308 | ALW | clr | clr | lf | lf | * | *
2024-12-14 | 349 | 1237 | ALW | rm lf | rm lf | lf | lf | * | *
2024-12-15 | 350 | 1248 | ALW | rm lf | rm lf | lf | lf | * | *
2024-12-16 | 351 | 1802 | ALW | rm lf | rm lf | lf | lf | * | *
2024-12-17 | 352 | 1334 | ALW | rm lf | rm lf | lf | lf |  |  | GML won't load
2024-12-18 | 353 | 1740 | ALW | rm lf | rm lf | lf | lf | * | *
2024-12-19 | 354 | 1627 | ALW | rm f | rm lf | lf | lf | * | *
2024-12-20 | 355 | 1643 | ALW | rm lf | rm lf | lf | lf | * | *
2024-12-21 | 356 | - | ALW | - | - | - | - | - | - | Con1- did not go to AWO
2024-12-22 | 357 | 1555 | ALW | rm lf | rm lf | lf | lf | - | - | GML won't load
2024-12-23 | 358 | 1805 | ALW | rm lf | rm lf | lf | lf | * | *
2024-12-24 | 359 | 1255 | ALW | rm lf | rm lf | lf | lf | * | *
2024-12-25 | 360 | 1635 | ALW | rm lf | rm lf | lf | lf | * | *
2024-12-26 | 361 | 1725 | ALW | rm lf | rm lf | lf | lf | * | *
2024-12-27 | 362 | 1507 | ALW | rm lf | rm lf | lf | lf | * | *
2024-12-28 | 363 | 1340 | ALW | rm hf | rm hf | lf | lf | * | * | Removed top shields and cleared out built up frost from fans 1340-1348Z.
2024-12-29 | 364 | 1445 | ALW | rm lf | rm lf | lf | lf | * | *
2024-12-30 | 365 | 1548 | ALW | rm lf | rm lf | lf | lf | * | *
2024-12-31 | 366 | 1505 | ALW | rm lf | rm lf | lf | lf | * | *
Note on E1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Shorter enclosure Reply: Aspirated temperature (short)
Note on F1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Longer enclosure Reply: Aspirated humidity (longer)
Note on G1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Ventus 2D Sonic Anemometer (V2) Reply: On NOAA boom with Wingbird
Note on H1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Heated Spinning Anemometer Reply: Secondary anemometer. There to record data if 9m Sonic primary anemometer is under maintenance. Wingbird on NOAA boom.
Note on I1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Refer to NOAA website T Logger Reply: NOAA enclosure (power enclosure on tower)- heater on and working Reply: NOAA enclosure - heater on and working but did not compensate with cooler temperatures very well. PI (Andrew Clarke) says that > -20 should be no cause for concern. Reply: Might have a fan running with a little heat coming out the top.
Note on J1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: NOAA website

## Inlet
Date | Julian Day | Time (UTC) | Observer's Initials | Neph/ Clap inlet | HATS inlet | Prototype inlet deicer | Sfc O3 inlet | Comments: climbs, filter changes, instrument work
2024-01-01 | 1 | 1525 | GM | rm lf | lf | - | rm lf
2024-01-02 | 2 | 1215 | GM | rm lf | lf | - | rm lf | Changed Sfc O3 particle filter.
2024-01-03 | 3 | 1530 | GM | rm lf | lf | - | rm lf
2024-01-04 | 4 | 1120 | GM | rm lf | lf | - | rm lf
2024-01-05 | 5 |  |  |  |  | - |  | Did not make it out to AWO after ICESat
2024-01-06 | 6 | 1255 | GM | rm lf | rm lf | - | rm lf
2024-01-07 | 7 |  |  |  |  | - |  | Did not go on roof due to 30kn winds
2024-01-08 | 8 |  |  |  |  | - |  | Did not go on roof due to 30kn winds
2024-01-09 | 9 | 1625 | GM | hf | clr | - | rm lf | Neph/clap inlet is clogged but it's too windy to take pot off and clean it.
2024-01-10 | 10 |  |  |  |  | - |  | Did not go on roof due to 30kn winds
2024-01-11 | 11 | 1300 | GM | rm hf | clr | - | rm lf | Took pot off inlet and cleared out snow clog.
2024-01-12 | 12 | 1530 | GM | clr | lf | - | rm lf
2024-01-13 | 13 | 1225 | GM | rm lf | lf | - | rm lf
2024-01-14 | 14 | 1450 | GM | rm lf | lf | - | rm lf
2024-01-15 | 15 | 1425 | GM | rm lf | lf | - | rm lf | The Snow Beast was ran for about 15min on station at the beginning of North winds 1600-1615Z
2024-01-16 | 16 | 1235 | GM | rm lf | lf | - | rm lf
2024-01-17 | 17 | 1530 | GM | clr | lf | - | rm lf
2024-01-18 | 18 | 1645 | GM | clr | lf | - | rm lf
2024-01-19 | 19 | 1220 | GM | rm lf | lf | - | rm lf
2024-01-20 | 20 | 1145 | GM | rm lf | lf | - | rm lf
2024-01-21 | 21 | 1705 | GM | rm lf | lf | - | rm lf
2024-01-22 | 22 | 1335 | GM | clr | lf | - | clr
2024-01-23 | 23 | 1225 | GM | clr | rm lf | - | rm lf
2024-01-24 | 24 | 1150 | GM | clr | rm lf | - | rm lf
2024-01-25 | 25 | 1630 | GM | rm lf | clr | - | rm lf
2024-01-26 | 26 | 1230 | GM | clr | lf | - | rm lf
2024-01-27 | 27 | 1300 | GM | clr | lf | - | rm lf
2024-01-28 | 28 | 1345 | GM | rm lf | lf | - | clr
2024-01-29 | 29 | 1310 | GM | rm lf | lf | - | rm lf
2024-01-30 | 30 | 1520 | GM | rm lf | lf | - | rm lf
2024-01-31 | 31 | 1600 | GM | rm lf | lf | - | rm lf
2024-02-01 | 32 | 1300 | GM | rm lf | lf | - | rm lf
2024-02-02 | 33 | [6 empty] | Did not make it out to AWO after ICESat
2024-02-03 | 34 | 1240 | GM | rm lf | lf | - | rm lf
2024-02-04 | 35 | 1450 | GM | rm lf | lf | - | rm lf
2024-02-05 | 36 | 1335 | GM | rm lf | lf | - | rm lf
2024-02-06 | 37 | 1630 | GM | rm lf | lf | - | rm lf
2024-02-07 | 38 | 1305 | GM | rm lf | rm lf | - | rm lf
2024-02-08 | 39 | 1920 | GM | rm lf | clr | - | rm lf
2024-02-09 | 40 | 1425 | GM DR | rm lf | clr | - | rm lf
2024-02-10 | 41 | 1310 | DR GM | rm lf | rm lf | - | rm lf
2024-02-11 | 42 | 1325 | DR GM | rm lf | rm lf | - | rm lf
2024-02-12 | 43 | 1320 | DR GM | rm lf | clr | - | rm lf
2024-02-13 | 44 | [6 empty] | Did not make it out to AWO after ICESat
2024-02-14 | 45 | [6 empty] | Did not make it out to AWO due to Con 1 conditions
2024-02-15 | 46 | 1345 | GM IG | rm hf | rm lf | - | rm lf
2024-02-16 | 47 | 1315 | DR | rm lf | rm lf | - | rm lf
2024-02-17 | 48 | 1825 | DR | rm lf | rm lf | - | rm lf
2024-02-18 | 49 | 1435 | DR | rm lf | rm lf | - | rm hf
2024-02-19 | 50 | 1325 | DR | rm lf | rm lf | - | rm lf
2024-02-20 | 51 | 1640 | DR | rm hf | rm hf | - | rm hf
2024-02-21 | 52 | 1530 | DR | rm hf | rm hf | - | rm lf
2024-02-22 | 53 | 1600 | DR | rm hf | rm lf | - | rm lf
2024-02-23 | 54 | 1315 | DR | rm hf | rm lf | - | rm lf
2024-02-24 | 55 | 1250 | DR | rm lf | rm lf | - | rm lf
2024-02-25 | 56 | 1700 | DR | rm lf | rm lf | - | rm lf | Light roof shoveling
2024-02-26 | 57 | 1340 | DR | rm lf | rm lf | - | rm lf
2024-02-27 | 58 | 1355 | DR | rm lf | rm lf | - | rm lf
2024-02-28 | 59 | 1310 | DR | rm lf | rm lf | - | rm lf | Light roof shoveling
2024-02-29 | 60 | 1530 | DR | rm lf | clr | - | rm lf
2024-03-01 | 61 | 1420 | DR | rm lf | rm lf | - | clr
2024-03-02 | 62 | 1100 | DR | rm lf | clr | - | rm lf
2024-03-03 | 63 | 1800 | DR | rm lf | clr | - | clr
2024-03-04 | 64 | 1610 | DR | rm lf | clr | - | rm lf
2024-03-05 | 65 | 1440 | DR | rm lf | clr | - | clr
2024-03-06 | 66 | 1315 | DR | rm lf | clr | - | rm lf
2024-03-07 | 67 | 1620 | DR | rm lf | clr | - | clr
2024-03-08 | 68 | 1205 | DR | rm lf | rm lf | - | rm lf | 1320 for Neph/Clap Inlet
2024-03-09 | 69 | 1315 | DR | rm lf | clr | - | clr
2024-03-10 | 70 | 1555 | DR | rm lf | clr | - | rm lf
2024-03-11 | 71 | 1600 | DR | rm lf | clr | - | rm lf
2024-03-12 | 72 |  | DR |  |  | - |  | Did not make it out to AWO after ICESat
2024-03-13 | 73 | 1350 | DR | rm lf | rm lf | - | rm lf
2024-03-14 | 74 | 1800 | DR | rm lf | clr | - | clr
2024-03-15 | 75 | 1445 | DR | rm lf | clr | - | rm lf
2024-03-16 | 76 | 1400 | DR | rm lf | clr | - | rm lf
2024-03-17 | 77 | 1935 | DR | rm lf | clr | - | rm lf
2024-03-18 | 78 | 1855 | DR | rm lf | clr | - | clr
2024-03-19 | 79 |  | DR |  |  | - |  | Did not make it out to AWO due to CON1
2024-03-20 | 80 | 2050 | DR | rm hf | rm lf | - | clr
2024-03-21 | 81 | 1645 | DR | rm lf | clr | - | clr
2024-03-22 | 82 | 1305 | DR | rm lf | lf | - | rm lf
2024-03-23 | 83 | 1750 | DR | rm lf | clr | - | rm lf
2024-03-24 | 84 | 1625 | DR | rm lf | clr | - | rm lf
2024-03-25 | 85 | 1700 | DR | rm lf | clr | - | rm lf
2024-03-26 | 86 | 1427 | DR | rm lf | clr | - | clr
2024-03-27 | 87 | 1835 | DR | rm lf | clr | - | rm lf
2024-03-28 | 88 | 1700 | DR | clr | clr | - | rm lf
2024-03-29 | 89 | 1340 | DR | rm lf | lf | - | rm lf
2024-03-30 | 90 | 1345 | DR | rm lf | clr | - | rm lf
2024-03-31 | 91 | 1555 | DR | rm lf | clr | - | rm lf
2024-04-01 | 92 | 1550 | DR | rm lf | rm lf | - | rm lf
2024-04-02 | 93 | 1640 | DR | clr | clr | - | clr
2024-04-03 | 94 |  | DR |  |  | - |  | Did not make it out to AWO after ICESat
2024-04-04 | 95 | 1300 | DR | rm lf | lf | - | rm lf
2024-04-05 | 96 | 1255 | DR | rm hf | rm lf | - | rm lf
2024-04-06 | 97 | 1250 | DR | rm lf | clr | - | rm lf
2024-04-07 | 98 | 1600 | DR | rm lf | clr | - | rm lf
2024-04-08 | 99 | 1550 | DR | rm lf | rm lf | - | rm lf
2024-04-09 | 100 | 1725 | DR | rm lf | clr | - | clr
2024-04-10 | 101 | 1505 | DR | rm lf | clr | - | clr
2024-04-11 | 102 | 1248 | DR | rm hf | rm lf | - | rm lf
2024-04-12 | 103 | 1245 | DR | rm lf | clr | - | rm lf
2024-04-13 | 104 | 1145 | DR | rm lf | clr | - | rm lf
2024-04-14 | 105 | 1915 | DR | rm lf | clr | - | rm lf
2024-04-15 | 106 | 1640 | DR | rm lf | clr | - | clr
2024-04-16 | 107 | 1245 | DR | rm lf | clr | - | clr
2024-04-17 | 108 | 2055 | DR | rm lf | clr | - | rm lf
2024-04-18 | 109 | 1520 | DR | rm lf | clr | - | rm lf
2024-04-19 | 110 | 1755 | DR | rm lf | clr | - | clr
2024-04-20 | 111 | 1625 | DR | rm lf | clr | - | clr
2024-04-21 | 112 | 1350 | DR | rm hf | clr | - | rm lf
2024-04-22 | 113 | 1135 | DR | rm lf | clr | - | clr | Shoveled roof
2024-04-23 | 114 | 1120 | DR | clr | clr | - | clr | Shoveled roof
2024-04-24 | 115 | 1600 | DR | rm lf | clr | - | clr
2024-04-25 | 116 |  | DR |  |  | - |  | HM - did not make it out to AWO.
2024-04-26 | 117 | 1135 | DR | rm lf | clr | - | rm lf
2024-04-27 | 118 | 1235 | DR | rm lf | clr | - | clr
2024-04-28 | 119 | 1510 | DR | clr | clr | - | clr
2024-04-29 | 120 | 1810 | DR | rm lf | clr | - | rm lf
2024-04-30 | 121 | 1240 | DR | clr | clr | - | clr
2024-05-01 | 122 | 1235 | DR | clr | clr | - | clr
2024-05-02 | 123 | 1250 | DR | clr | clr | - | clr
2024-05-03 | 124 | 1950 | DR | clr | clr | - | rm lf | Took POPS enclosure down at 1453UTC
2024-05-04 | 125 | 1650 | DR | rm lf | clr | - | rm lf
2024-05-05 | 126 |  | DR |  |  | - |  | HW - did not make it out to AWO.
2024-05-06 | 127 | 1800 | DR | clr | clr | - | clr
2024-05-07 | 128 | 1730 | DR | rm lf | clr | - | rm lf
2024-05-08 | 129 | 1010 | DR | rm lf | rm lf | - | rm lf
2024-05-09 | 130 | 2000 | DR | rm lf | clr | - | rm lf
2024-05-10 | 131 |  | DR |  |  | - |  | Did not make it out to AWO after ICESat
2024-05-11 | 132 | 1705 | DR | clr | clr | - | rm lf
2024-05-12 | 133 | 1405 | DR | rm lf | lf | - | rm lf
2024-05-13 | 134 | 1755 | DR | clr | clr | - | clr
2024-05-14 | 135 | 1620 | DR | clr | clr | - | clr
2024-05-15 | 136 |  | DR |  |  | - |  | Did not make it out to AWO after ICESat
2024-05-16 | 137 | 1740 | DR | clr | clr | - | clr
2024-05-17 | 138 | 1230 | DR | clr | clr | - | clr
2024-05-18 | 139 | 1505 | DR | rm hf | clr | - | rm lf
2024-05-19 | 140 | 1505 | DR | rm lf | clr | - | rm lf
2024-05-20 | 141 | 1220 | DR | rm lf | clr | - | rm lf
2024-05-21 | 142 |  | DR |  |  | - |  | HM - did not make it out to AWO.
2024-05-22 | 143 | 1115 | DR | rm lf | clr | - | rm lf
2024-05-23 | 144 | 1215 | DR | rm lf | clr | - | rm lf
2024-05-24 | 145 | 945 | DR | rm lf | rm lf | - | rm lf
2024-05-25 | 146 | 1600 | DR | clr | clr | - | clr
2024-05-26 | 147 | 1220 | DR | clr | clr | - | clr | Unplanned power outage - storm conditions.
2024-05-27 | 148 | 800 | DR | clr | clr | - | clr
2024-05-28 | 149 |  | DR |  |  | - |  | Day off.
2024-05-29 | 150 |  | DR |  |  | - |  | Con 1 - didn't go out to AWO.
2024-05-30 | 151 | 1450 | DR | clr | clr | - | clr
2024-05-31 | 152 |  | DR |  |  | -
2024-06-01 | 153 | 1015 | DR | rm lf | clr | - | rm lf
2024-06-02 | 154 | 1105 | DR | rm lf | clr | - | rm lf
2024-06-03 | 155 | 1645 | DR | clr | clr | - | clr
2024-06-04 | 156 |  | DR |  |  | - |  | Did not make it out to AWO after ICESat
2024-06-05 | 157 | 1100 | DR | rm lf | lf | - | rm lf
2024-06-06 | 158 | 1645 | CW | clr | clr | - | clr
2024-06-07 | 159 | 1500 | CW | clr | clr | - | clr | Powered down surface Ozone instrument for welding operations at 1458 UTC, bagged and sealed inlet.
2024-06-08 | 160 | 835 | CW | clr | clr | - | clr
2024-06-09 | 161 | 1145 | CW | clr | clr | - | clr
2024-06-10 | 162 | 935 | CW | clr | clr | - | clr
2024-06-11 | 163 | 1100 | CW | clr | clr | - | clr
2024-06-12 | 164 | 1050 | CW | clr | clr | - | clr
2024-06-13 | 165 | 1034 | CW | clr | clr | - | clr
2024-06-14 | 166 | 1005 | CW | clr | clr | - | clr
2024-06-15 | 167 | 1245 | CW | clr | clr | - | clr | Removed cap from Surface ozone inlet and powered on instrument.
2024-06-16 | 168 | 1518 | CW | clr | clr | - | clr
2024-06-17 | 169 | 1235 | CW | clr | clr | - | clr
2024-06-18 | 170 | 1018 | CW | rm lf | clr | - | rm lf
2024-06-19 | 171 | 1516 | CW | rm lf | clr | - | rm lf
2024-06-20 | 172 |  | CW |  |  |  |  | Did not go onto AWO roof due to lift operations.
2024-06-21 | 173 |  | CW |  |  |  |  | Did not go onto AWO roof due to lift operations.
2024-06-22 | 174 |  | CW |  |  |  |  | Did not make it to AWO, waiting for completion of stairs.
2024-06-23 | 175 | 1404 | CW | rm lf | clr | - | rm lf | 1405 O3 inlet.
2024-06-24 | 176 | 1413 | CW | rm lf | clr | - | rm lf | Brushed O3 inlet at 1614. cleared Neph/CLAP inlet at 1613.
2024-06-25 | 177 | 1559 | CW | rm lf | clr | - | rm lf
2024-06-26 | 178 | 1110 | CW | clr | clr | - | clr
2024-06-27 | 179 | 1148 | CW | clr | clr | - | rm lf
2024-06-28 | 180 | 1030 | CW | clr | clr | - | clr
2024-06-29 | 181 | 1101 | CW | clr | clr | - | clr | Power outage during UPS maintenance, UPS instruments affected from 1128-1134Z.
2024-06-30 | 182 | 1250 | CW | clr | clr | - | clr
2024-07-01 | 183 |  | CW |  |  |  |  | Did not go onto AWO roof due to High winds
2024-07-02 | 184 | 1615 | CW | clr | clr | - | clr | Changed O3 inlet filter at 1615.
2024-07-03 | 185 | 1015 | CW | clr | clr | - | clr
2024-07-04 | 186 |  | CW |  |  |  |  | Did not make it to AWO after ICESAT.
2024-07-05 | 187 | 1710 | CW | clr | clr | - | clr
2024-07-06 | 188 | 1805 | CW | clr | clr | - | clr
2024-07-07 | 189 | 1246 | CW | rm lf | clr | - | clr
2024-07-08 | 190 | 1038 | CW | rm lf | clr | - | rm lf | 1038 - brushed aerolosol inlet. 1039 - brushed O3.
2024-07-09 | 191 | 1114 | CW | rm lf | clr | - | rm lf | 1114 - brushed aerosol inlet. 1015 - brushed O3 inlet.
2024-07-10 | 192 | 1445 | CW | rm lf | clr | - | rm lf
2024-07-11 | 193 | 1211 | CW | clr | clr | - | clr
2024-07-12 | 194 | 1510 | CW | clr | clr | - | clr
2024-07-13 | 195 | 1455 | CW | clr | clr | - | clr
2024-07-14 | 196 | 1702 | CW | clr | clr | - | clr
2024-07-15 | 197 | 1124 | CW | clr | clr | - | clr
2024-07-16 | 198 | 1241 | CW | clr | clr | - | clr | 1108-1248: Digging around Met tower in preparation for tightening of guy lines.
2024-07-17 | 199 | 1436 | CW | clr | clr | - | clr
2024-07-18 | 200 | 1726 | CW | rm lf | clr | - | rm lf
2024-07-19 | 201 | 1739 | CW | clr | clr | - | clr
2024-07-20 | 202 | 1258 | CW | clr | clr | - | rm lf
2024-07-21 | 203 | 1009 | CW | rm lf | clr | - | rm lf
2024-07-22 | 204 | 1255 | CW | rm lf | clr | - | rm lf | rm lf from neph & O3 inlet at 1259Z.
2024-07-23 | 205 | 1547 | CW | rm lf | clr | - | rm lf
2024-07-24 | 206 | 1510 | CW | rm lf | clr | - | clr
2024-07-25 | 207 | 1130 | CW | clr | clr | - | clr
2024-07-26 | 208 | 1551 | CW | rm lf | clr | - | clr
2024-07-27 | 209 | 1041 | CW | clr | clr | - | clr
2024-07-28 | 210 | 1204 | CW | clr | clr | - | clr
2024-07-29 | 211 | 1227 | CW | clr | clr | - | clr
2024-07-30 | 212 | 1400 | CW | clr | clr | - | clr
2024-07-31 | 213 | 1709 | CW | clr | clr | - | clr
2024-08-01 | 214 | 1455 | CW | rm lf | clr | - | rm lf
2024-08-02 | 215 | 1628 | CW | rm lf | clr | - | rm lf
2024-08-03 | 216 | 1558 | CW | rm lf | clr | - | rm lf | 1506 brushed O3. 1557 brushed neph/CLAP.
2024-08-04 | 217 | 1742 | CW | rm lf | clr | - | rm lf | Brushed neph inlet at 1743. Brushed O3 inlet at 1744.
2024-08-05 | 218 |  | CW |  |  |  |  | 1049-1217: shut down carbon vane pump to change inlet. Did not make it to AWO after ICESat.
2024-08-06 | 219 | 1128 | CW | rm lf | lf | - | rm lf | Brushed aerosol inlet at 1128Z. Brushed O3 inlet at 1131.
2024-08-07 | 220 | 1130 | CW | rm lf | clr | - | rm lf | Brushed O3 inlet at 1130. Brushed aerosol inlet at 1131.
2024-08-08 | 221 | 1025 | CW | clr | clr | - | clr
2024-08-09 | 222 | 1204 | CW | clr | clr | - | rm lf | Noticed two gaps in data from 2055-2211 & 2234-2337z
2024-08-10 | 223 | 1054 | CW | rm lf | lf | - | rm lf | Brushed aerosol inlet at 1054z. Brushed O3 inlet at 1053z.
2024-08-11 | 224 | 1843 | CW | clr | clr | - | clr
2024-08-12 | 225 | 1150 | CW | rm lf | clr | - | clr
2024-08-13 | 226 | 1000 | CW | clr | clr | - | rm lf
2024-08-14 | 227 |  | CW |  |  |  |  | ICESat
2024-08-15 | 228 | 1512 | CW | clr | clr | - | clr
2024-08-16 | 229 | 1004 | CW | rm lf | clr | - | clr
2024-08-17 | 230 | 1518 | CW | clr | clr | - | clr
2024-08-18 | 231 |  | CW |  |  |  |  | Bear protocols, did not go to awo
2024-08-19 | 232 | 1528 | CW | clr | clr | - | rm lf
2024-08-20 | 233 | 1356 | CW | rm lf | clr | - | rm lf
2024-08-21 | 234 | 1215 | CW | clr | clr | - | rm hf
2024-08-22 | 235 | 1906 | CW | clr | clr | - | rm lf
2024-08-23 | 236 | 905 | CW | rm lf | lf | - | rm lf
2024-08-24 | 237 | 1030 | CW | rm lf | lf | - | rm lf
2024-08-25 | 238 | 1417 | CW | rm lf | lf | - | rm lf
2024-08-26 | 239 | 1215 | CW | clr | lf | - | clr
2024-08-27 | 240 | 1240 | CW | clr | lf | - | rm lf
2024-08-28 | 241 | 1355 | CW | rm lf | lf | - | rm lf
2024-08-29 | 242 | 1337 | CW | clr | clr | - | clr
2024-08-30 | 243 | 917 | CW | clr | clr | - | clr
2024-08-31 | 244 | 1222 | CW | clr | clr | - | clr
2024-09-01 | 245 | 1252 | CW | clr | clr | - | clr
2024-09-02 | 246 | 1643 | CW | clr | clr | - | clr
2024-09-03 | 247 |  | CW |  |  |  |  | Did not make it to AWO after ICESat.
2024-09-04 | 248 | 1432 | CW | clr | clr | - | rmlf
2024-09-05 | 249 | 1235 | CW | rm lf | clr | - | rm lf
2024-09-06 | 250 | 1416 | CW | clr | clr | - | rm lf
2024-09-07 | 251 |  | CW |  |  |  |  | Did not go to AWO due to poor visibility
2024-09-08 | 252 | 1850 | CW | clr | clr | - | clr
2024-09-09 | 253 | 1235 | CW | clr | clr | - | clr
2024-09-10 | 254 |  | CW |  |  |  |  | Did not go to AWO due to poor visibility
2024-09-11 | 255 | 1321 | CW | rm hf | rm hf | - | rm hf
2024-09-12 | 256 | 912 | CW | rm hf | hf | - | rm hf
2024-09-13 | 257 | 1012 | CW | rm hf | rm hf | - | rm hf
2024-09-14 | 258 | 1102 | CW | rm lf | lf | - | rm lf
2024-09-15 | 259 | 1136 | CW | clr | clr | - | rm lf
2024-09-16 | 260 | 1201 | CW | rm lf | clr | - | rm lf
2024-09-17 | 261 | 1752 | CW | rm lf | rm lf | - | rm lf
2024-09-18 | 262 | 1439 | CW | rm hf | lf | - | rm lf
2024-09-19 | 263 | 1215 | CW | clr | clr | - | rm lf
2024-09-20 | 264 | 1553 | CW | clr | clr | - | rm lf
2024-09-21 | 265 | 1008 | CW | clr | clr | - | rm lf
2024-09-22 | 266 | 1252 | CW | clr | clr | - | clr
2024-09-23 | 267 | 1325 | CW | clr | clr | - | clr
2024-09-24 | 268 | 1050 | CW | lf | lf | - | lf
2024-09-25 | 269 | 1414 | CW | rm hf | rm lf | - | rm hf
2024-09-26 | 270 | 1442 | CW | rm hf | rm lf | - | rm hf
2024-09-27 | 271 | 956 | CW | rm lf | lf | - | rm lf
2024-09-28 | 272 |  | CW |  |  |  |  | Did not make it to AWO after raising sector lines
2024-09-29 | 273 | 1259 | CW | rm lf | lf | - | rm lf
2024-09-30 | 274 | 1109 | CW | clr | clr | - | clr
2024-10-01 | 275 | 1242 | CW | rm hf | lf | - | rm lf
2024-10-02 | 276 | 1505 | CW | rm hf | rm lf | - | rm lf
2024-10-03 | 277 | 1458 | CW | rm lf | clr | - | clr
2024-10-04 | 278 | 1652 | CW | clr | clr | - | clr
2024-10-05 | 279 |  | CW |  |  |  |  | Did not go to AWO after ICESat
2024-10-06 | 280 | 1720 | CW | rm lf | clr | - | rm lf
2024-10-07 | 281 | 1325 | CW | rm hf | clr | - | rm lf
2024-10-08 | 282 | 1010 | CW | rm lf | lf | - | rm lf
2024-10-09 | 283 | 1209 | CW | rm lf | lf | - | rm lf
2024-10-10 | 284 | 1127 | CW | rm lf | lf | - | rm lf
2024-10-11 | 285 | 1321 | CW | rm lf | lf | - | rm lf
2024-10-12 | 286 | 1328 | CW | rm lf | lf | - | clr
2024-10-13 | 287 | 1411 | ALW | rm lf | rm lf | - | clr
2024-10-14 | 288 | 1648 | ALW | rm lf | clr | - | rm lf
2024-10-15 | 289 | 1506 | ALW | rm lf | rm lf | - | rm lf
2024-10-16 | 290 | 1738 | ALW | rm lf | lf | - | rm lf
2024-10-17 | 291 | 1202 | ALW | rm lf | lf | - | rm lf
2024-10-18 | 292 | 1242 | ALW | rm lf | lf | - | rm lf
2024-10-19 | 293 | 1114 | ALW | rm lf | lf | - | rm lf
2024-10-20 | 294 | 1522 | ALW | rm lf | lf | - | rm lf | Sfc 03 might have frost inside the tube
2024-10-21 | 295 | 1556 | ALW | rm lf | rm lf | - | rm lf
2024-10-22 | 296 | 1407 | ALW | rm lf | lf | - | rm lf
2024-10-23 | 297 | 1751 | ALW | rm lf | lf | - | rm lf
2024-10-24 | 298 | 1502 | ALW | rm lf | rm lf | - | rm lf
2024-10-25 | 299 | 1250 | ALW | rm lf | clr | - | rm lf
2024-10-26 | 300 | 1259 | ALW | rm lf | lf | - | rm lf
2024-10-27 | 301 | 1455 | ALW | rm lf | rm lf | - | rm lf
2024-10-28 | 302 | 1530 | ALW | rm lf | rm lf | - | rm lf
2024-10-29 | 303 | - | ALW | - | - | - | - | Con 1 - didn't go out to AWO.
2024-10-30 | 304 | 1505 | ALW | rm lf | rm lf | - | rm lf
2024-10-31 | 305 | 1510 | ALW | rm lf | lf | - | rm lf
2024-11-01 | 306 | 1509 | ALW | rm lf | lf | - | rm lf
2024-11-02 | 307 | - | ALW | - | - | - | - | ICESat
2024-11-03 | 308 | 1515 | ALW | rm lf | lf | - | rm lf
2024-11-04 | 309 | - | ALW | - | - | - | - | Too windy to go on roof
2024-11-05 | 310 | 1436 | ALW | rm lf | lf | - | rm lf
2024-11-06 | 311 | - | ALW | - | - | - | - | Too windy to go on roof
2024-11-07 | 312 | - | ALW | - | - | - | - | Did not go to roof after tower climb
2024-11-08 | 313 | 1730 | ALW | rm lf | rm lf | - | rm lf
2024-11-09 | 314 | 1308 | ALW | rm lf | rm lf | - | rm lf
2024-11-10 | 315 | 1500 | ALW | rm lf | lf | - | rm lf
2024-11-11 | 316 | 1340 | ALW | rm lf | rm lf | - | rm lf
2024-11-12 | 317 | 1439 | ALW | rm lf | rm lf | - | rm lf
2024-11-13 | 318 | - | ALW | - | - | - | - | ICESat- did not go to AWO
2024-11-14 | 319 | - | ALW | - | - | - | - | Too windy to go on roof
2024-11-15 | 320 | 1320 | ALW | rm lf | rm lf | - | rm lf
2024-11-16 | 321 | 1215 | ALW | rm lf | rm lf | - | rm lf
2024-11-17 | 322 | 1345 | ALW | rm lf | rm lf | - | rm lf
2024-11-18 | 323 | 1538 | ALW | rm lf | rm lf | - | rm lf
2024-11-19 | 324 | 1447 | ALW | rm lf | rm lf | - | rm lf
2024-11-20 | 325 | 1508 | ALW | rm lf | rm lf | - | rm lf
2024-11-21 | 326 | 1602 | ALW | rm lf | rm lf | - | rm lf
2024-11-22 | 327 | 1353 | ALW | rm lf | rm lf | - | rm lf
2024-11-23 | 328 | 1225 | ALW | rm lf | lf | - | rm lf
2024-11-24 | 329 | 1352 | ALW | rm lf | rm lf | - | rm lf | HATS sampled: 1341-1505Z.
2024-11-25 | 330 | - | ALW | - | - | - | - | CON 1- did not go to AWO
2024-11-26 | 331 | 1720 | ALW | rm lf | lf | - | rm lf
2024-11-27 | 332 | - | ALW | - | - | - | - | CON 1- did not go to AWO
2024-11-28 | 333 | - | ALW | - | - | - | - | Holiday- did not go to AWO
2024-11-29 | 334 | 1459 | ALW | rm lf | rm lf | - | rm lf
2024-11-30 | 335 | 1300 | ALW | rm lf | rm lf | - | rm lf
2024-12-01 | 336 | - | ALW | - | - | - | -
2024-12-02 | 337 | 1223 | ALW | rm lf | rm lf | - | rm lf | Showed up to AWO with orange cap on the ground. Reattached 1225Z.
2024-12-03 | 338 | 1402 | ALW | rm lf | rm lf | - | rm lf
2024-12-04 | 339 | - | ALW | - | - | - | - | ICESAT- did not go to AWO
2024-12-05 | 340 | - | ALW | - | - | - | - | Sick- did not go to AWO
2024-12-06 | 341 | 1524 | ALW | rm lf | rm lf | - | rm lf
2024-12-07 | 342 | 1538 | ALW | rm lf | rm lf | - | rm lf
2024-12-08 | 343 | 1524 | ALW | rm lf | rm lf | - | rm lf | Showed up to AWO with orange cap on the ground. Reattached 1524Z.
2024-12-09 | 344 | 1447 | ALW | rm lf | lf | - | rm lf
2024-12-10 | 345 | 1344 | ALW | rm lf | rm lf | - | rm lf
2024-12-11 | 346 | - | ALW | - | - | - | - | Orange cap blown to the ground about 1530Z. Too windy to go to roof.
2024-12-12 | 347 | 1557 | ALW | rm lf | rm lf | - | reattached | Orange cap replaced @ 1612Z.
2024-12-13 | 348 | - | ALW | - | - | - | - | Too windy to go on roof
2024-12-14 | 349 | 1217 | ALW | rm lf | rm lf | - | rm lf
2024-12-15 | 350 | 1326 | ALW | rm lf | rm lf | - | rm lf
2024-12-16 | 351 | 1516 | ALW | rm lf | rm lf | - | rm lf
2024-12-17 | 352 | 1428 | ALW | rm lf | lf | - | rm lf
2024-12-18 | 353 | 1820 | ALW | reattached | lf | - | rm lf | CON 1- did not go to AWO
2024-12-19 | 354 | 1736 | ALW | rm lf | rm lf | - | rm lf
2024-12-20 | 355 | 1615 | ALW | rm lf | rm lf | - | rm lf
2024-12-21 | 356 | - | ALW | - | - | - | - | CON 1- did not go to AWO
2024-12-22 | 357 | 1512 | ALW | rm lf | rm lf | - | rm lf
2024-12-23 | 358 | 1653 | ALW | rm lf | rm lf | - | rm lf
2024-12-24 | 359 | 1210 | ALW | rm lf | rm lf | - | rm lf
2024-12-25 | 360 | 1412 | ALW | rm lf | rm lf | - | rm lf
2024-12-26 | 361 | 1642 | ALW | rm lf | rm lf | - | rm lf
2024-12-27 | 362 | 1342 | ALW | rm lf | rm lf | - | rm lf
2024-12-28 | 363 | 1215 | ALW | rm lf | rm lf | - | rm lf
2024-12-29 | 364 | 1343 | ALW | rm lf | rm lf | - | rm lf
2024-12-30 | 365 | 1455 | ALW | rm lf | rm lf | - | rm lf
2024-12-31 | 366 | 1400 | ALW | rm lf | rm lf | - | rm lf
Note on E1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Black and Gray Inlet Reply: Now has the pot over top Reply: aerosol inlet on s side railing Reply: Has mesh covering the intake to help keep rime out. Reach hand up pot and inside intake to remove rime from mesh.
Note on F1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Top Inlet- looks like bells Reply: On left side of mini tower on roof.
Note on G1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: On right side of mini tower on roof. Usually surrounded by rime. Leave be.
Note on H1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Orange Cap Inlet Reply: Has paperish filter inside to keep frost free and be replaced ~January and July 1st

## NEPH & CLAP
Date | Julian Day | Time (UTC) | Observer's Initials | Sample inlet? | Check Daily Aerosol Email | CLAP flow ~0.60 lpm | CLAP green tr>0.7 | CLAP Spot # | MAGIC 210 cm-3 > 0 | MAGIC 210 dewpoint > 10 | MAGIC 250 cm-3 > 0 | MAGIC 250 dewpoint > ? | Refilled MAGIC Reservoir | Monthly Span Check | NEPH Lamp On/Current (<7A) | NEPH Flag Status | NEPH Flow (~18 lpm) | Ecotech Flag status | Ecotech Flow (~4.5lpm) | Ecotech flag check on instrument (every other week) | Comments
2024-01-01 | 1 | 1612 | GM | rm lf | * | 0.6 | 0.91195 | 6 | Y | Y |  |  |  |  | 6 | 100 | 18
2024-01-02 | 2 |  |  | rm lf | * | [8 empty] | X | [6 empty] | Span Check: 1707-1738Z, Error:2.0%, Tank Pressure: 650psi
2024-01-03 | 3 | 1702 | GM | rm lf | * | 0.59 | 0.90902 | 6 | Y | Y |  |  | Y |  | 6 | 100 | 18
2024-01-04 | 4 |  |  | rm lf | *
2024-01-05 | 5 |  |  |  | * | [15 empty] | Did not make it out to AWO after ICESat
2024-01-06 | 6 | 1351 | GM | rm lf | * | 0.58 | 0.9073 | 6 | Y | Y |  |  |  |  | 6 | 100 | 18 |  |  |  | Adjusted clap flow from 0.58 to 0.60
2024-01-07 | 7 |  |  |  | *
2024-01-08 | 8 | 1730 | GM |  | * | 0.6 | 0.90516 | 6 | N | Y |  |  |  |  | 6 | 0000 | 18 |  |  |  | Inlet likely clogged but it's too windy to head up to the roof right now
2024-01-09 | 9 |  |  | hf | *
2024-01-10 | 10 | 1311 | GM |  | * | 0.61 | 0.90427 | 6 | N | Y |  |  |  |  | 6 | 0000 | 18 |  |  |  | Inlet clogged but it's too windy to go on roof. Adjusted CLAP flow from 0.61 to 0.60
2024-01-11 | 11 |  |  | rm hf | *
2024-01-12 | 12 | 1617 | GM | clr | * | 0.62 | 0.88949 | 6 | Y | Y |  |  |  |  | 6 | 100 | 18 |  |  |  | Adjusted clap flow from 0.62 to 0.60
2024-01-13 | 13 |  |  | rm lf | *
2024-01-14 | 14 | 1524 | GM | rm lf | * | 0.6 | 0.88337 | 6 | Y | Y |  |  |  |  | 6 | 100 | 18 |  |  |  | Magic is showing "Moderator (3rd stage) and Initiator (2nd stage): Temperature out of range" and "Status: initiator temperature out of range, moderator temperature out of range", but seems to have corrected itself after taking in some water from the reservoir.
2024-01-15 | 15 |  |  | rm lf | * | [15 empty] | The Snow Beast was ran for about 15min on station at the beginning of North winds 1600-1615Z
2024-01-16 | 16 | 1339 | GM | rm lf | * | 0.6 | 0.82179 | 6 | Y | Y |  |  |  |  | 6 | 100 | 18 |  |  |  | Magic is showing "Initiator (2nd stage): temperature out of range" again
2024-01-17 | 17 |  |  | clr | * | [15 empty] | Magic was restarted 1519-1521Z
2024-01-18 | 18 | 1656 | GM | clr | * | 0.57 | 0.81388 | 6 | Y | Y |  |  |  |  | 6 | 100 | 18 |  |  |  | Adjusted clap flow from 0.57 to 0.60
2024-01-19 | 19 |  |  | rm lf | *
2024-01-20 | 20 | 1158 | GM | rm lf | * | 0.57 | 0.80847 | 6 | Y | Y |  |  | Y |  | 6 | 100 | 18 |  |  |  | Adjusted clap flow from 0.57 to 0.60
2024-01-21 | 21 |  |  | rm lf | * | [15 empty] | Today's email noted there were 10 lost communications events, which is much higher than the usual none or 1
2024-01-22 | 22 | 1428 | GM | clr | * | 0.6 | 0.79586 | 6 | Y | Y |  |  |  |  | 6 | 100 | 18
2024-01-23 | 23 |  |  | clr | *
2024-01-24 | 24 | 1243 | GM | clr | * | 0.6 | 0.788 | 6 | Y | Y |  |  |  |  | 6 | 100 | 18
2024-01-25 | 25 |  |  | rm lf | *
2024-01-26 | 26 | 1311 | GM | clr | * | 0.6 | 0.7789 | 6 | Y | Y |  |  |  |  | 6 | 100 | 18
2024-01-27 | 27 |  |  | clr | *
2024-01-28 | 28 | 1244 | GM | rm lf | * | 0.0 | 0.7755 | 6 | Y | Y |  |  |  |  | 6 | 100 | 0 |  |  |  | Agnes Pump appears to have shut down, possibly from over heating (it's a bit hot to the touch). Unplugged it at 1248Z to let cool and investigate further. Plots show flow dropped to zero on 2024-01-28 around 0405Z. After cooling, pump was plugged back in and did start up but had an internal grinding sound, awaiting further thoughts from others before diving deeper.
2024-01-29 | 29 |  |  | rm lf | * | [15 empty] | Neph and CLAP powered down at 1739Z due to pump failure
2024-01-30 | 30 | 1258 | GM | rm lf | * | 0.6 | 0.77522 | 6 | Y | Y |  |  |  |  | 6 | 100 | 18 |  |  |  | Neph and CLAP powered on at 1230Z, diaphragm pump replaced with carbon vane pump. Neph flow adjusted to 18 lpm and CLAP flow adjusted to .60 lpm. On roof swapping inlet hats from pot to old type from 1630-1720Z.
2024-01-31 | 31 |  |  | rm lf | *
2024-02-01 | 32 | 1351 | GM | rm lf | * | 0.6 | 0.76955 | 6 | Y | Y |  |  |  |  | 6 | 100 | 18
2024-02-02 | 33 |  |  |  | *
2024-02-03 | 34 | 1256 | GM | rm lf | * | 0.61 | 0.76131 | 6 | Y | Y |  |  | Y | X | 6 | 100 | 18 |  |  |  | Adjusted Clap flow from 0.61 to 0.60. Span Check: 1329-1400Z, Error: 1.5%, Tank Pressure: 650psi
2024-02-04 | 35 |  |  | rm lf | *
2024-02-05 | 36 | 1421 | GM | rm lf | * | 0.61 | 0.75126 | 6 | Y | Y |  |  |  |  | 6 | 100 | 18 |  |  |  | Adjusted clap flow from 0.61 to 0.60
2024-02-06 | 37 |  |  | rm lf | *
2024-02-07 | 38 | 1433 | GM | rm lf | * | 0.61 | 0.74228 | 6 | Y | Y |  |  |  |  | 6 | 100 | 18 |  |  |  | Adjusted clap flow from 0.61 to 0.60
2024-02-08 | 39 |  |  | rm lf | *
2024-02-09 | 40 | 1534 | GM DR | rm lf | * | 0.59 | 0.73614 | 6 | Y | Y |  |  |  |  | 6 | 100 | 18 |  |  |  | Adjusted clap flow from 0.59 to 0.60
2024-02-10 | 41 |  |  | rm lf | *
2024-02-11 | 42 | 1430 | DR | rm lf | * | 0.59 | 0.73395 | 6 | Y | Y |  |  |  |  | 5.9 | 100 | 18
2024-02-12 | 43 |  |  | rm lf | * | [15 empty] | Clap filter change complete. Magic 200 powered down per Erin until further notice.
2024-02-13 | 44 |  |  |  | * | [15 empty] | Did not make it out to AWO after ICESat
2024-02-14 | 45 |  |  |  | * | [15 empty] | Did not make it out to AWO due to CON1
2024-02-15 | 46 | 1357 | GM | rm hf | * | 0.6 | 0.99102 | 1 | N/A | N/A |  |  | N/A |  | 6 | 100 | 18
2024-02-16 | 47 |  |  | rm lf | * | [15 empty] | Adjusted clap flow from 0.613 to 0.60
2024-02-17 | 48 | 1815 | DR | rm lf | * | 0.604 | 0.98147 | 1 | N/A | N/A |  |  | N/A |  | 6 | 100 | 18
2024-02-18 | 49 |  |  | rm lf | *
2024-02-19 | 50 | 1233 | DR | rm lf | * | 0.599 | 0.97738 | 1 | N/A | N/A |  |  | N/A |  | 6 | 100 | 18
2024-02-20 | 51 |  |  | rm hf | *
2024-02-21 | 52 | 1520 | DR | rm hf | * | 0.6 | 0.97608 | 1 | N/A | N/A |  |  | N/A |  | 6 | 100 | 18 |  |  |  | Adjusted clap flow from 0.587 to 0.60
2024-02-22 | 53 | 1613 | DR | rm hf | * | 0.6 | 0.96975 | 1 | N/A | N/A |  |  | N/A |  | 6 | 100 | 18
2024-02-23 | 54 | 1304 | DR | rm hf | * | 0.604 | 0.94092 | 1 | N/A | N/A |  |  | N/A |  | 6 | 100 | 18
2024-02-24 | 55 | 1259 | DR | rm lf | * | 0.601 | 0.93881 | 1 | N/A | N/A |  |  | N/A |  | 5.9 | 100 | 18
2024-02-25 | 56 | 1631 | DR | rm lf | * | 0.595 | 0.93299 | 1 | N/A | N/A |  |  | N/A |  | 6 | 100 | 18
2024-02-26 | 57 | 1245 | DR | rm lf | * | 0.6 | 0.92954 | 1 | N/A | N/A |  |  | N/A |  | 5.9 | 100 | 18 |  |  |  | Adjusted clap flow from 0.585 to 0.60
2024-02-27 | 58 | 1407 | DR | rm lf | * | 0.605 | 0.92329 | 1 | N/A | N/A |  |  | N/A |  | 5.9 | 100 | 18
2024-02-28 | 59 | 1259 | DR | rm lf | * | 0.603 | 0.9173 | 1 | N/A | N/A |  |  | N/A |  | 5.9 | 100 | 18
2024-02-29 | 60 | 1413 | DR | rm lf | * | 0.6 | 0.90815 | 1 | N/A | N/A |  |  | N/A |  | 5.9 | 100 | 18 |  |  |  | Adjusted clap flow from 0.613 to 0.60 (N winds from ground level)
2024-03-01 | 61 | 1306 | DR | rm lf | * | 0.6 | 0.88641 | 1 | N/A | N/A |  |  | N/A |  | 6 | 100 | 18 |  |  |  | Adjusted clap flow from 0.612 to 0.60
2024-03-02 | 62 | 1038 | DR | rm lf | * | 0.607 | 0.87894 | 1 | N/A | N/A |  |  | N/A |  | 5.9 | Zero | 18 |  |  |  | Very warm day (5°F)
2024-03-03 | 63 | 1831 | DR | rm lf | * | 0.602 | 0.8739 | 1 | N/A | N/A |  |  | N/A |  | 5.9 | 100 | 18 |  |  |  | Neph Span Check 1630-1823 Error: 39.1%, 38.6%, 60.6% (w/o co2) CO2 tank @ 720psi - 700psi - sent email
2024-03-04 | 64 | 1230 | DR | rm lf | * | 0.603 | 0.86921 | 1 | N/A | N/A |  |  | N/A | 1.5% | 5.9 | 100 | 18 |  |  |  | Neph Span Check 1529-1600 Error: 40.3% & 1814-1844 Error %: 1.5%
2024-03-05 | 65 | 1423 | DR | rm lf | * | 0.605 | 0.86724 | 1 | N/A | N/A |  |  | N/A |  | 6 | 100 | 18 |  |  |  | Adjusted clap flow from 0.609 to 0.60
2024-03-06 | 66 | 1300 | DR | rm lf | * | 0.6 | 0.86585 | 1 | N/A | N/A |  |  | N/A |  | 5.9 | 100 | 18
2024-03-07 | 67 | 1541 | DR | rm lf | * | 0.602 | 0.86363 | 1 | N/A | N/A |  |  | N/A |  | 5.9 | Zero | 18
2024-03-08 | 68 | 1338 | DR | rm lf | * | 0.603 | 0.86336 | 1 | N/A | N/A |  |  | N/A |  | 5.9 | 100 | 18
2024-03-09 | 69 | 1238 | DR | rm lf | * | 0.599 | 0.86252 | 1 | N/A | N/A |  |  | N/A |  | 5.9 | 100 | 18
2024-03-10 | 70 | 1539 | DR | rm lf | * | 0.598 | 0.85867 | 1 | N/A | N/A |  |  | N/A |  | 5.9 | 100 | 18
2024-03-11 | 71 | 1425 | DR | rm lf | * | 0.601 | 0.85642 | 1 | N/A | N/A |  |  | N/A |  | 5.9 | Zero | 18
2024-03-12 | 72 |  |  |  | * |  |  |  | N/A | N/A |  |  | N/A | [7 empty] | Did not make it out to AWO after ICESat
2024-03-13 | 73 | 1251 | DR | rm lf | * | 0.595 | 0.85152 | 1 | N/A | N/A |  |  | N/A |  | 5.9 | 100 | 18
2024-03-14 | 74 | 1631 | DR | rm lf | * | 0.599 | 0.85052 | 1 | N/A | N/A |  |  | N/A |  | 5.9 | 100 | 18 |  |  |  | Adjusted clap flow from 0.584 to 0.599
2024-03-15 | 75 | 1402 | DR | rm lf | * | 0.595 | 0.84922 | 1 | N/A | N/A |  |  | N/A |  | 5.9 | 100 | 18
2024-03-16 | 76 | 1315 | DR | rm lf | * | 0.595 | 0.84659 | 1 | N/A | N/A |  |  | N/A |  | 5.9 | 100 | 18
2024-03-17 | 77 | 1905 | DR | rm lf | * | 0.596 | 0.84054 | 1 | N/A | N/A |  |  | N/A |  | 5.9 | 100 | 18
2024-03-18 | 78 | 1741 | DR | rm lf | * | 0.598 | 0.83367 | 1 | N/A | N/A |  |  | N/A |  | 5.9 | 100 | 18
2024-03-19 | 79 |  |  |  | * |  |  |  | N/A | N/A |  |  | N/A | [7 empty] | Did not make it out to AWO due to CON1
2024-03-20 | 80 | 1128 | DR | rm hf | * | 0.599 | 0.82808 | 1 | N/A | N/A |  |  | N/A |  | 5.9 | 100 | 18 |  |  |  | Pump turned off from 1231-2052 for aerosol inlet maintenance - took long due to being locked out of building. Aerosol inlet cleared off snow buildup.
2024-03-21 | 81 | 1634 | DR | rm lf | * | 0.599 | 0.82204 | 1 | N/A | N/A |  |  | N/A |  | 5.9 | 100 | 18 |  |  |  | Adjusted clap flow from 0.590 to 0.599
2024-03-22 | 82 | 1216 | DR | rm lf | * | 0.6 | 0.81865 | 1 | N/A | N/A |  |  | N/A |  | 5.9 | 100 | 18 |  |  |  | Adjusted clap flow from 0.607 to 0.600. Received MAGIC humidifier.
2024-03-23 | 83 | 1802 | DR | rm lf | * | 0.599 | 0.81557 | 1 | N/A | N/A |  |  | N/A |  | 5.9 | 100 | 18 |  |  |  | Adjusted clap flow from 0.590 to 0.599
2024-03-24 | 84 | 1636 | DR | rm lf | * | 0.602 | 0.81077 | 1 | N/A | N/A |  |  | N/A |  | 5.9 | 100 | 18 |  |  |  | Adjusted clap flow from 0.590 to 0.602
2024-03-25 | 85 | 1544 | DR | rm lf | * | 0.603 | 0.80621 | 1 | N/A | N/A |  |  | N/A |  | 5.9 | 100 | 18
2024-03-26 | 86 | 1331 | DR | rm lf | * | 0.602 | 0.79426 | 1 | 126/131 | 18.6/-9.9 |  |  | Y |  | 5.9 | 100 | 18 |  |  |  | Both Magic210 & Magic250 are installed.
2024-03-27 | 87 |  | DR | rm lf | * | [15 empty] | Pump turned off at 1743 - removed aerosol inlet head again to determine if there is an obstruction in the pipe.
2024-03-28 | 88 | 1705 | DR | clr | * | 0.599 | 0.77256 | 1 | 299/434 | 17.6/-9.9 |  |  | N |  | 5.9 | 100 | 18 |  |  |  | Pump turned back on - frost removed from inlet line - aerosol head reinstalled. Adjusted clap flow from 0.625 to 0.599. MAGIC concentrations high due to pump being off.
2024-03-29 | 89 | 1301 | DR | rm lf | * | 0.6 | 0.74259 | 1 | 253/397 | 17.4/15.7 |  |  | N |  | 5.9 | 100 | 18 |  |  |  | Adjusted clap flow from 0.615 to 0.600
2024-03-30 | 90 | 1309 | DR | rm lf | * | 0.6 | 0.73228 | 1 | 189/455 | 16.9/15.0 |  |  | N |  | 5.9 | 100 | 18 |  |  |  | Adjusted clap flow from 0.607 to 0.600
2024-03-31 | 91 | 1542 | DR | rm lf | * | 0.601 | 0.71426 | 1 | 127/537 | 17.4/15.7 |  |  | Y |  | 5.9 | 100 | 18 |  |  |  | Adjusted clap flow from 0.607 to 0.601
2024-04-01 | 92 | 1504 | DR | rm lf | * | 0.602 | 0.99606 | 2 | 187/743 | 17.5/15.7 |  |  | N |  | 5.9 | 100 | 18
2024-04-02 | 93 | 1629 | DR | clr | * | 0.598 | 0.96919 | 2 | 163/222 | 18.1/16.3 |  |  | N |  | 5.9 | Zero | 18
2024-04-03 | 94 |  | DR |  | * | [15 empty] | Did not make it out to AWO after ICESat
2024-04-04 | 95 | 1156 | DR | rm lf | * | 0.6 | 0.93162 | 2 | 129/190 | 17.6/15.2 |  |  | Y |  | 5.9 | 100 | 18 |  |  |  | Adjusted clap flow from 0.584 to 0.600. Uninstalled MAGIC250 at 1246 per PI request.
2024-04-05 | 96 | 1208 | DR | rm hf | * | 0.6 | 0.93012 | 2 | 105 | 17 |  |  | N |  | 5.9 | 100 | 18
2024-04-06 | 97 | 1242 | DR | rm lf | * | 0.6 | 0.91502 | 2 | 116 | 17.1 |  |  | N | 1.5% | 5.9 | 100 | 18 |  |  |  | Adjusted clap flow from 0.593 to 0.600. Neph Span Check 1204-1239 Error 1.5% CO2 Tank ~690psi
2024-04-07 | 98 | 1527 | DR | rm lf | * | 0.6 | 0.90981 | 2 | 77 | 17.5 |  |  | N |  | 5.9 | 100 | 18 |  |  |  | Adjusted clap flow from 0.592 to 0.600.
2024-04-08 | 99 | 1141 | DR | rm lf | * | 0.601 | 0.90817 | 2 | 88 | 16.6 |  |  | N |  | 5.9 | 100 | 18
2024-04-09 | 100 | 1717 | DR | rm lf | * | 0.6 | 0.9004 | 2 | 145 | 17.5 |  |  | Y |  | 5.9 | 100 | 18
2024-04-10 | 101 | 1423 | DR | rm lf | * | 0.598 | 0.89479 | 2 | 121 | 16.5 |  |  | N |  | 5.8 | 100 | 18
2024-04-11 | 102 | 1201 | DR | rm hf | * | 0.597 | 0.88671 | 2 | 141 | 16.1 |  |  | N |  | 5.9 | 100 | 18
2024-04-12 | 103 | 1127 | DR | rm lf | * | 0.597 | 0.86177 | 2 | 196 | 16.1 |  |  | N |  | 5.8 | 100 | 18
2024-04-13 | 104 | 1106 | DR | rm lf | * | 0.597 | 0.77251 | 2 | 890 | 16.7 |  |  | N |  | 5.9 | 100 | 18 |  |  |  | North winds.
2024-04-14 | 105 | 1855 | DR | rm lf | * | 0.603 | 0.7519 | 2 | 143 | 17.6 |  |  | N |  | 5.9 | 100 | 18
2024-04-15 | 106 | 1618 | DR | rm lf | * | 0.603 | 0.74159 | 2 | 215 | 17.5 |  |  | N |  | 5.9 | 100 | 18
2024-04-16 | 107 | 1205 | DR | rm lf | * | 0.599 | 0.73098 | 2 | 155 | 16.1 |  |  | N |  | 5.8 | 100 | 18
2024-04-17 | 108 | 2039 | DR | rm lf | * | 0.6 | 0.71775 | 2 | 333 | 17.7 |  |  | Y |  | 5.8 | 100 | 18
2024-04-18 | 109 | 1513 | DR | rm lf | * | 0.602 | 0.71362 | 2 | 419 | 16.7 |  |  | N |  | 5.8 | 100 | 18
2024-04-19 | 110 | 1742 | DR | rm lf | * | 0.601 | 0.70581 | 2 | 231 | 17.5 |  |  | N |  | 5.9 | 100 | 18 |  |  |  | Adjusted clap flow from 0.591 to 0.601.
2024-04-20 | 111 | 1621 | DR | rm lf | * | 0.601 | 0.99589 | 3 | 221 | 17.9 |  |  | N |  | 5.8 | 100 | 18 |  |  |  | Adjusted clap flow from 0.613 to 0.601.
2024-04-21 | 112 | 1342 | DR | rm hf | * | 0.599 | 0.97846 | 3 | 188 | 18.3 |  |  | N |  | 5.8 | 100 | 18 |  |  |  | Adjusted clap flow from 0.611 to 0.599.
2024-04-22 | 113 | 1214 | DR | rm lf | * | 0.599 | 0.96548 | 3 | 123 | 18.3 |  |  | N |  | 5.8 | Zero | 18 |  |  |  | Adjusted clap flow from 0.613 to 0.599.
2024-04-23 | 114 | 1020 | DR | clr | * | 0.6 | 0.95042 | 3 | 175 | 18 |  |  | Y |  | 5.9 | Zero | 18
2024-04-24 | 115 | 1728 | DR | rm lf | * | 0.596 | 0.93187 | 3 | 133 | 19 |  |  | N |  | 5.8 | Zero | 18
2024-04-25 | 116 |  | DR |  | * | [15 empty] | HM - did not make it out to AWO.
2024-04-26 | 117 | 1045 | DR | rm lf | * | 0.596 | 0.9137 | 3 | 110 | 18.4 |  |  | N |  | 5.8 | Zero | 18
2024-04-27 | 118 | 1130 | DR | rm lf | * | 0.6 | 0.89386 | 3 | 293 | 20.2 |  |  | N |  | 5.8 | Zero | 18 |  |  |  | Adjusted clap flow from 0.590 to 0.599.
2024-04-28 | 119 | 1454 | DR | clr | * | 0.599 | 0.8746 | 3 | 271 | 19.7 |  |  | N |  | 5.8 | Zero | 18
2024-04-29 | 120 | 1714 | DR | rm lf | * | 0.598 | 0.86423 | 3 | 271 | 20 |  |  | N |  | 5.8 | Zero | 18
2024-04-30 | 121 | 1230 | DR | clr | * | 0.596 | 0.84748 | 3 | 362 | 18.3 |  |  | Y |  | 5.9 | Zero | 18
2024-05-01 | 122 | 1532 | DR | clr | * | 0.599 | 0.81697 | 3 | 204 | 19.9 |  |  | N |  | 5.8 | 100 | 18 |  |  |  | Adjusted clap flow from 0.589 to 0.599. N wineds earlier.
2024-05-02 | 123 | 1203 | DR | clr | * | 0.598 | 0.80582 | 3 | 177 | 19.1 |  |  | N |  | 5.9 | 100 | 18
2024-05-03 | 124 | 1644 | DR | rm lf | * | 0.598 | 0.79857 | 3 | 124 | 20.2 |  |  | N |  | 5.8 | 100 | 18
2024-05-04 | 125 | 1637 | DR | rm lf | * | 0.598 | 0.79232 | 3 | 126 | 18.4 |  |  | N |  | 5.8 | 100 | 18
2024-05-05 | 126 |  | DR |  | * | [7 empty] | N | [7 empty] | HW - did not make it out to AWO.
2024-05-06 | 127 | 1743 | DR | clr | * | 0.6 | 0.78153 | 3 | 4281 | 17.5 |  |  | N | 1.7% | 5.8 | Zero | 18 |  |  |  | Neph Span Check 1227-1300 Error: 1.7% CO2 Tank ~700psi. Lowish winds - MAGIC reading 4k concentrations? Adjusted clap flow from 0.581 to 0.600.
2024-05-07 | 128 | 1533 | DR | rm lf | * | 0.602 | 0.77825 | 3 | 531 | 18.3 |  |  | Y |  | 5.8 | 100 | 18
2024-05-08 | 129 | 1028 | DR | rm lf | * | 0.6 | 0.77591 | 3 | 1357 | 16.9 |  |  | N |  | 5.8 | 100 | 18
2024-05-09 | 130 | 2003 | DR | rm lf | * | 0.6 | 0.76759 | 3 | 131 | 17.4 |  |  | N |  | 5.8 | 100 | 18
2024-05-10 | 131 |  | DR |  | * | [15 empty] | Did not make it out to AWO after ICESat
2024-05-11 | 132 | 1631 | DR | clr | * | 0.603 | 0.74881 | 3 |  |  |  |  | N |  | 5.8 | Zero | 18
2024-05-12 | 133 | 1345 | DR | rm lf | * | 0.604 | 0.74208 | 3 | 115 | 18.6 |  |  | N |  | 5.8 | Zero | 18
2024-05-13 | 134 | 1452 | DR | clr | * | 0.604 | 0.73627 | 3 | 258 | 18.2 |  |  | N |  | 5.8 | Zero | 18
2024-05-14 | 135 | 1534 | DR | clr | * | 0.597 | 0.72503 | 3 | 227 | 18 |  |  | Y |  | 5.8 | 100 | 18
2024-05-15 | 136 |  | DR |  | * | [15 empty] | Did not make it out to AWO after ICESat
2024-05-16 | 137 | 1448 | DR | clr | * | 0.6 | 0.71095 | 3 | 510 | 18.2 |  |  | N |  | 5.8 | Zero | 18 |  |  |  | Adjusted clap flow from 0.609 to 0.600.
2024-05-17 | 138 | 1208 | DR | clr | * | 0.601 | 0.94847 | 4 | 1913 | 17.1 |  |  | N |  | 5.7 | 100 | 18 |  |  |  | North winds
2024-05-18 | 139 | 1448 | DR | rm hf | * | 0.599 | 0.76719 | 4 | 1764 | 18 |  |  | N |  | 5.7 | Zero | 18 |  |  |  | North winds
2024-05-19 | 140 | 1846 | DR | rm lf | * | 0.598 | 0.72431 | 4 | 179 | 17.7 |  |  | N |  | 5.8 | 100 | 18
2024-05-20 | 141 | 1221 | DR | rm lf | * | 0.596 | 0.71564 | 4 | 232 | 17.5 |  |  | N |  | 5.7 | 100 | 18
2024-05-21 | 142 |  | DR |  | * | [15 empty] | HM - did not make it out to AWO.
2024-05-22 | 143 | 1119 | DR | rm lf | * | 0.598 | 0.99569 | 5 | 250 | 17.4 |  |  | Y |  | 5.7 | 100 | 18
2024-05-23 | 144 | 1151 | DR | rm lf | * | 0.6 | 0.98788 | 5 | 213 | 17.8 |  |  | N |  | 5.7 | 100 | 18 |  |  |  | AWO Raise Prep started today - track packing started at 0920L/1020Z - 1100L/1200Z.
2024-05-24 | 145 | 1000 | DR | rm lf | * | 0.595 | 0.97963 | 5 | 224 | 17.1 |  |  | N |  | 5.7 | 100 | 18
2024-05-25 | 146 | 1557 | DR | clr | * | 0.6 | 0.9616 | 5 | 2882 | 18 |  |  | N |  | 5.7 | Zero | 18 |  |  |  | Adjusted clap flow from 0.583 to 0.599. Adjusted neph flow a bit higher.
2024-05-26 | 147 | 731 | DR | clr | * | 0.599 | 0.96038 | 5 | 137 | 17.8 |  |  | N |  | 5.7 | Zero | 18 |  |  |  | Power outage - storm conditions.
2024-05-27 | 148 |  | DR | clr | * |  |  | 5
2024-05-28 | 149 |  | DR |  | * |  |  | 5 | [12 empty] | Day off.
2024-05-29 | 150 |  | DR |  | * |  |  | 5 | [12 empty] | Con 1 - didn't go out to AWO.
2024-05-30 | 151 | 1445 | DR | clr | * | 0.601 | 0.82999 | 5 | 201 | 18.1 |  |  | N |  | 5.7 | Zero | 18 |  |  |  | Adjusted clap flow from 0.581 to 0.601. North winds.
2024-05-31 | 152 |  | DR |  | * |  |  | 5
2024-06-01 | 153 | 955 | DR | rm lf | * | 0.601 | 0.77516 | 5 | 285 | 17 |  |  | Y |  | 5.7 | 100 | 18
2024-06-02 | 154 | 1103 | DR | rm lf | * | 0.605 | 0.73522 | 5 | 1532 | 17.4 |  |  | N |  | 5.7 | 100 | 18
2024-06-03 | 155 | 1629 | DR | clr | * | 0.6 | 0.71363 | 5 | 330 | 17.9 |  |  | N |  | 5.7 | 100 | 18 |  |  |  | Adjusted clap flow from 0.617 to 0.600. North winds.
2024-06-04 | 156 |  | DR |  | * | [15 empty] | Did not make it out to AWO after ICESat
2024-06-05 | 157 | 1056 | DR | rm lf | * | 0.601 | 0.84827 | 6 | 536 | 17.4 |  |  | N |  | 5.7 | Zero | 18
2024-06-06 | 158 | 1500 | CW | clr | * | 0.6 | 0.81164 | 6 | 849 | 18.7 |  |  | Y |  | 5.7 | Zero | 18
2024-06-07 | 159 | 1556 | CW | clr | * | 0.596 | 0.78063 | 6 | 834 | 17.6 |  |  | N | 1.2 | 5.6 | Zero | 18 |  |  |  | NephSpan check 1637-1714 Error 1.2% Tank ~700psi.
2024-06-08 | 160 | 950 | CW | clr | * | 0.596 | 0.7666 | 6 | 380 | 16.8 |  |  | N |  | 5.6 | Zero | 18
2024-06-09 | 161 | 1117 | CW | clr | * | 0.596 | 0.75306 | 6 | 110 | 17.1 |  |  | N |  | 5.6 | Zero | 18
2024-06-10 | 162 | 1003 | CW | clr | * | 0.593 | 0.74911 | 6 | 330 | 16.9 |  |  | N |  | 5.8 | Zero | 18
2024-06-11 | 163 | 1042 | CW | clr | * | 0.59 | 0.74768 | 6 | 88 | 16.9 |  |  | N |  | 5.9 | Zero | 18 |  |  |  | Adjusted CLAP flow from 0.59lmp to 0.601lpm.
2024-06-12 | 164 | 1008 | CW | clr | * | 0.596 | 0.7468 | 6 | 61 | 17.1 |  |  | Y |  | 5.9 | Zero | 18
2024-06-13 | 165 | 1028 | CW | clr | * | 0.601 | 0.74516 | 6 | 74 | 17.9 |  |  | N |  | 5.9 | Zero | 18
2024-06-14 | 166 | 940 | CW | clr | * | 0.603 | 0.74342 | 6 | 1550 | 17.6 |  |  | N |  | 5.9 | Zero | 18
2024-06-15 | 167 | 1212 | CW | clr | * | 0.604 | 0.73174 | 6 | 226 | 18.1 |  |  | N |  | 5.9 | Zero | 18
2024-06-16 | 168 | 1459 | CW | clr | * | 0.601 | 0.7028 | 6 | 335 | 18.9 |  |  | N |  | 6.2 | Zero | 18 |  |  |  | Adjusted CLAP flor from 0.606 to 0.601.
2024-06-17 | 169 | 1146 | CW | clr | * | 0.601 | 0.92556 | 7 | 182 | 19.3 |  |  | N |  | 6.2 | Zero | 18 |  |  |  | Adjusted CLAP flow from 0.616 to 0.601
2024-06-18 | 170 | 1004 | CW | rm lf | * | 0.596 | 0.83127 | 7 | 510 | 18.3 |  |  | N |  | 6.3 | Zero | 18
2024-06-19 | 171 | 1630 | CW | rm lf | * | 0.591 | 0.70977 | 7 | 99999.9 | 18.6 |  |  | Y |  | 6.3 | Zero | 18 |  |  |  | Adjusted CLAP flow from 0.591 to 0.600.
2024-06-20 | 172 | 945 | CW | n/a | * | 0.6 | 0.70446 | 7 | 113 | 18.2 |  |  | N |  | 6.2 | 0100 | 18 |  |  |  | CLAP flag: 004. NEPH flag 0100. Did not go onto AWO roof due to lift operations.
2024-06-21 | 173 | 1224 | CW | n/a | * | 0.6 | 0.7016 | 7 | 143 | 20.2 |  |  | N |  | 6.2 | Zero | 18 |  |  |  | Did not go onto AWO roof due to lift operations.
2024-06-22 | 174 |  |  |  | * | [7 empty] | N | [7 empty] | Did not go out to AWO, awaiting completion of stairs.
2024-06-23 | 175 | 1337 | CW | clr | * | 0.601 | 0.95298 | 8 | 214 | 18.8 |  |  | N |  | 6.2 | Zero | 18 |  |  |  | Adjusted clap flow from 0.583 to 0.601.
2024-06-24 | 176 | 1450 | CW | clr | * | 0.6 | 0.91165 | 8 | 669 | 18.3 |  |  | N |  | 6.2 | Zero | 18 |  |  |  | Changed CLAP filter. Adjusted CLAP flow from 0.613 to 0.601.
2024-06-25 | 177 | 1414 | CW | rm lf |  | 0.602 | 0.99627 | 1 | 330 | 18.9 |  |  | N |  | 6.2 | Zero | 18 |  |  |  | Did not receive daily aerosol email.
2024-06-26 | 178 | 1125 | CW | clr |  | 0.6 | 0.95874 | 1 | 241 | 18.4 |  |  | N |  | 6.2 | Zero | 18 |  |  |  | Adjusted CLAP flow from 0.610 to 0.600. Have not received daily aerosol email.
2024-06-27 | 179 | 1055 | CW | clr |  | 0.6 | 0.86086 | 1 | 2214 | 18.3 |  |  | Y |  | 6.1 | Zero | 18 |  |  |  | Adjusted CLAP flow from 0.605 to 0.600.
2024-06-28 | 180 | 1049 | CW | clr | * | 0.598 | 0.84277 | 1 | 83 | 18.8 |  |  | N |  | 6.1 | Zero | 18
2024-06-29 | 181 | 946 | CW | clr | * | 0.6 | 0.83889 | 1 | 254 | 18 |  |  | N |  | 6.1 | 1 | 18 |  |  |  | Adjusted CLAP flor from 0.591 to 0.600. Power outage during UPS maintenance, UPS instruments affected from 1128-1134Z.
2024-06-30 | 182 | 1318 | CW | clr | * | 0.597 | 0.83162 | 1 | 137 | 18.6 |  |  | N |  | 6.1 | Zero | 18
2024-07-01 | 183 | 1538 | CW |  |  | 0.6 | 0.82752 | 1 | 119 | 18.1 |  |  | N |  | 6.5 | Zero | 18 |  |  |  | Adjusted CLAP flow from 0.596 to 0.600.
2024-07-02 | 184 | 1511 | CW | clr | * | 0.604 | 0.82685 | 1 | 181 | 19.3 |  |  | N |  | 6.5 | Zero | 18 |  |  |  | Changed neph bulb, powered down from 1527 to 1600; lamp: 12.7V, 6.0A, 75W.
2024-07-03 | 185 | 1007 | CW | clr | * | 0.6 | 0.82171 | 1 | 271 | 17.9 |  |  | Y | 2.1 | 5.9 | Zero | 18 |  |  |  | Neph Span check: 0933 - 1003, 2.1% error, Tank pressure ~700psi. Adjusted CLAP flow from 0.608 to 0.600.
2024-07-04 | 186 |  |  |  | * | [15 empty] | Did not make it to AWO after ICESAT.
2024-07-05 | 187 | 1422 | CW | clr | * | 0.606 | 0.79206 | 1 | 510 | 20.2 |  |  | N |  | 5.9 | Zero | 18
2024-07-06 | 188 | 1756 | CW | clr | * | 0.599 | 0.77398 | 1 | 276 | 19.9 |  |  | N |  | 5.9 | Zero | 18
2024-07-07 | 189 | 1159 | CW | rm lf | * | 0.597 | 0.76319 | 1 | 215 | 18.6 |  |  | N |  | 5.9 | Zero | 18
2024-07-08 | 190 | 946 | CW | rm lf | * | 0.596 | 0.75453 | 1 | 138 | 17.6 |  |  | N |  | 5.9 | Zero | 18
2024-07-09 | 191 | 951 | CW | rm lf | * | 0.6 | 0.74192 | 1 | 237 | 18.1 |  |  | Y |  | 5.9 | Zero | 18 |  |  |  | Adjusted CLAP flow from 0.589 to 0.600.
2024-07-10 | 192 | 1406 | CW | rm lf | * | 0.595 | 0.73356 | 1 | 204 | 18.4 |  |  | N |  | 5.9 | Zero | 18
2024-07-11 | 193 | 1212 | CW | clr | * | 0.596 | 0.72521 | 1 | 107 | 17.7 |  |  | N |  | 5.9 | Zero | 18
2024-07-12 | 194 | 1559 | CW | clr | * | 0.599 | 0.71852 | 1 | 89 | 17.9 |  |  | N |  | 6 | Zero | 18
2024-07-13 | 195 |  | CW | clr
2024-07-14 | 196 | 1707 | CW | clr | * | 0.597 | 0.71102 | 1 | 113 | 18.5 |  |  | N |  | 6 | Zero | 18
2024-07-15 | 197 | 1124 | CW | clr | *
2024-07-16 | 198 | 1150 | CW | clr | * | 0.6 | 0.70641 | 1 | 56 | 18.6 |  |  | Y |  | 5.9 | Zero | 18 |  |  |  | Adjusted CLAP flow from 0.610 to 0.601
2024-07-17 | 199 | 1429 | CW | clr | * | 0.603 | 0.99145 | 2 | 254 | 19.8 |  |  | N |  | 5.9 | Zero | 18
2024-07-18 | 200 | 1726 | CW | rm lf | *
2024-07-19 | 201 | 1309 | CW | clr | * | 0.605 | 0.96168 | 2 | 180 | 19.1 |  |  | N |  | 6 | Zero | 18
2024-07-20 | 202 | 1349 | CW | clr | * | 0.588 | 0.91319 | 2 | 1829 | 19.1 |  |  | N |  | 5.9 | Zero | 18
2024-07-21 | 203 | 1005 | CW | clr | * | [15 empty] | Erin adjusted CLAP flow to 0.685 slpm.
2024-07-22 | 204 | 1305 | CW | rm lf | * | 0.656 | 0.85669 | 2 | 339 | 17.5 |  |  | N |  | 6 | Zero | 18
2024-07-23 | 205 | 1547 | CW | rm lf | *
2024-07-24 | 206 | 1511 | CW | rm lf | * | 0.653 | 0.98373 | 1 | 120 | 17.4 |  |  | N |  | 6 | Zero | 18 |  |  |  | Erin changed filter on CLAP. New neph installed, not yet connected to intake. New software acquisition system (LoggerNet) installed for instruments connected to Campbell CR1000 data logger.
2024-07-25 | 207 | 1217 | CW | clr | * | 0.654 | 0.97128 | 1 | [6 empty] | 6 | Zero | 18
2024-07-26 | 208 | 1551 | CW | clr | * | [15 empty] | Erin installing new neph, new tubing from the pentapus to old/new neph.
2024-07-27 | 209 | 1032 | CW | clr | * | 0.662 | 0.93765 | 1 | 139 | 17.8 |  |  | N |  | 5.9 | Zero | 18
2024-07-28 | 210 | 1204 | CW | clr | *
2024-07-29 | 211 | 1218 | CW | clr | * | 0.754 | 0.92806 | 1 | 1636 | 16.8 |  |  | Y |  | n/a | n/a | n/a | Zero | 4.5 | Pass | Neph turned off 1030Z for maintenance by Erin. Repowered 1245Z.
2024-07-30 | 212 | 1400 | CW | clr | * | [15 empty] | Magic 250 reinitialized, now collecting good data.
2024-07-31 | 213 | 1657 | CW | clr | * | 0.678 | 0.90229 | 1 | 424 | 18 | 424 | -6.9 | N | Eco : 1.6 | 5.9 | Zero | 18 | Zero | 4.5
2024-08-01 | 214 | 1459 | CW | rm lf | *
2024-08-02 | 215 | 1656 | CW | rm lf | * | 0.662 | 0.86703 | 1 | 143 | 18 | 142 | -7.5 | N |  | 5.9 | Zero | 18 | Zero | 4.5
2024-08-03 | 216 | 1557 | CW | rm lf
2024-08-04 | 217 | 1734 | CW | rm lf |  | 0.657 | 0.72116 | 1 | 1463 | 20.4 | 1028 | -6.1 | N |  | 5.9 | Zero | 18 | Zero | 4.5
2024-08-05 | 218
2024-08-06 | 219 | 1039 | CW | rm lf | * | 0.647 | 0.99318 | 2 | 138 | 17.3 | 138 | -8.8 | Y |  | 5.9 | Zero | 18 | Zero | 4.5
2024-08-07 | 220 | 1130 | CW | rm lf | *
2024-08-08 | 221 | 1014 | CW | clr | * | 0.639 | 0.9713 | 2 | 120 | 18.2 | 124 | -9.1 | N |  | 5.9 | Zero | 18 | Zero | 4.5
2024-08-09 | 222 | 1302 | CW | clr | * | [15 empty] | Display on Magic 200 not illuminated, reading concentrations from the aerosol acquisition software instead.
2024-08-10 | 223 | 1115 | CW | rm lf | * | 0.633 | 0.91115 | 2 | 85 | 17.9 | 85 | -9.8 | N |  | 5.9 | Zero | 18 | Zero | 4.5
2024-08-11 | 224 | 1843 | CW | clr | *
2024-08-12 | 225 | 1140 | CW | rm lf | * | 0.641 | 0.8992 | 2 | 96 | 18.2 | 92 | -8.5 | N |  | 5.9 | Zero | 18 | Zero | 4.5 | Pass
2024-08-13 | 226 | 1000 | CW | clr | * | [15 empty] | Blip in data @ 0658-0727Z. Changed CLAP filter - unsure why the instrument needed filter change but could be related to the brief power outage this morning? Depowered Magic 210 for retro, hid columns.
2024-08-14 | 227 |  |  |  | * | [15 empty] | ICESat
2024-08-15 | 228 | 1501 | CW | clr | * | 0.609 | 0.97438 | 1 |  |  | 326 | -9.1 | Y |  | 5.9 | Zero | 18 | Zero | 4.5
2024-08-16 | 229 | 1004 | CW | rm lf | *
2024-08-17 | 230 | 1537 | CW | clr | * | 0.617 | 0.93825 | 1 |  |  | 98 | -9 | N |  | 5.9 | Zero | 18 | Zero | 4.5 |  | Changed CLAP filter for Nathan - approved by Erin.
2024-08-18 | 231 | 1845 | CW |  | *
2024-08-19 | 232 | 1541 | CW | clr | * | 0.889 | 1.2539 | 1 |  |  | 644 | -10.3 | N |  | 5.9 | Zero | 18 | Zero | 4.5
2024-08-20 | 233 | 1356 | CW | rm lf | * | 0.599 | [14 empty] | Adjusted CLAP flow to 0.599
2024-08-21 | 234 | 1226 | CW | clr | * | 0.596 | 1.3118 | 1 |  |  | 570 | -8.9 | N |  | 5.9 | Zero | 18 | Zero | 4.5
2024-08-22 | 235 | 1909 | CW | clr | * | 0.6 | [14 empty] | Adjusted CLAP flow from 0.474 to 0.601
2024-08-23 | 236 | 1029 | CW | rm lf | * | 0.597 | 1.2882 | 1 |  |  | 190 | -9.6 | N |  | 5.9 | Zero | 18 | Zero | 4.5 |  | AWO internet outage from 2024-08-22 at 2224Z- 2024-08-23 at 0907Z.
2024-08-24 | 237 | 1030 | CW | clr | *
2024-08-25 | 238 | 1428 | CW | rm lf | * | 0.6 | 1.1511 | 1 |  |  | 119 | -11.1 | N |  | 5.9 | Zero | 18 | Zero | 4.5 | Pass
2024-08-26 | 239 | 1215 | CW | clr | *
2024-08-27 | 240 | 1231 | CW | clr | * | 0.592 | 1.094 | 1 |  |  | 184 | -10.6 | Y |  | 5.9 | Zero | 18 | Zero | 4.5
2024-08-28 | 241 | 1355 | CW | rm lf | *
2024-08-29 | 242 | 1325 | CW | clr | * | 0.6 | 1.015 | 1 |  |  | 134 | -9.5 | N |  | 5.9 | Zero | 18 | Zero | 4.5 |  | Adjusted CLAP flow to 0.60.
2024-08-30 | 243 | 917 | CW | clr | *
2024-08-31 | 244 | 1218 | CW | clr | * | 0.603 | 0.9505 | 1 |  |  | 72 | -10.3 | N |  | 5.9 | Zero | 18 | Zero | 4.5
2024-09-01 | 245 | 1252 | CW | clr | *
2024-09-02 | 246 | 1505 | CW | clr | * | 0.604 | 0.90996 | 1 |  |  | 98 | -10.3 | N |  | 5.9 | Zero | 18 | Zero | 4.5
2024-09-03 | 247 |  |  |  | * | [15 empty] | Did not make it to AWO after ICESat.
2024-09-04 | 248 | 1357 | CW | clr | * | 0.6 | 0.86926 | 1 |  |  | 163 | -9.6 | N | TSI: 1.3 | 5.9 | Zero | 18 | Zero | 4.5 |  | Magic concentration dropped to Zero, and CLAP flow dropped to zero during tsi spancheck.
2024-09-05 | 249 | 1416 | CW | rm lf | *
2024-09-06 | 250 | 1428 | CW | clr | * | 0.6 | 0.86316 | 1 |  |  | 111 | -10.8 | Y |  | 5.9 | Zero | 18 | Zero | 4.5
2024-09-07 | 251 |  |  |  | * | [15 empty] | Did not go to AWO due to poor visibility
2024-09-08 | 252 | 1834 | CW | clr | * | 0.613 | 0.82071 | 1 |  |  | 89 | -12.2 | N |  | 5.9 | Zero | 18 | Zero | 4.5
2024-09-09 | 253 | 1235 | CW | clr | * | [14 empty] | Pass
2024-09-10 | 254 |  |  |  | * | [15 empty] | Did not go to AWO due to poor visibility
2024-09-11 | 255 | 1311 | CW | rm hf | * | 0.615 | 0.74902 | 1 |  |  | 34 | -11.6 | N |  | 5.9 | Zero | 18 | Zero | 4.5
2024-09-12 | 256 | 912 | CW | rm hf | *
2024-09-13 | 257 | 1004 | CW | rm hf | * | 0.61 | 0.73659 | 1 |  |  | 103 | -10.7 | N |  | 5.9 | Zero | 18 | Zero | 4.5
2024-09-14 | 258 | 1102 | CW | rm lf | *
2024-09-15 | 259 | 1839 | CW | clr | * | 0.6 | 0.84958 | 2 |  |  | 70 | -12.2 | N |  | 5.9 | Zero | 18 | Zero | 4.5
2024-09-16 | 260 | 1201 | CW | rm lf | *
2024-09-17 | 261 | 1846 | CW | rm lf | * | 0.59 | 0.83051 | 2 |  |  | 84 | -12.2 | Y |  | 5.9 | Zero | 18 | Zero | 4.5 |  | TSI neph relative humidity out of range, reading 668 (actual 671.4). Ecotech reading 671.1.
2024-09-18 | 262 | 1439 | CW | rm hf | *
2024-09-19 | 263 | 1236 | CW | clr | * | 0.6 | 0.81799 | 2 |  |  | 48 | -12.2 | N |  | 5.9 | Zero | 18 | Zero | 4.5 |  | Relative humidity still out of range on TSI
2024-09-20 | 264 | 1008 | CW | clr | *
2024-09-21 | 265 | 944 | CW | clr | * | 0.61 | 0.80522 | 2 |  |  | 150 | -12.2 | N |  | 5.9 | Zero | 18 | Zero | 4.5
2024-09-22 | 266 | 1252 | CW | clr | *
2024-09-23 | 267 | 1255 | CW | clr | * | 0.607 | 0.77786 | 2 |  |  | 204 | -10.3 | N |  | 5.9 | Zero | 18 | Zero | 4.5 | Pass
2024-09-24 | 268 | 1050 | CW | lf | *
2024-09-25 | 269 | 1452 | CW | rm hf | * | 0.6 | 0.75845 | 2 |  |  | 37 | -9.9 | N |  | 5.9 | Zero | 18 | Zero | 4.5
2024-09-26 | 270 | 1442 | CW | rm lf | *
2024-09-27 | 271 | 1001 | CW | rm lf | * | 0.6 | 0.74103 | 2 |  |  | 116 | -11.3 | N |  | 5.9 | Zero | 18 | Zero | 4.5 |  | Adjusted CLAP flow from 0.501 to 0.600.
2024-09-28 | 272 |  |  |  | *
2024-09-29 | 273 | 1303 | CW | rm lf | * | 0.6 | 0.7278 | 2 |  |  | 9 | -12.2 | Y |  | 5.9 | Zero | 18 | Zero | 4.5 |  | Adjusted CLAP flow from 0.504 to 0.600.
2024-09-30 | 274 | 1109 | CW | clr | *
2024-10-01 | 275 | 1158 | CW | rm lf | * | 0.6 | 0.70791 | 2 |  |  | 227 | -12 | N |  | 5.9 | Zero | 18 | Zero | 4.5 |  | Adjusted CLAP flow from 0.607 to 0.601.
2024-10-02 | 276 | 1505 | CW | rm lf | *
2024-10-03 | 277 | 1504 | CW | rm lf | * | 0.6 | 0.97253 | 3 |  |  | 210 | -10 | N |  | 5.9 | Zero | 18 | Zero | 4.5 |  | Adjusted CLAP flow from 0.640 to 0.600.
2024-10-04 | 278 | 1652 | CW | clr | *
2024-10-05 | 279 |  | CW |  | * | [15 empty] | Did not go to AWO after ICESat
2024-10-06 | 280 | 1720 | CW | rm lf | * | 0.588 | 0.90362 | 3 |  |  | 176 | -10.9 | N |  | 5.9 | Zero | 18 | Zero | 4.5
2024-10-07 | 281 | 1325 | CW | rm hf | *
2024-10-08 | 282 | 1010 | CW | rm lf | * | 0.6 | 0.87761 | 3 |  |  | 38 | -12.2 | Y |  | 5.9 | Zero | 18 | Zero | 4.5
2024-10-09 | 283 | 1209 | CW | rm lf | * | [14 empty] | Pass
2024-10-10 | 284 | 1127 | CW | rm lf | * | 0.6 | 0.86707 | 3 |  |  | 198 | -12.2 | N |  | 5.9 | Zero | 18 | Zero | 4.5
2024-10-11 | 285 | 1321 | CW | rm lf | *
2024-10-12 | 286 | 1235 | ALW | rm lf | * | 0.6 | 0.85431 | 3 |  |  | 77 | -12.2 | N |  | 5.8 | Zero | 18 | Zero | 4.5
2024-10-13 | 287 | 1343 | ALW | rm lf | * | 0.59 | 0.84463 | 3 |  |  | 59 | -12.2 | Y |  | 5.9 | zero | 18 | zero | 4.5
2024-10-14 | 288 | 1535 | ALW | rm lf | * | 0.59 | 0.83894 | 3 |  |  | 54 | -12.2 | n | E: 1.3 | 5.8 | zero | 18 | zero | 4.5
2024-10-15 | 289 | 1309 | ALW | rm lf | * | 0.6 | 0.88057 | 3 |  |  | 87 | -12.2 | n |  | 5.8 | zero | 18 | zero | 4.5
2024-10-16 | 290 | 1731 | ALW | rm lf | * | 0.6 | 0.7921 | 3 |  |  | 96 | -12.2 | n |  | 5.9 | zero | 18 | zero | 4.5
2024-10-17 | 291 | 1147 | ALW | rm lf | * | 0.596 | 0.78512 | 3 |  |  | 92 | -12.2 | n |  | 5.9 | zero | 18 | zero | 4.5
2024-10-18 | 292 | 1343 | ALW | rm lf | * | - | - | - |  |  | - | - | n |  | - | - | 18 | zero | 4.5 |  | Dell computer has "!LOST CONNECTION!", have contacted Erin.
2024-10-19 | 293 | 1047 | ALW | rm lf | * | - | - | - |  |  | - | - | n |  | - | - | 18 | zero | 4.5 |  | Dell computer has "!LOST CONNECTION!", have contacted Erin. Restarted Dell @ 1047
2024-10-20 | 294 | 1413 | ALW | rm lf | * | - | - | - |  |  | 38 | -12.2 | n |  | - | - | 18 | zero | 4.5 | pass | CLAP and TSI have no connection in Forge
2024-10-21 | 295 | 1511 | ALW | rm lf | * | 0.6 | 0.78317 | 3 |  |  | 117 | -12.2 | n |  | 5.8 | zero | 18 | zero | 4.5 | pass | CLAP and TSI reconnected to Forge @ 1636.
2024-10-22 | 296 | 1250 | ALW | rm lf | * | 0.608 | 0.77814 | 3 |  |  | 120 | -12.2 | n |  | 5.9 | zero | 18 | zero | 4.5
2024-10-23 | 297 | 1649 | ALW | rm lf | * | 0.613 | 0.77193 | 3 |  |  | 123 | -12.2 | n |  | 5.8 | zero | 18 | zero | 4.5
2024-10-24 | 298 | 1514 | ALW | rm lf | * | 0.609 | 0.75732 | 3 |  |  | 17 | -12.2 | n |  | 5.8 | zero | 18 | zero | 4.5
2024-10-25 | 299 | 1252 | ALW | rm lf | * | 0.608 | 0.75517 | 3 |  |  | 59 | -12.2 | n |  | 5.8 | zero | 18 | zero | 4.5
2024-10-26 | 300 | 1254 | ALW | rm lf | * | 0.614 | 0.75058 | 3 |  |  | 26 | -12.2 | n |  | 5.8 | zero | 18 | zero | 4.5
2024-10-27 | 301 | 1455 | ALW | rm lf | * | 0.611 | 0.73868 | 3 |  |  | 31 | -12.2 | y |  | 5.8 | zero | 18 | zero | 4.5
2024-10-28 | 302 | 1530 | ALW | rm lf | * | 0.603 | 0.73646 | 3 |  |  | 18 | -12.2 | n |  | 5.8 | zero | 18 | zero | 4.5
2024-10-29 | 303 | - | ALW | - | * | - | - | - |  |  | - | - | - |  | - | - | - | - | - |  | Con 1- didn't go to AWO
2024-10-30 | 304 | 1505 | ALW | rm lf | * | 0.606 | 0.72923 | 3 |  |  | 22 | -12.2 | n |  | 5.8 | zero | 18 | zero | 4.5
2024-10-31 | 305 | 1815 | ALW | rm lf | * | 0.612 | 0.7265 | 3 |  |  | 40 | -12.2 | n |  | 5.8 | zero | 18 | zero | 4.5
2024-11-01 | 306 | 1509 | ALW | rm lf | * | 0.604 | 0.71439 | 3 |  |  | 61 | -12.2 | n | echotech | 5.8 | zero | 18 | zero | 4.5 |  | Echotech spancheck 1442-1519Z with 1.5% error
2024-11-02 | 307 | - | ALW | - | * | - | - | - |  |  | - | - | n |  | - | - | - | - | - |  | ICESat
2024-11-03 | 308 | 1515 | ALW | rm lf | * | 0.595 | 0.70975 | 3 |  |  | 18 | -12.2 | N |  | 5.8 | Zero | 18 | Zero | 4.5
2024-11-04 | 309 | 1342 | ALW | - | * | 0.594 | 0.70771 | 3 |  |  | 11 | -12.2 | n |  | 5.8 | zero | 18 | zero | 4.5 |  | Too windy to go to roof
2024-11-05 | 310 | 1438 | ALW | rm lf | * | 0.6 | 0.70645 | 3 |  |  | 23 | -12.2 | n |  | 5.8 | zero | 18 | zero | 4.5
2024-11-06 | 311 | 1347 | ALW | - | * | 0.589 | 0.70504 | 3 |  |  | 10 | -12.2 | n |  | 5.8 | zero | 18 | zero | 4.5 |  | Too windy to go to roof
2024-11-07 | 312 | 1923 | ALW | rm lf | * | 0.588 | 0.70354 | 3 |  |  | 23 | -12.2 | n |  | 5.8 | zero | 18 | zero | 4.5
2024-11-08 | 313 | 1738 | ALW | rm lf | * | 0.594 | 0.70131 | 3 |  |  | 42 | -12.2 | y |  | 5.8 | zero | 18 | zero | 4.5
2024-11-09 | 314 | 1258 | ALW | rm lf | * | 0.603 | 0.99734 | 4 |  |  | 20 | -12 | n |  | 5.8 | zero | 18 | zero | 4.5
2024-11-10 | 315 | 1500 | ALW | rm lf | * | 0.6 | 0.99575 | 4 |  |  | 17 | -12.2 | n |  | 5.8 | zero | 18 | zero | 4.5 | pass
2024-11-11 | 316 | 1259 | ALW | rm lf | * | 0.594 | 0.99509 | 4 |  |  | 1 | -12.2 | n |  | 5.8 | zero | 18 | zero | 4.5
2024-11-12 | 317 | 1453 | ALW | rm lf | * | 0.591 | 0.99273 | 4 |  |  | 24 | -9.4 | n |  | 5.8 | zero | 18 | zero | 4.5
2024-11-13 | 318 | - | ALW | - | * | - | - | - |  |  | - | - | n |  | - | - | - | - | - |  | ICESat- did not go to AWO
2024-11-14 | 319 | 1553 | ALW | - | * | 0.597 | 0.98775 | 4 |  |  | 51 | -12.2 | n |  | 5.8 | Zero | 18 | Zero | 4.5 | pass | Too windy to go to roof
2024-11-15 | 320 | 1341 | ALW | rm lf | * | 0.608 | 0.98509 | 4 |  |  | 10 | -12.2 | n |  | 5.8 | zero | 18 | zero | 4.5
2024-11-16 | 321 | 1238 | ALW | rm lf | * | 0.607 | 0.96318 | 4 |  |  | 12 | -12.2 | y |  | 5.8 | zero | 18 | zero | 4.5
2024-11-17 | 322 | 1432 | ALW | rm lf | * | 0.615 | 0.94258 | 4 |  |  | 92 | -12.2 | n |  | 5.8 | zero | 18 | zero | 4.5
2024-11-18 | 323 | 1549 | ALW | rm lf | * | 0.622 | 0.93433 | 4 |  |  | 0 | -12.2 |  |  | 5.8 | zero | 18 | zero | 4.5 |  | Magic CPC claims to have a dry wick
2024-11-19 | 324 | 1324 | ALW | rm lf | * | 0.598 | 0.92339 | 4 |  |  | 76.8 | -12.2 | y |  | 5.8 | zero | 18 | zero | 4.5 |  | Magic CPC wick rewet. Turned off 1651-1719Z.
2024-11-20 | 325 | 1627 | ALW | rm lf | * | 0.594 | 0.91591 | 4 |  |  | 111 | -12.2 | n |  | 5.8 | zero | 18 | zero | 4.5 |  | Magic CPC wick rewet. Turned off 1710-1727Z.
2024-11-21 | 326 | 1612 | ALW | rm lf | * | 0.589 | 0.91044 | 4 |  |  | 73 | -12.2 | n |  | 5.8 | zero | 18 | zero | 4.5 | pass | Magic CPC wick rewet. Turned off 1720-1735Z.
2024-11-22 | 327 | 1320 | ALW | rm lf | * | 0.585 | 0.90522 | 4 |  |  | 16 | -12.2 | n |  | 5.8 | Zero | 18 | Zero | 4.5 |  | Magic CPC wick rewet. Turned off 1442-1506Z.
2024-11-23 | 328 | 1234 | ALW | rm lf | * | 0.59 | 0.90271 | 4 |  |  | 46 | -12.2 | n |  | 5.8 | Zero | 18 | Zero | 4.5 |  | Adjusted CLAP lpm from .584-.598. Magic CPC wick rewet. Turned off 1238-1253.
2024-11-24 | 329 | 1404 | ALW | rm lf | * | 0.595 | 0.89626 | 4 |  |  | 81 | -12.2 | n |  | 5.8 | zero | 18 | zero | 4.5 |  | Magic CPC wick rewet. Turned off 1432-1452Z. Won't recommunicate.
2024-11-25 | 330 | - | ALW | - | * | - | - | - |  |  | - | - | - | - | - | - | - | - | - | - | Con 1- did not go to AWO
2024-11-26 | 331 | 1811 | ALW | rm lf | * | 0.578 | 0.88698 | 4 |  |  | 82 | -12.2 | n |  | 5.8 | Zero | 18 | Zero | 4.5 |  | Magic CPC wick rewet. Humidifier re-prepped. Vapor lock loosened. Turned off 1816-1906. Bubbles purged from 1907-1910.
2024-11-27 | 332 | - | ALW | - | * | - | - | - |  |  | - | - | - | - | - | - | - | - | - | - | Con 1- did not go to AWO
2024-11-28 | 333 | - | ALW | - | * | - | - | - |  |  | - | - | - | - | - | - | - | - | - | - | Holiday- did not go to AWO
2024-11-29 | 334 | 1520 | ALW | rm lf | * | 0.594 | 0.85636 | 4 |  |  | 67 | -12.2 | n |  | 5.8 | zero | 18 | Zero | 4.5
2024-11-30 | 335 | 1410 | ALW | rm | * | 0.594 | 0.84743 | 4 |  |  | 38 | -12.2 | n |  | 5.7 | zero | 18 | zero | 4.5
2024-12-01 | 336 | - | ALW | - | * | - | - | - |  |  | - | - | - | - | - | - | - | - | - |  | Con 1- did not go to AWO
2024-12-02 | 337 | 1207 | ALW | rm lf | * | 0.587 | 0.83016 | 4 |  |  | 25 | -12.2 | n |  | 5.7 | zero | 18 | Zero | 4.5
2024-12-03 | 338 | 1346 | ALW | rm lf | * | 0.581 | 0.82645 | 4 |  |  | 14 | -12.2 | y |  | 5.8 | Zero | 18 | Zero | 4.5 | pass
2024-12-04 | 339 | - | ALW | - | * | - | - | - |  |  | - | - | - | - | - | - | - | - | - | - | ICESAT- did not go to AWO
2024-12-05 | 340 | - | ALW | - | * | - | - | - |  |  | - | - | - | - | - | - | - | - | - | - | Sick- did not go to AWO
2024-12-06 | 341 | 1702 | ALW | rm lf | * | 0.59 | 0.82298 | 4 |  |  | 29 | -12.2 | n | TSI: 1% | 5.7 | zero | 18 | Zero | 4.5 |  | TSI Neph Spancheck 1553-1623. 1% Error
2024-12-07 | 342 | 1238 | ALW | rm lf | * | 0.594 | 0.82229 | 4 |  |  | 21 | -12.2 | n |  | 5.7 | Zero | 18 | Zero | 4.5
2024-12-08 | 343 | 1536 | ALW | rm lf | * | 0.588 | 0.81921 | 4 |  |  | 32 | -11.6 | n |  | 5.7 | zero | 18 | zero | 4.5
2024-12-09 | 344 | 1401 | ALW | rm lf | * | 0.588 | 0.81731 | 4 |  |  | 38 | -12.2 | y |  | 5.7 | Zero | 18 | Zero | 4.5
2024-12-10 | 345 | 1623 | ALW | rm lf | * | 0.604 | 0.81388 | 4 |  |  | 37 | -12.2 | n |  | 5.7 | Zero | 18 | Zero | 4.5 | pass
2024-12-11 | 346 | 1553 | ALW | - | * | 0.595 | 0.80941 | 4 |  |  | 47 | -12.2 | n |  | 5.7 | zero | 18 | zero | 4.5
2024-12-12 | 347 | 1557 | ALW | rm lf | * | 0.593 | 0.80663 | 4 |  |  | 52 | -12.2 | n |  | 5.7 | Zero | 18 | Zero | 4.5
2024-12-13 | 348 | 1258 | ALW | - | * | 0.601 | 0.80287 | 4 |  |  | 39 | -12.2 | n |  | 5.7 | Zero | 18 | Zero | 4.5 |  | Too windy to go to roof
2024-12-14 | 349 | 1309 | ALW | rm lf | * | 0.588 | 0.80047 | 4 |  |  | 18 | -12.2 | n |  | 5.7 | Zero | 18 | Zero | 4.5
2024-12-15 | 350 | 1404 | ALW | rm lf | * | 0.59 | 0.79944 | 4 |  |  | 10 | -12.2 | n |  | 5.7 | Zero | 18 | Zero | 4.5
2024-12-16 | 351 | 1530 | ALW | rm lf | * | 0.597 | 0.7988 | 4 |  |  | 13 | -12.2 | n |  | 5.7 | Zero | 18 | Zero | 4.5
2024-12-17 | 352 | 1402 | ALW | rm lf | * | 0.597 | 0.79736 | 4 |  |  | 12 | -12.2 | y |  | 5.7 | Zero | 18 | Zero | 4.5 | pass
2024-12-18 | 353 | - | ALW | - | * | - | - | - |  |  | - | - | - | - | - | - | - | - | - | - | Con 1- did not go to AWO
2024-12-19 | 354 | 1614 | ALW | rm lf | * | 0.599 | 0.7951 | 4 |  |  | 13 | -12.2 | n |  | 5.7 | Zero | 18 | Zero | 4.5
2024-12-20 | 355 | 1531 | ALW | rm lf | * | 0.602 | 0.78732 | 4 |  |  | 35 | -12.2 | n |  | 5.7 | Zero | 18 | Zero | 4.5
2024-12-21 | 356 | - | ALW | - | * | - | - | - |  |  | - | - | - | - | - | - | - | - | - | - | Con 1- did not go to AWO
2024-12-22 | 357 | 1533 | ALW | rm lf | * | 0.612 | 0.77908 | 4 |  |  | 25 | -12.2 | n |  | 5.7 | Zero | 18 | Zero | 4.5
2024-12-23 | 358 | 1721 | ALW | rm lf | * | 0.596 | 0.77655 | 4 |  |  | 12 | -12.2 | y |  | 5.7 | Zero | 18 | Zero | 4.5 | pass
2024-12-24 | 359 | 1326 | ALW | rm lf | * | 0.592 | 0.77485 | 4 |  |  | 19 | -12.2 | n |  | 5.7 | Zero | 18 | Zero | 4.5
2024-12-25 | 360 | 1426 | ALW | rm lf | * | 0.588 | 0.77108 | 4 |  |  | 15 | -12.2 | n |  | 5.7 | Zero | 18 | Zero | 4.5
2024-12-26 | 361 | 1652 | ALW | rm lf | * | 0.609 | 0.76688 | 4 |  |  | 17 | -12.2 | n |  | 5.6 | Zero | 18 | Zero | 4.5
2024-12-27 | 362 | 1429 | ALW | rm lf | * | 0.593 | 0.76009 | 4 |  |  | 29 | -12.2 | n |  | 5.7 | Zero | 18 | Zero | 4.5
2024-12-28 | 363 | 1156 | ALW | rm lf | * | 0.599 | 0.7521 | 4 |  |  | 8 | -12.2 | n |  | 5.6 | Zero | 18 | Zero | 4.5
2024-12-29 | 364 | 1359 | ALW | rm lf | * | 0.607 | 0.74899 | 4 |  |  | 35 | -12.2 | n |  | 5.6 | Zero | 18 | Zero | 4.5
2024-12-30 | 365 | 1512 | ALW | rm lf | * | 0.607 | 0.74191 | 4 |  |  | 52 | -12.2 | n |  | 5.6 | Zero | 18 | Zero | 4.5
2024-12-31 | 366 | 1301 | ALW | rm lf | * | 0.604 | 0.73903 | 4 |  |  | 54 | -12.2 | y |  | 5.6 | Zero | 18 | Zero | 4.5 | pass
Note on G1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: A11 GMD CLAP Box: Flow (Between 0.59-0.61 lpm) Reply: Update: 0.59-0.63 lpm Reply: Update: ~0.7 (>0.65/<0.75)
Note on H1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Green Transmittance: (When it drops below 0.7 it should change to a new spot) Reply: Has a total of 8 spots. Filter should be changed when spot 8 is nearing 0.7.
Note on I1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Spot
Note on J1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Concentration
Note on K1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Dewpoint
Note on L1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Concentration Reply: Per Erin: above 0 is normal. Concentration seems to trend lower in the winter. If hits 0, ask Erin and consider opening the lines to room air (search concentration in email and consult Erin)
Note on M1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Dewpoint Reply: Should remain above 10 per protocol
Note on O1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Alternate between Nephelometers so that each neph gets a span check every other month but you complete a span check every month. Per protocol
Note on P1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: S11 TSI Box: Lamp (A) Reply: Should be <7.0. If bulb exceeds 7, it is time to change the bulb. You should receive an automate message about this in the ersl. Per protocol
Note on Q1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Flags Reply: Forge Acquisition box title will flash and identify flag.
Note on R1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: On physical Neph stucture Reply: Per protocol. 18 is marked with sharpie.
Note on S1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Check the flag status on Forge (computer)
Note on T1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Check rotameter on Ecotech. Okay if it changes 5> flow >4. Reply: Per protocol
Note on U1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Check sys status on ecotech screen, verify everything is “pass”

## Surface Ozone
Date | Julian Day | Time (UTC) | Observer's Initials | Inlet | Panel Display (30-60 ppb) | Bench Temp (<45C) | Pressure (>620 hPA) | Flow rate (LPM) (>0.35) | Flow rate (LPM) (>.35) | Intensity Cell A (>45,000) | Intensity Cell B (>45,000) | Comments (Alarms, Calibration Levels...)
2024-01-01 | 1 | 1614 | GM | rm lf | 38.7 | 32.3 | 667.9 | 0.461 | 0.495 | 83802 | 103121
2024-01-02 | 2 | [10 empty] | Replaced inlet particle filter. O3 Calibration: 1303Z-1348Z, Zero:0.9, L1:168.5, L2:41.2
2024-01-03 | 3 | 1705 | GM | rm lf | 35.9 | 33.6 | 663.4 | 0.459 | 0.493 | 84259 | 103618
2024-01-04 | 4
2024-01-05 | 5 | [10 empty] | Did not make it out to AWO after ICESat
2024-01-06 | 6 | 1353 | GM | rm lf | 45.5 | 33.3 | 662.5 | 0.459 | 0.493 | 84114 | 103516
2024-01-07 | 7
2024-01-08 | 8 | 1732 | GM | lf | 45.1 | 34.2 | 668.6 | 0.462 | 0.494 | 84263 | 103599 | Did not go up on roof to clean inlet due to 30kn winds
2024-01-09 | 9 | [10 empty] | O3 Calibration: 1730-1816Z: Zero:0.1, L1:168.8, L2:40.5
2024-01-10 | 10 | 1313 | GM | lf | 41.5 | 33.6 | 674.6 | 0.464 | 0.498 | 84060 | 103411 | Did not go up on roof to clean inlet due to 30kn winds
2024-01-11 | 11
2024-01-12 | 12 | 1618 | GM | rm lf | 44.8 | 32.9 | 680.2 | 0.466 | 0.5 | 83711 | 103010
2024-01-13 | 13
2024-01-14 | 14 | 1534 | GM | rm lf | 46.8 | 33.9 | 681.9 | 0.467 | 0.501 | 84060 | 103418
2024-01-15 | 15 | [10 empty] | The Snow Beast was ran for about 15min on station at the beginning of North winds 1600-1615Z
2024-01-16 | 16 | 1346 | GM | rm lf | 47.3 | 33.4 | 676.3 | 0.464 | 0.498 | 83823 | 103171 | O3 Calibration: 1753-1839Z: Zero:-0.2, L1:168.7, L2:41.5
2024-01-17 | 17
2024-01-18 | 18 | 1657 | GM | rm lf | 44.8 | 32.6 | 656.9 | 0.456 | 0.489 | 83596 | 102941
2024-01-19 | 19
2024-01-20 | 20 | 1201 | GM | rm lf | 40.2 | 33.4 | 640.6 | 0.449 | 0.483 | 83862 | 103256
2024-01-21 | 21
2024-01-22 | 22 | 1429 | GM | rm lf | 41.8 | 33.5 | 641.3 | 0.45 | 0.483 | 83848 | 103187
2024-01-23 | 23 | [10 empty] | O3 Calibration: 1409-1455Z: Zero:0.0, L1:169.5, L2:41.4
2024-01-24 | 24 | 1243 | GM | rm lf | 44.1 | 33.1 | 641.4 | 0.45 | 0.483 | 83869 | 103257
2024-01-25 | 25
2024-01-26 | 26 | 1313 | GM | rm lf | 41.3 | 33.2 | 638.6 | 0.448 | 0.482 | 83817 | 103184
2024-01-27 | 27
2024-01-28 | 28 | 1420 | GM | clr | 42.4 | 33.3 | 646.2 | 0.451 | 0.485 | 83858 | 103216
2024-01-29 | 29
2024-01-30 | 30 | 1300 | GM | rm lf | 42.3 | 31.9 | 645.5 | 0.451 | 0.485 | 83502 | 102888 | O3 Calibration: 1407-1453Z: Zero:-0.2, L1:168.8, L2: 41.0
2024-01-31 | 31
2024-02-01 | 32 | 1352 | GM | rm lf | 36.4 | 32.4 | 645.5 | 0.453 | 0.486 | 83484 | 102834
2024-02-02 | 33
2024-02-03 | 34 | 1259 | GM | rm lf | 38.4 | 33 | 654.1 | 0.455 | 0.489 | 83469 | 102723
2024-02-04 | 35
2024-02-05 | 36 | 1425 | GM | rm lf | 42.5 | 32.9 | 665.3 | 0.46 | 0.493 | 83477 | 102731 | O3 Calibration: 1747-1833Z: Zero:-0.1, L1:168.0, L2: 41.2
2024-02-06 | 37
2024-02-07 | 38 | 1435 | GM | rm lf | 41.4 | 32.4 | 670.1 | 0.46 | 0.495 | 83225 | 102478
2024-02-08 | 39
2024-02-09 | 40 | 1538 | GM | rm lf | 43.9 | 32.9 | 659.7 | 0.457 | 0.49 | 83646 | 102959
2024-02-10 | 41
2024-02-11 | 42 | 1433 | DR | rm lf | 41.5 | 33.7 | 661.3 | 0.457 | 0.491 | 83734 | 103064
2024-02-12 | 43 | [10 empty] | O3 Calibration: 1339-1427Z: Zero: 0.3, L1: 169.7, L2: 42.5
2024-02-13 | 44 | [10 empty] | Did not make it out to AWO after ICESat
2024-02-14 | 45 | [10 empty] | Did not make it out to AWO due to Con 1 conditions
2024-02-15 | 46 | 1401 | GM | rm lf | 43.3 | 32.3 | 658.6 | 0.457 | 0.49 | 83147 | 102421
2024-02-16 | 47
2024-02-17 | 48 | 1816 | DR |  | 48.2 | 34.2 | 668.1 | 0.461 | 0.494 | 83627 | 102907
2024-02-18 | 49
2024-02-19 | 50 | [10 empty] | O3 Calibration: 1214-1311Z: Zero: 0.2, L1: 169.8, L2: 41.8
2024-02-20 | 51
2024-02-21 | 52 | 1522 | DR | rm lf | 42.7 | 34.9 | 655.8 | 0.455 | 0.489 | 83903 | 103254
2024-02-22 | 53 | 1614 | DR | rm lf | 47.8 | 34.6 | 658.5 | 0.456 | 0.49 | 83911 | 103256
2024-02-23 | 54 | 1303 | DR | rm lf | 47.5 | 34.7 | 659.9 | 0.457 | 0.491 | 83607 | 102924
2024-02-24 | 55 | 1258 | DR | rm lf | 45.8 | 33.7 | 657.3 | 0.456 | 0.49 | 83348 | 102587
2024-02-25 | 56 | 1632 | DR | rm lf | 46.1 | 34.4 | 653 | 0.454 | 0.487 | 83540 | 102810
2024-02-26 | 57 | 1345 | DR | rm lf | 45.5 | 34.9 | 645.3 | 0.451 | 0.484 | 83812 | 103141 | O3 Calibration: 1245-1330Z Zero: -0.2, L1: 170.6, L2: 41.4
2024-02-27 | 58 | 1406 | DR | rm lf | 42.7 | 33.6 | 649.5 | 0.453 | 0.486 | 83384 | 102688
2024-02-28 | 59 | 1300 | DR | rm lf | 40.3 | 32.7 | 648.6 | 0.452 | 0.486 | 83260 | 102559
2024-02-29 | 60 | 1415 | DR | rm lf | 45.6 | 33.4 | 657 | 0.455 | 0.489 | 83377 | 102655
2024-03-01 | 61 | 1309 | DR | clr | 45.4 | 33.8 | 666.3 | 0.459 | 0.493 | 83370 | 102650
2024-03-02 | 62 | 1041 | DR | rm lf | 46.7 | 34.8 | 671.9 | 0.461 | 0.495 | 83569 | 102813
2024-03-03 | 63 | 1639 | DR | clr | 52.5 | 35 | 668.9 | 0.461 | 0.495 | 83624 | 102942
2024-03-04 | 64 | 1617 | DR | rm lf | 44.1 | 34.9 | 667.3 | 0.459 | 0.493 | 83580 | 102887 | O3 Calibration: 1230-1319Z Zero: 0.4, L1: 171.0, L2: 42.0
2024-03-05 | 65 | 1424 | DR | clr | 44.4 | 34.5 | 670 | 0.461 | 0.495 | 83463 | 102744
2024-03-06 | 66 | 1259 | DR | rm lf | 47.3 | 34.7 | 673.1 | 0.461 | 0.495 | 83556 | 102839
2024-03-07 | 67 | 1540 | DR | clr | 47.6 | 34.8 | 675.3 | 0.463 | 0.497 | 83605 | 102928
2024-03-08 | 68 | 1339 | DR | rm lf | 43.6 | 34.8 | 674.3 | 0.462 | 0.497 | 83407 | 102641
2024-03-09 | 69 | 1237 | DR | clr | 46.1 | 34.9 | 673.5 | 0.462 | 0.496 | 83517 | 102797
2024-03-10 | 70 | 1538 | DR | rm lf | 46.2 | 34.8 | 672.1 | 0.462 | 0.495 | 83419 | 102678
2024-03-11 | 71 | 1542 | DR | rm lf | 46.6 | 35.5 | 675.3 | 0.462 | 0.497 | 83697 | 103046 | O3 Calibration: 1424-1514Z Zero: -0.2, L1: 169.5, L2: 42.1
2024-03-12 | 72 | [10 empty] | Did not make it out to AWO after ICESat
2024-03-13 | 73 | 1250 | DR | rm lf | 45.7 | 34 | 670.5 | 0.461 | 0.494 | 83144 | 102347
2024-03-14 | 74 | 1634 | DR | clr | 45.2 | 34.1 | 661.9 | 0.456 | 0.49 | 83151 | 102392
2024-03-15 | 75 | 1403 | DR | rm lf | 46.9 | 33.7 | 660.1 | 0.456 | 0.49 | 83119 | 102362
2024-03-16 | 76 | 1314 | DR | rm lf | 46.3 | 33 | 658.9 | 0.456 | 0.489 | 82905 | 102167
2024-03-17 | 77 | 1904 | DR | rm lf | 43.7 | 34.1 | 659.5 | 0.456 | 0.49 | 83248 | 102517
2024-03-18 | 78 | 1739 | DR | clr | 46.9 | 34.9 | 661.7 | 0.457 | 0.491 | 83640 | 102990 | O3 Calibration: 1624-1732Z Zero: -0.5, L1: 170.6, L2: 41.4
2024-03-19 | 79 | [10 empty] | Did not make it out to AWO due to CON1
2024-03-20 | 80 | 1130 | DR | clr | 43.8 | 34.7 | 663.1 | 0.457 | 0.491 | 83341 | 102586
2024-03-21 | 81 | 1635 | DR | clr | 47.6 | 34.8 | 662.6 | 0.457 | 0.491 | 83241 | 102440
2024-03-22 | 82 | 1218 | DR | rm lf | 41.1 | 34.4 | 671.1 | 0.46 | 0.495 | 83165 | 102388
2024-03-23 | 83 | 1803 | DR | rm lf | 43 | 34.5 | 673.4 | 0.461 | 0.495 | 83107 | 102312
2024-03-24 | 84 | 1638 | DR | rm lf | 45.5 | 34.6 | 673.3 | 0.462 | 0.495 | 83265 | 102514
2024-03-25 | 85 | 1543 | DR | rm lf | 49.1 | 34.7 | 671.5 | 0.461 | 0.495 | 83272 | 102553 | O3 Calibration: 1546-1632Z Zero: 0.2, L1: 170.6, L2: 42.0
2024-03-26 | 86 | 1330 | DR | clr | 50.6 | 34.3 | 670.2 | 0.46 | 0.494 | 83077 | 102290
2024-03-27 | 87 |  | DR | rm lf
2024-03-28 | 88 | 1708 | DR | rm lf | 50 | 34.5 | 665 | 0.458 | 0.492 | 83223 | 102471
2024-03-29 | 89 | 1303 | DR | rm lf | 43.9 | 33.9 | 670.6 | 0.46 | 0.494 | 83007 | 102221
2024-03-30 | 90 | 1308 | DR | rm lf | 49.2 | 33.6 | 675.9 | 0.462 | 0.497 | 82862 | 102073
2024-03-31 | 91 | 1542 | DR | rm lf | 53.6 | 34 | 681.5 | 0.464 | 0.499 | 83116 | 102381
2024-04-01 | 92 | 1504 | DR | rm lf | 54.1 | 34.1 | 685.5 | 0.467 | 0.5 | 82987 | 102186 | O3 Calibration: 1556-1641Z Zero: 0.0, L1: 170.1, L2: 42.5
2024-04-02 | 93 | 1627 | DR | clr | 56.2 | 34.6 | 682.6 | 0.465 | 0.499 | 83108 | 102303
2024-04-03 | 94 |  | DR | [8 empty] | Did not make it out to AWO after ICESat
2024-04-04 | 95 | 1155 | DR | rm lf | 42.7 | 34 | 671.4 | 0.46 | 0.494 | 82967 | 102162
2024-04-05 | 96 | 1209 | DR | rm lf | 43.8 | 33.9 | 672.9 | 0.461 | 0.494 | 82894 | 102090
2024-04-06 | 97 | 1209 | DR | rm lf | 49.3 | 34.2 | 667.4 | 0.458 | 0.492 | 82970 | 102192
2024-04-07 | 98 | 1526 | DR | rm lf | 50.7 | 34.6 | 661 | 0.456 | 0.49 | 83025 | 102206
2024-04-08 | 99 | 1140 | DR | rm lf | 48.6 | 33.7 | 661.8 | 0.456 | 0.49 | 82942 | 102162 | O3 Calibration: 1209-1300Z Zero: -0.0, L1: 170.9, L2: 42.1
2024-04-09 | 100 | 1717 | DR | clr | 46.8 | 35.1 | 661.5 | 0.455 | 0.49 | 82915 | 102081
2024-04-10 | 101 | 1422 | DR | clr | 46.5 | 34 | 659.9 | 0.455 | 0.489 | 82701 | 101786
2024-04-11 | 102 | 1201 | DR | rm lf | 46.6 | 33.5 | 659.8 | 0.455 | 0.49 | 82631 | 101772
2024-04-12 | 103 | 1126 | DR | rm lf | 49.1 | 33.6 | 660.2 | 0.456 | 0.489 | 82574 | 101628
2024-04-13 | 104 | 1105 | DR | rm lf | 42.2 | 33.7 | 660.9 | 0.455 | 0.49 | 82661 | 101792
2024-04-14 | 105 | 1854 | DR | rm lf | 46.2 | 35 | 668.5 | 0.459 | 0.493 | 82963 | 102060
2024-04-15 | 106 | 1618 | DR | clr | 49.1 | 35.2 | 669.5 | 0.459 | 0.493 | 83145 | 102345 | O3 Calibration: 1530-1616Z Zero: 0.9, L1: 170.9, L2: 42.8
2024-04-16 | 107 | 1204 | DR | clr | 50.6 | 33.5 | 665.5 | 0.458 | 0.491 | 82568 | 101612
2024-04-17 | 108 | 2038 | DR | rm lf | 47.2 | 35.5 | 669.3 | 0.458 | 0.493 | 83170 | 102341
2024-04-18 | 109 | 1513 | DR | rm lf | 47.2 | 34.3 | 668.5 | 0.459 | 0.492 | 82875 | 101970
2024-04-19 | 110 | 1742 | DR | rm lf | 49.6 | 35.2 | 661.8 | 0.455 | 0.489 | 82911 | 102010
2024-04-20 | 111 | 1620 | DR | clr | 46.7 | 35.3 | 660.2 | 0.455 | 0.489 | 83228 | 102353
2024-04-21 | 112 | 1341 | DR | rm lf | 47.9 | 35.5 | 669.3 | 0.459 | 0.493 | 83060 | 102189
2024-04-22 | 113 | 1213 | DR | clr | 45.4 | 35.4 | 682.2 | 0.464 | 0.498 | 83189 | 102325 | O3 Calibration: 1114-1212Z Zero: -0.3, L1: 170.0, L2: 41.5
2024-04-23 | 114 | 1019 | DR | clr | 52.9 | 35 | 684.1 | 0.465 | 0.499 | 82804 | 101878
2024-04-24 | 115 | 1729 | DR | clr | 51.8 | 36 | 682.3 | 0.463 | 0.498 | 83014 | 102011
2024-04-25 | 116 |  | DR | [8 empty] | HM - did not make it out to AWO.
2024-04-26 | 117 | 1046 | DR | rm lf | 50 | 35.1 | 681.9 | 0.463 | 0.498 | 82884 | 101908
2024-04-27 | 118 | 1129 | DR | clr | 55.1 | 37.2 | 680.2 | 0.462 | 0.497 | 83259 | 102360
2024-04-28 | 119 | 1454 | DR | clr | 52.1 | 36.5 | 679.3 | 0.462 | 0.497 | 83215 | 102207
2024-04-29 | 120 | 1713 | DR | rm lf | 57.1 | 37.2 | 680.5 | 0.463 | 0.497 | 83416 | 102504 | O3 Calibration: 1534-1627Z Zero: 0.0, L1: 170.9, L2: 42.4
2024-04-30 | 121 | 1229 | DR | clr | 54.9 | 35.4 | 677.7 | 0.461 | 0.496 | 83007 | 102041 | Line has some ice build up
2024-05-01 | 122 | 1532 | DR | clr | 53.4 | 37.1 | 672.7 | 0.459 | 0.494 | 82909 | 101897
2024-05-02 | 123 | 1202 | DR | clr | 48.2 | 36 | 674.5 | 0.46 | 0.494 | 82543 | 101516
2024-05-03 | 124 | 1644 | DR | rm lf | 47.4 | 36.5 | 676.1 | 0.461 | 0.495 | 83307 | 102380
2024-05-04 | 125 | 1636 | DR | rm lf | 47.1 | 35.7 | 677.4 | 0.461 | 0.495 | 82859 | 101875
2024-05-05 | 126 |  | DR | [8 empty] | HW - did not make it out to AWO.
2024-05-06 | 127 | 1745 | DR | clr | 43.5 | 35 | 663.3 | 0.455 | 0.49 | 82860 | 101657 | O3 Calibration: 1056-1145Z Zero: 0.3, L1: 170.1, L2: 41.8
2024-05-07 | 128 | 1532 | DR | rm lf | 44.5 | 34.8 | 663.9 | 0.456 | 0.49 | 82530 | 101468
2024-05-08 | 129 | 1027 | DR | rm lf | 43.8 | 34.2 | 663.5 | 0.456 | 0.49 | 82387 | 101278
2024-05-09 | 130 | 2002 | DR | rm lf | 41.8 | 34.8 | 665.8 | 0.456 | 0.49 | 82457 | 101339
2024-05-10 | 131 |  | DR | [8 empty] | Did not make it out to AWO after ICESat
2024-05-11 | 132 | 1630 | DR | rm lf | 47.4 | 34.8 | 669.4 | 0.457 | 0.492 | 82374 | 101287
2024-05-12 | 133 | 1344 | DR | rm lf | 50.5 | 34.9 | 671.9 | 0.458 | 0.493 | 82616 | 101574
2024-05-13 | 134 | 1451 | DR | clr | 44.3 | 35.1 | 672.3 | 0.459 | 0.493 | 82496 | 101377 | O3 Calibration: 1453-1551Z Zero: -0.3, L1: 170.8, L2: 42.5
2024-05-14 | 135 | 1533 | DR | clr | 46.4 | 35.5 | 666.3 | 0.456 | 0.49 | 82607 | 101487
2024-05-15 | 136 |  | DR | [8 empty] | Did not make it out to AWO after ICESat
2024-05-16 | 137 | 1447 | DR | clr | 50.4 | 34.6 | 678.2 | 0.461 | 0.496 | 82412 | 101312
2024-05-17 | 138 | 1207 | DR | clr | 43.1 | 33.9 | 676.6 | 0.461 | 0.495 | 81939 | 100724
2024-05-18 | 139 | 1447 | DR | rm lf | 45 | 34.7 | 676.2 | 0.461 | 0.494 | 82386 | 101194
2024-05-19 | 140 | 1847 | DR | rm lf | 41 | 35 | 675.1 | 0.46 | 0.494 | 82295 | 101090
2024-05-20 | 141 | 1220 | DR | rm lf | 38.45 | 34.4 | 674.5 | 0.459 | 0.494 | 82094 | 100890
2024-05-21 | 142 |  | DR | [8 empty] | HM - did not make it out to AWO.
2024-05-22 | 143 | 1118 | DR | rm lf | 39.7 | 34.7 | 667.9 | 0.456 | 0.491 | 82252 | 101061 | O3 Calibration: 1014-1110Z Zero: -0.1, L1: 169.2, L2: 42.0
2024-05-23 | 144 | 1150 | DR | rm lf | 43.2 | 34.2 | 663.5 | 0.455 | 0.489 | 82055 | 100777
2024-05-24 | 145 | 959 | DR | rm lf | 42.5 | 34.4 | 665.8 | 0.456 | 0.49 | 82348 | 101129
2024-05-25 | 146 | 1556 | DR | clr | 40.8 | 35.2 | 669.1 | 0.457 | 0.491 | 82311 | 101077
2024-05-26 | 147 | 730 | DR | clr | 43.7 | 35.1 | 670.9 | 0.456 | 0.49 | 82368 | 101103 | Unplanned power outage - storm conditions.
2024-05-27 | 148 |  | DR | clr | [7 empty] | O3 Calibration: 1211-1256Z Zero: -0.1, L1: 171.8, L2: 42.7
2024-05-28 | 149 |  | DR | [8 empty] | Day off.
2024-05-29 | 150 |  | DR | [8 empty] | Con 1 - didn't go out to AWO.
2024-05-30 | 151 | 1444 | DR | clr | 48.7 | 35.3 | 660.3 | 0.453 | 0.487 | 82775 | 100997
2024-05-31 | 152 |  | DR
2024-06-01 | 153 | 954 | DR | rm lf | 26.2 | 34.8 | 662.5 | 0.454 | 0.489 | 81925 | 100573
2024-06-02 | 154 | 1103 | DR | rm lf | 23.03 | 34.8 | 667 | 0.456 | 0.49 | 82191 | 100904
2024-06-03 | 155 | 1629 | DR | clr | 27.3 | 35 | 676.9 | 0.46 | 0.494 | 82344 | 101050 | O3 Calibration: 1631-1722Z Zero: 0.4, L1: 169.5, L2: 41.9
2024-06-04 | 156 |  | DR | [8 empty] | Did not make it out to AWO after ICESat
2024-06-05 | 157 | 1055 | DR | rm lf | 44 | 34.2 | 680.2 | 0.462 | 0.496 | 81790 | 100364
2024-06-06 | 158 | 1455 | CW | clr | 46.8 | 34.8 | 680.9 | 0.461 | 0.496 | 82182 | 100857
2024-06-07 | 159 | 1439 | CW | clr | 47.04 |  | 680.3 | 0.461 | 0.496 | 81739 | 100234 | Powered off instrument at 1458 UTC for Welding operations, bagged and sealed intake.
2024-06-08 | 160 | 950 | CW | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a
2024-06-09 | 161 | 1116 | CW | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a
2024-06-10 | 162 | 1005 | CW | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a
2024-06-11 | 163 | 1053 | CW | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a
2024-06-12 | 164 | 1006 | CW | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a
2024-06-13 | 165 | 1027 | CW | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a
2024-06-14 | 166 | 942 | CW | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a
2024-06-15 | 167 | 1255 | CW | clr | 43.1 | 18.4 | 682.1 | 0.471 | 0.502 | 85524 | 103680 | Powered on inlet at 1255 and powered off at 1208 UTC. Powered on instrument at 1308. O3 calibration: 1322-1410Z Zero: -2.4, L1:165.4, L2: 40.3
2024-06-16 | 168 | 1508 | CW | clr | 45.7 | 35.6 | 682.7 | 0.46 | 0.495 | 82377 | 100782
2024-06-17 | 169 | 1155 | CW | clr | 46 | 35.4 | 683.4 | 0.461 | 0.495 | 82330 | 100834
2024-06-18 | 170 | 1001 | CW | rm lf | 49.13 | 35.2 | 678.7 | 0.459 | 0.493 | 82028 | 100431
2024-06-19 | 171 | 1638 | CW | rm lf | 45.92 | 35.4 | 675.8 | 0.457 | 0.492 | 82179 | 100669 | Flag 1C100000
2024-06-20 | 172 | 945 | CW | n/a | 51.37 | 35.2 | 675.3 | 0.458 | 0.493 | 82054 | 100577 | Flag 1C100000. Did not go onto AWO roof due to lift operations
2024-06-21 | 173 | 1224 | CW | n/a | 46.9 | 36.3 | 672.1 | 0.456 | 0.491 | 82337 | 100889
2024-06-22 | 174
2024-06-23 | 175 | 1333 | CW | rm lf | 41.91 | 35.6 | 688.7 | 0.455 | 0.489 | 82344 | 100902 | O3 calibration: 1222-1330Z Zero: 0.0, L1: 169.0, L2: 41.4. Removed light frost along intake line and cap.
2024-06-24 | 176 | 1453 | CW | rm lf | 45.8 | 34.8 | 667.5 | 0.454 | 0.489 | 81943 | 100444
2024-06-25 | 177 | 1416 | CW | rm lf | 48.2 | 35.2 | 669.4 | 0.456 | 0.49 | 82043 | 100533
2024-06-26 | 178 | 1129 | CW | clr | 45.4 | 34.8 | 673.4 | 0.457 | 0.491 | 81982 | 100451
2024-06-27 | 179 | 1140 | CW | rm lf | 44.5 | 31.1 | 676.6 | 0.458 | 0.493 | 81915 | 100331
2024-06-28 | 180 | 1047 | CW | clr | 40.2 | 35.3 | 676.1 | 0.458 | 0.493 | 82039 | 100518
2024-06-29 | 181 | 955 | CW | clr | 59 | 35 | 671.3 | 0.456 | 0.491 | 81864 | 100277 | O3 Calibration: 0957- 0943. Zero: -0.5, L1: 169.0, L2: 41.3. Power outage during UPS maintenance, UPS instruments affected from 1128-1134Z.
2024-06-30 | 182 | 1320 | CW | clr | 47.3 | 34.4 | 667.9 | 0.455 | 0.489 | 81700 | 100215
2024-07-01 | 183 | 1541 | CW | clr | 39.6 | 34.5 | 666.6 | 0.454 | 0.489 | 81640 | 100041
2024-07-02 | 184 | 1510 | CW | clr | 48.58 | 35.1 | 671 | 0.456 | 0.491 | 81930 | 100355 | Changed O3 inlet filter at 1615.
2024-07-03 | 185 | 947 | CW | clr | 50.5 | 34.7 | 680.1 | 0.46 | 0.495 | 81854 | 100255
2024-07-04 | 186 | [10 empty] | Did not make it to AWO after ICESAT.
2024-07-05 | 187 | 1422 | CW | clr | 43.62 | 35.5 | 688.1 | 0.463 | 0.498 | 82041 | 100425
2024-07-06 | 188 | 1756 | CW | clr | 52.86 | 36 | 683.9 | 0.461 | 0.496 | 82166 | 100602 | O3 flag 1C100000
2024-07-07 | 189 | 1300 | CW | clr | 65.5 | 35.7 | 680.9 | 0.46 | 0.494 | 82210 | 100620 | O3 Calibration: 1202-1255. Zero: 0.0, L1: 167.2, L2: 41.9
2024-07-08 | 190 | 944 | CW | rm lf | 52.58 | 34.4 | 679.1 | 0.46 | 0.494 | 81557 | 99883
2024-07-09 | 191 | 955 | CW | rm lf | 55.5 | 34.8 | 675.3 | 0.458 | 0.492 | 81693 | 100014
2024-07-10 | 192 | 1406 | CW | rm lf | 47.39 | 35 | 671.1 | 0.456 | 0.49 | 81740 | 100025
2024-07-11 | 193 | 1213 | CW | clr | 55.06 | 34.4 | 672.3 | 0.456 | 0.491 | 81691 | 100037
2024-07-12 | 194 | 1606 | CW | clr | 34.94 | 34.6 | 675.1 | 0.458 | 0.492 | 81536 | 99774
2024-07-13 | 195 | 1702 | CW | clr
2024-07-14 | 196 | 1720 | CW | clr | 44.01 | 35.3 | 675.9 | 0.457 | 0.492 | 81770 | 100069
2024-07-15 | 197 | 1124 | CW | clr
2024-07-16 | 198 | 1304 | CW | clr | 35.57 | 35.5 | 687.7 | 0.462 | 0.497 | 81798 | 100055 | O3 calibration: 1146. Zero: 0.2, L1: 168.5, L2: 41.3.
2024-07-17 | 199 | 1431 | CW | clr | 44.88 | 35.3 | 691.5 | 0.464 | 0.5 | 81758 | 99969
2024-07-18 | 200 | 1726 | CW | rm lf
2024-07-19 | 201 | 1309 | CW | clr | 50.23 | 35.1 | 691.8 | 0.464 | 0.499 | 81697 | 99874
2024-07-20 | 202 | 1349 | CW | clr | 46.29 | 35.2 | 687 | 0.462 | 0.497 | 81662 | 99872
2024-07-21 | 203 | 1307 | CW
2024-07-22 | 204 | 1256 | CW | rm lf | 44.12 | 34.7 | 679.8 | 0.459 | 0.494 | 81203 | 99284
2024-07-23 | 205 | 1446 | CW | rm lf | [7 empty] | O3 calibration: 1350-1438. Zero: -0.2, L1: 169.6, L2: 48.6.
2024-07-24 | 206 | 1518 | CW | clr | 49.04 | 35.3 | 671.5 | 0.455 | 0.49 | 81562 | 99706
2024-07-25 | 207 | 1240 | CW | clr | 50.83 | 33.2 | 672.1 | 0.456 | 0.491 | 81347 | 99466
2024-07-26 | 208 | 1551 | CW | clr
2024-07-27 | 209 | 1033 | CW | clr | 37.81 | 35.8 | 672.3 | 0.454 | 0.488 | 81846 | 100027 | Pressure looks 5.6 off from NOAA measurment
2024-07-28 | 210 | 1205 | CW | clr
2024-07-29 | 211 | 1244 | CW | clr | 35.95 | 35.3 | 671.8 | 0.452 | 0.487 | 81421 | 99509
2024-07-30 | 212 | 1400 | CW | clr
2024-07-31 | 213 | 1657 | CW | clr | 41.44 | 35.6 | 685.3 | 0.451 | 0.485 | 81477 | 99540
2024-08-01 | 214 | 1459 | CW | rm lf | [7 empty] | O3 calibration @ 1457. Zero: 0.1, L1: 168.3, L2 38.0
2024-08-02 | 215 | 1656 | CW | rm lf | 12.01 | 34.9 | 691.8 | Zero | 0.165 | 81340 | 99357 | Pressure ~8 off, inlet is clear of frost. Cell A flow 0lpm, concentration below lower threshold.
2024-08-03 | 216 | 1506 | CW | rm lf | 6 |  |  | Zero |  |  |  | Concentration below lower threshold. Cell A & B flow exceeding lower alarm threshold.
2024-08-04 | 217 | 1747 | CW | rm lf | -3.89 | 37.2 | 692.3 | Zero | 0.135 | 82016 | 100072
2024-08-05 | 218 | [10 empty] | O3 pump maintenance: 1448- instrument off for pump repair 1448- pump was making noise and spinning smoothly so we'll replace the diaphram and see if that fixes things 1630- O3 back on 1715- O3 back to normal sampling ICESat
2024-08-06 | 219 | 1039 | CW | rm lf | 51.2 | 35.3 | 678.2 | 0.469 | 0.504 | 81341 | 99342 | O3 calibration: 1036-1123. Zero: 0.5, L1: 168.4 L2: 42.4.
2024-08-07 | 220 | 1130 | CW | rm lf
2024-08-08 | 221 | 1017 | CW | clr | 57.61 | 35.4 | 675.4 | 0.467 | 0.583 | 81355 | 99341
2024-08-09 | 222 | 1302 | CW | rm lf
2024-08-10 | 223 | 1118 | CW | rm lf | 54.31 | 35.6 | 671.1 | 0.465 | 0.501 | 81540 | 99536
2024-08-11 | 224 | 1843 | CW | clr
2024-08-12 | 225 | 1146 | CW | clr | 35.6 | 35.6 | 680.7 | 0.47 | 0.505 | 81205 | 99130 | Flag 1C10000
2024-08-13 | 226 | 1000 | CW | rm lf | [7 empty] | Blip in data @ 0658-0727Z. O3 calibration: 1627-1713. Zero: 0.4, L1: 167.6, L2: 42.7.
2024-08-14 | 227 | [10 empty] | ICESat
2024-08-15 | 228 | 1507 | CW | clr | 47.62 | 34.5 | 670.5 | 0.465 | 0.501 | 81019 | 98853
2024-08-16 | 229 | 1004 | CW | clr
2024-08-17 | 230 | 1537 | CW | clr | 39.55 | 35.2 | 676.1 | 0.468 | 0.503 | 81283 | 99156
2024-08-18 | 231 | 1845 | CW | rm lf
2024-08-19 | 232 | 1541 | CW | rm lf | 40.08 | 35 | 674.6 | 0.468 | 0.503 | 81034 | 98855
2024-08-20 | 233 | 1356 | CW | rm lf | [7 empty] | O3 calibration: 1215-1245. Zero:-0.5, L1:168.5, L2: 42.9
2024-08-21 | 234 | 1226 | CW | rm hf | 51.31 | 34.6 | 675.5 | 0.468 | 0.503 | 80847 | 98525
2024-08-22 | 235 | 1909 | CW | rm lf
2024-08-23 | 236 | 1035 | CW | rm lf | 46.18 | 35.5 | 676.2 | 0.468 | 0.503 |  |  | AWO internet outage from 2024-08-22 at 2224Z- 2024-08-23 at 0907Z.
2024-08-24 | 237 | 1030 | CW | rm lf
2024-08-25 | 238 | 1428 | CW | rm lf | 46.88 | 35 | 674.5 | 0.468 | 0.503 | 80951 | 98642
2024-08-26 | 239 | 1215 | CW | rm lf
2024-08-27 | 240 | 1241 | CW | rm lf | 46.9 | [6 empty] | O3 calibration: 1229 - 1329. Zero: -0.1, L1: 169.6, L2:42.5
2024-08-28 | 241 | 1355 | CW | rm lf
2024-08-29 | 242 | 1325 | CW | clr | 41.88 | 35.7 | 665.7 | 0.464 | 0.499 | 81165 | 98789
2024-08-30 | 243 | 918 | CW | clr
2024-08-31 | 244 | 1220 | CW | clr | 40.38 | 36.2 | 669.4 | 0.465 | 0.501 | 81268 | 98916
2024-09-01 | 245 | 1252 | CW | clr
2024-09-02 | 246 | 1505 | CW | clr | 169.2 | 35.7 | 675 | 0.468 | 0.504 | 81141 | 98840 | O3 calibration: 1440 - . Zero: 0.1, L1: 169.6, L2: 43.2
2024-09-03 | 247 | [10 empty] | Did not make it to AWO after ICESat.
2024-09-04 | 248 | 1356 | CW | rm lf | 47.97 | 35.9 | 674.7 | 0.468 | 0.503 | 81016 | 99582
2024-09-05 | 249 | 1416 | CW | rm lf
2024-09-06 | 250 | 1428 | CW | rm lf | 44.53 | 36.2 | 672.7 | 0.466 | 0.502 | 81162 | 98678
2024-09-07 | 251 | [10 empty] | Did not go to AWO due to poor visibility
2024-09-08 | 252 | 1835 | CW | clr | 38.53 | 36.2 | 681 | 0.471 | 0.506 | 81169 | 98786
2024-09-09 | 253 | 1235 | CW | clr
2024-09-10 | 254 | [10 empty] | Did not go to AWO due to poor visibility
2024-09-11 | 255 | 1311 | CW | rm hf | 37.97 | 34 | 685.4 | 0.473 | 0.508 | 80517 | 97895 | 03 calibration check: 1313-1358 Zero: 0.7, L1: 169.2, L2:32.9.
2024-09-12 | 256 | 912 | CW | rm hf
2024-09-13 | 257 | 1004 | CW | rm hf | 39.05 | 34.8 | 682.9 | 0.471 | 0.507 | 80566 | 97902
2024-09-14 | 258 | 1104 | CW | rm lf
2024-09-15 | 259 | 1840 | CW | rm lf | 42.7 | 34.7 | 677 | 0.47 | 0.504 | 80572 | 97878
2024-09-16 | 260 | 1201 | CW | rm lf
2024-09-17 | 261 | 1846 | CW | rm lf | 44.08 | 35.4 | 675.4 | 0.469 | 0.503 | 80928 | 98294 | 03 calibration check: Zero: -0.3, L1: 169.9, L2: 42.0
2024-09-18 | 262 | 1439 | CW | rm lf
2024-09-19 | 263 | 1226 | CW | rm lf | 38.71 | 34.4 | 683.1 | 0.472 | 0.507 | 80319 | 97556
2024-09-20 | 264 | 1008 | CW | rm lf
2024-09-21 | 265 | 944 | CW | rm lf | 49.84 | 33.8 | 694.1 | 0.477 | 0.512 | 80010 | 97120
2024-09-22 | 266 | 1252 | CW | clr
2024-09-23 | 267 | 1255 | CW | clr | 60.43 | 34.2 | 695.5 | 0.478 | 0.512 | 80245 | 97359
2024-09-24 | 268 | 1050 | CW | lf
2024-09-25 | 269 | 1452 | CW | rm hf | 51.02 | 34.7 | 687.4 | 0.474 | 0.509 | 80491 | 97631 | O3 calibration check: Zero: -0.2, L1: 168.5, L2: 50.5
2024-09-26 | 270 | 1442 | CW | rm hf
2024-09-27 | 271 | 1001 | CW | rm lf | 52.27 | 34.9 | 677.8 | 0.469 | 0.505 | 80463 | 97600
2024-09-28 | 272 | [10 empty] | Did not make it to AWO after raising sector lines
2024-09-29 | 273 | 1303 | CW | rm lf | 33.9 | 34.8 | 668.9 | 0.466 | 0.501 | 80496 | 97611
2024-09-30 | 274 | 1109 | CW | clr
2024-10-01 | 275 | 1158 | CW | rm lf | 28.1 | 34.6 | 683.3 | 0.473 | 0.507 | 80333 | 97362 | O3 calibration check: Zero: 0.5, L1: 168.5, L2: 41.08
2024-10-02 | 276 | 1505 | CW | rm lf
2024-10-03 | 277 | 1504 | CW | clr | 40.42 | 35 | 689.7 | 0.475 | 0.51 | 80290 | 97214
2024-10-04 | 278 | 1652 | CW | clr
2024-10-05 | 279 |  | CW | [8 empty] | Did not go to AWO after ICESat
2024-10-06 | 280 | 1720 | CW | rm lf | 44.31 | 34.9 | 683.8 | 0.472 | 0.507 | 80247 | 97118
2024-10-07 | 281 | 1325 | CW | rm lf
2024-10-08 | 282 | 1011 | CW | rm lf | 38.45 | 34.5 | 686.1 | 0.474 | 0.508 | 80102 | 96990
2024-10-09 | 283 | 1211 | CW | rm lf
2024-10-10 | 284 | 1127 | CW | rm lf | 46.99 | 35.4 | 680.5 | 0.478 | 0.506 | 80196 | 97033
2024-10-11 | 285 | 1321 | CW | rm lf
2024-10-12 | 286 | 1248 | ALW | clr | 38.69 | 34.8 | 672.6 | 0.468 | 0.583 | 80069 | 96855
2024-10-13 | 287 | 1359 | ALW | clr | 46.39 | 34.8 | 668.3 | 0.466 | 0.501 | 80299 | 97135
2024-10-14 | 288 | 1557 | ALW | rm lf | 42.88 | 35 | 665.9 | 0.464 | 0.5 | 80431 | 97263
2024-10-15 | 289 | 1307 | ALW | rm lf | 43.69 | 35.6 | 667.7 | 0.466 | 0.501 | 80535 | 97440 | O3 calibration check: 1228- 1320, Zero: 0.36, L1: 170.1, L2: 42.5
2024-10-16 | 290 | 1753 | ALW | rm lf | 47.98 | 35.5 | 658.5 | 0.461 | 0.497 | 80311 | 97157
2024-10-17 | 291 | 1151 | ALW | rm lf | 45.59 | 35.2 | 655.8 | 0.46 | 0.496 | 80409 | 97252
2024-10-18 | 292 | 1343 | ALW | rm lf | - | - | - | - | - | - | - | Dell computer has "!LOST CONNECTION!", have contacted Erin.
2024-10-19 | 293 | 1047 | ALW | rm lf | 44.69 | 35.6 | 659.1 | 0.462 | 0.497 | 80090 | 96850 | Restarted Dell computer at 1047 and Forge reconnected with only the surface ozone.
2024-10-20 | 294 | 1435 | ALW | rm lf | 41.86 | 35.7 | 654.6 | 0.459 | 0.459 | 80381 | 97158
2024-10-21 | 295 | 1516 | ALW | rm lf | 42.25 | 35.8 | 658.5 | 0.461 | 0.496 | 80375 | 97134
2024-10-22 | 296 | 1351 | ALW | rm lf | 44.3 | 36.1 | 658.7 | 0.461 | 0.497 | 80302 | 97051 | O3 Calibration Check. 1250-1338Z. Zero:-0.3. L1: 170.9. L2: 43.1.
2024-10-23 | 297 | 1751 | ALW | rm lf | 35.7 | 36.1 | 662.9 | 0.463 | 0.499 | 80455 | 97179
2024-10-24 | 298 | 1502 | ALW | rm lf | 31.87 | 36 | 661.9 | 0.463 | 0.498 | 80467 | 97176
2024-10-25 | 299 | 1250 | ALW | rm lf | 37.88 | 35.7 | 657.5 | 0.461 | 0.496 | 80500 | 97196
2024-10-26 | 300 | 1254 | ALW | rm lf | 35.57 | 35 | 663.8 | 0.464 | 0.499 | 80101 | 96751
2024-10-27 | 301 | 1500 | ALW | rm lf | 36.42 | 34.3 | 664.2 | 0.464 | 0.5 | 79896 | 96487
2024-10-28 | 302 | 1530 | ALW | rm lf | 36.28 | 32.5 | 659 | 0.462 | 0.497 | 79420 | 95936
2024-10-29 | 303 | - | ALW | - | - | - | - | - | - | - | - | Con 1 - didn't go out to AWO.
2024-10-30 | 304 | 1505 | ALW | rm lf | 35.27 | 31.7 | 668.7 | 0.467 | 0.502 | 78947 | 95377 | Callibration completed 1403-1450Z. Zero: 0.2. L1: 167.1. L2: 41.7.
2024-10-31 | 305 | 1815 | ALW | rm lf | 35.54 | 32.2 | 667.8 | 0.466 | 0.501 | 79193 | 95636
2024-11-01 | 306 | 1509 | ALW | rm lf | 42 | 34.6 | 663.4 | 0.464 | 0.499 | 79922 | 96437
2024-11-02 | 307 | - | ALW | - | - | - | - | - | - | - | - | ICESat
2024-11-03 | 308 | 1515 | ALW | rm lf | 40.79 | 34.9 | 665.1 | 0.464 | 0.5 | 79957 | 96423
2024-11-04 | 309 | 1343 | ALW | - | 39.74 | 34.7 | 660.6 | 0.462 | 0.498 | 79763 | 96211 | Too windy to go to the roof
2024-11-05 | 310 | 1440 | ALW | rm lf | 36.44 | 34.9 | 666.3 | 0.465 | 0.5 | 79624 | 96011 | Weekly callibration 1400-1450Z. Zero: 0.4, L1: 171.1, L2: 42.7.
2024-11-06 | 311 | 1347 | ALW | - | 36.84 | 34.3 | 657.4 | 0.462 | 0.497 | 79772 | 96212 | Too windy to go to the roof
2024-11-07 | 312 | 1912 | ALW | - | 42.51 | 34.7 | 658.1 | 0.462 | 0.497 | 79710 | 96133 | Did not go to roof after tower climb
2024-11-08 | 313 | 1738 | ALW | rm lf | 44.84 | 34.4 | 664.9 | 0.465 | 0.5 | 79764 | 96182
2024-11-09 | 314 | 1258 | ALW | rm lf | 42.82 | 34.8 | 668.7 | 0.467 | 0.502 | 79704 | 96041
2024-11-10 | 315 | 1500 | ALW | rm lf | 43.25 | 35 | 667.8 | 0.466 | 0.501 | 79871 | 96293
2024-11-11 | 316 | 1259 | ALW | rm lf | 40.37 | 34.8 | 666.2 | 0.465 | 0.5 | 79801 | 96154
2024-11-12 | 317 | 1455 | ALW | rm lf | 42.01 | 34.9 | 659.4 | 0.462 | 0.497 | 79780 | 96129 | Weekly calibration 1535-1625Z. Z:0.9, L1: 107.8, L2: 42.5.
2024-11-13 | 318 | - | ALW | - | - | - | - | - | - | - | - | ICESat- did not go to AWO
2024-11-14 | 319 | 1558 | ALW | - | 45.45 | 32.8 | 664.7 | 0.466 | 0.501 | 79008 | 95208 | Too windy to go to the roof
2024-11-15 | 320 | 1343 | ALW | rm lf | 37.39 | 34 | 676.1 | 0.47 | 0.505 | 79405 | 95640
2024-11-16 | 321 | 1241 | ALW | rm lf | 37.67 | 34.5 | 675.9 | 0.47 | 0.505 | 79637 | 95902
2024-11-17 | 322 | 1435 | ALW | rm lf | 37.56 | 34.2 | 681.7 | 0.473 | 0.508 | 79321 | 95474
2024-11-18 | 323 | 1552 | ALW | rm lf | 41.64 | 34.2 | 685.9 | 0.475 | 0.509 | 79533 | 95750
2024-11-19 | 324 | 1300 | ALW | rm lf | 44.56 | 34.4 | 689.3 | 0.476 | 0.51 | 79546 | 95734 | Weekly calibration 1319-1405. Z1: -0.3. L1: 170.6. L2:42.6.
2024-11-20 | 325 | 1627 | ALW | rm lf | 43.05 | 34.7 | 681 | 0.472 | 0.507 | 79620 | 95802
2024-11-21 | 326 | 1613 | ALW | rm lf | 42.23 | 34.5 | 678.6 | 0.472 | 0.507 | 79589 | 95774
2024-11-22 | 327 | 1322 | ALW | rm lf | 36.91 | 34.7 | 677.8 | 0.471 | 0.505 | 79651 | 95835
2024-11-23 | 328 | 1233 | ALW | rm lf | 37.6 | 34.6 | 673.8 | 0.469 | 0.504 | 79363 | 95483
2024-11-24 | 329 | 1402 | ALW | rm lf | 47.71 | 34.5 | 670.2 | 0.467 | 0.502 | 79575 | 95674
2024-11-25 | 330 | - | ALW | - | - | - | - | - | - | - | - | Con 1- did not go to AWO
2024-11-26 | 331 | 1810 | ALW | rm lf | 38.62 | 34.1 | 659 | 0.463 | 0.498 | 79277 | 95321 | Surface Ozone calibration 1626-1713Z. Z: 1.1, L1: 170.9, L2:42.8.
2024-11-27 | 332 | - | ALW | - | - | - | - | - | - | - | - | Con 1- did not go to AWO
2024-11-28 | 333 | - | ALW | - | - | - | - | - | - | - | - | Holiday- did not go to AWO
2024-11-29 | 334 | 1535 | ALW | rm lf | 41.14 | 34.3 | 669.4 | 0.467 | 0.502 | 79352 | 95308
2024-11-30 | 335 | 1416 | ALW | rm lf | 38.68 | 34.3 | 669.9 | 0.467 | 0.502 | 79355 | 95326
2024-12-01 | 336 | - | ALW | - | - | - | - | - | - | - | - | Con 1- did not go to AWO
2024-12-02 | 337 | 1241 | ALW | rm lf | 42.45 | 31.8 | 664.1 | 0.465 | 0.501 | 78606 | 94479 | Showed up to AWO with orange cap on the ground. Reattached 1225.
2024-12-03 | 338 | 1430 | ALW | rm lf | 39.09 | 35.5 | 662.1 | 0.463 | 0.499 | 79681 | 95626 | Surface Ozone Calibration 1338-1427Z. Z: 0.2, L1: 172.3, L2: 43.4.
2024-12-04 | 339 | - | ALW | - | - | - | - | - | - | - | - | ICESat- did not go to AWO
2024-12-05 | 340 | - | ALW | - | - | - | - | - | - | - | - | Sick- did not go to AWO
2024-12-06 | 341 | 1559 | ALW | rm lf | 36.2 | 35 | 667.7 | 0.466 | 0.502 | 79469 | 95383
2024-12-07 | 342 | 1245 | ALW | rm lf | 36.95 | 33.9 | 671.5 | 0.468 | 0.503 | 79083 | 94902
2024-12-08 | 343 | 1539 | ALW | rm lf | 40.08 | 34.5 | 667.5 | 0.467 | 0.502 | 79264 | 95115
2024-12-09 | 344 | 1403 | ALW | rm lf | 39.03 | 34.5 | 668.5 | 0.467 | 0.503 | 79332 | 95165
2024-12-10 | 345 | 1621 | ALW | rm lf | 38.03 | 34.7 | 678.6 | 0.471 | 0.506 | 79280 | 95086 | Surface Ozone Calibration 1435-1521Z. Z: -0.2, L1: 170.3, L2: 42.6.
2024-12-11 | 346 | 1552 | ALW | - | 40.73 | 34.6 | 677.7 | 0.472 | 0.507 | 79276 | 95085 | Orange cap blown to the ground about 1530Z. Too windy to go to roof.
2024-12-12 | 347 | 1716 | ALW | rm lf | 40.93 | 35 | 673.9 | 0.469 | 0.504 | 79403 | 95219 | Orange cap replaced @ 1612Z.
2024-12-13 | 348 | 1257 | ALW | - | 40.95 | 33.9 | 664.9 | 0.466 | 0.501 | 78956 | 94664 | Too windy to go to the roof
2024-12-14 | 349 | 1307 | ALW | rm lf | 38.4 | 34.6 | 653.9 | 0.46 | 0.497 | 79054 | 94760
2024-12-15 | 350 | 1408 | ALW | rm lf | 31.24 | 34.7 | 655.8 | 0.461 | 0.497 | 79264 | 94990
2024-12-16 | 351 | 1529 | ALW | rm lf | 35.23 | 33.6 | 659.7 | 0.464 | 0.499 | 78780 | 94412
2024-12-17 | 352 | 1405 | ALW | rm lf | 40.77 | 33.8 | 660.6 | 0.464 | 0.5 | 78903 | 94552 | Surface Ozone Calibration 1435-1520Z. Z: -0.2, L1: 167.7, L2: 42.6.
2024-12-18 | 353 | - | ALW | - | - | - | - | - | - | - | - | Con 1- did not go to AWO
2024-12-19 | 354 | 1616 | ALW | rm lf | 40.14 | 34.3 | 653.4 | 0.461 | 0.496 | 79091 | 94718
2024-12-20 | 355 | 1505 | ALW | rm lf | 39.87 | 33.8 | 651 | 0.46 | 0.495 | 78867 | 94466
2024-12-21 | 356 | - | ALW | - | - | - | - | - | - | - | - | Con 1- did not go to AWO
2024-12-22 | 357 | 1532 | ALW | rm lf | 41.16 | 33.9 | 659 | 0.464 | 0.499 | 78797 | 94323
2024-12-23 | 358 | 1718 | ALW | rm lf | 41.36 | 34.2 | 646.6 | 0.458 | 0.492 | 78901 | 94437 | Surface Ozone Calibration 1547-1633Z. Z: 0.4, L1: 168.8, L2: 42.5.
2024-12-24 | 359 | 1327 | ALW | rm lf | 40.43 | 3.6 | 664.1 | 0.456 | 0.492 | 78831 | 94335
2024-12-25 | 360 | 1422 | ALW | rm lf | 41.3 | 33.4 | 641.8 | 0.456 | 0.491 | 78590 | 94079
2024-12-26 | 361 | 1651 | ALW | rm lf | 35.99 | 28.9 | 654.6 | 0.463 | 0.497 | 77444 | 92751
2024-12-27 | 362 | 1428 | ALW | rm lf | 31.7 | 29.8 | 666.5 | 0.467 | 0.502 | 77660 | 92980
2024-12-28 | 363 | 1156 | ALW | rm lf | 34.22 | 32.3 | 671.1 | 0.469 | 0.504 | 78303 | 93673
2024-12-29 | 364 | 1400 | ALW | rm lf | 39.63 | 29.2 | 676.6 | 0.472 | 0.506 | 77417 | 92679
2024-12-30 | 365 | 1512 | ALW | rm lf | 45 | 33.5 | 680.7 | 0.474 | 0.508 | 78611 | 94013
2024-12-31 | 366 | 1431 | ALW | rm lf | 44.18 | 34 | 681.1 | 0.473 | 0.508 | 78682 | 94039 | Surface Ozone calibration 1258-1345Z. Z: 0.1, L1: 168.4, L2: 44.2.
Note on F1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: G81 Thermo Box: Ozone (ppb)
Note on G1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Sample (C)
Note on H1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Sample (hPa) Reply: If pressure off from atmospheric pressure, inlet is plugged
Note on I1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Cell A (lpm)
Note on J1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Cell B (lpm)
Note on K1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Cell A (Hz)
Note on L1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Cell B (Hz)