## MET
Date | Julian Day | Time (UTC) | Observer's Initials | 2m T | 2m TD | 9m SONIC | 9m RM Young | Check Enclosure Temp (~12 C) | Check date and time | Comments (Climbing, removing instruments, etc)
2025-01-01 | 1 | 1621 | ALW | rm lf | rm lf | lf | lf | * | *
2025-01-02 | 2 | 1511 | ALW | rm lf | rm lf | lf | lf | * | *
2025-01-03 | 3 | - | ALW | - | - | - | - | * | * | ICEsat- did not go to AWO
2025-01-04 | 4 | 1358 | ALW | rm lf | rm lf | lf | lf | * | *
2025-01-05 | 5 | 1510 | ALW | rm lf | rm lf | lf | lf | * | *
2025-01-06 | 6 | 1745 | ALW | rm lf | rm lf | lf | lf | * | *
2025-01-07 | 7 | 1814 | ALW | rm lf | rm lf | lf | lf | * | *
2025-01-08 | 8 | 1556 | ALW | clr | clr | lf | lf | * | *
2025-01-09 | 9 | 1559 | ALW | clr | clr | lf | lf | * | *
2025-01-10 | 10 | - | ALW | - | - | - | - | -2 | * | Con 1- did not go to AWO
2025-01-11 | 11 | 1335 | ALW | rm lf | rm lf | lf | lf | * | *
2025-01-12 | 12 | 1319 | ALW | rm lf | rm lf | lf | lf |  |  | GML page down
2025-01-13 | 13 | 1640 | ALW | rm lf | rm lf | lf | lf | * | *
2025-01-14 | 14 | 1429 | ALW | clr | rm lf | lf | lf | * | *
2025-01-15 | 15 | 1830 | ALW | rm lf | rm lf | lf | lf | * | *
2025-01-16 | 16 | - | ALW | - | - | - | - | * | -2 | Con 1- did not go to AWO
2025-01-17 | 17 | - | ALW | - | - | - | - | * | -2.5 | Con 1- did not go to AWO
2025-01-18 | 18 | 1347 | ALW | rm lf | rm lf | lf | lf | * | *
2025-01-19 | 19 | 1542 | ALW | rm lf | rm lf | lf | lf | * | *
2025-01-20 | 20 | 1557 | ALW | rm lf | rm lf | lf | lf | * | *
2025-01-21 | 21 | 1730 | ALW | rm lf | rm lf | lf | lf | * | *
2025-01-22 | 22 | 1603 | ALW | rm lf | rm lf | lf | lf | * | *
2025-01-23 | 23 | 1617 | ALW | rm lf | rm lf | lf | lf | * | * | Tops taken off and cleaned of snow @ 1641
2025-01-24 | 24 | 1454 | ALW | clr | clr | lf | lf | * | *
2025-01-25 | 25 | 1252 | ALW | rm lf | rm lf | lf | lf | * | *
2025-01-26 | 26 | 1505 | ALW | rm lf | rm lf | lf | lf | * | *
2025-01-27 | 27 | 1638 | ALW | rm lf | rm lf | lf | lf | * | *
2025-01-28 | 28 | 1720 | ALW | rm lf | rm lf | lf | lf | * | *
2025-01-29 | 29 | 1742 | ALW | rm lf | rm lf | lf | lf | * | *
2025-01-30 | 30 | 1319 | ALW | rm lf | rm lf | lf | lf | * | *
2025-01-31 | 31 | 1348 | ALW | rm lf | rm lf | lf | lf | * | *
2025-02-01 | 32 | - | ALW | - | - | - | - | * | *
2025-02-02 | 33 | 1538 | ALW | clr | clr | lf | lf | * | *
2025-02-03 | 34 | 1611 | ALW | rm lf | rm lf | lf | lf | * | *
2025-02-04 | 35 | - | ALW | - | - | - | - | * | * | ICEsat- did not go to AWO
2025-02-05 | 36 | 1557 | ALW | rm lf | rm lf | lf | lf | - | - | GML page down
2025-02-06 | 37 | 1352 | ALW | rm lf | rm lf | lf | lf | * | *
2025-02-07 | 38 | - | ALW | - | - | - | - | -1 | * | Con 1- did not go to AWO
2025-02-08 | 39 | 1818 | ALW | rm lf | rm lf | lf | lf | * | *
2025-02-09 | 40 | 1400 | ALW | - | - | - | - | * | *
2025-02-10 | 41 | - | ALW | - | - | - | - | * | * | ICEsat- did not go to AWO
2025-02-11 | 42 | 1620 | ALW | rm lf | rm fl | lf | lf | * | *
2025-02-12 | 43 | 1125 | ALW | rm lf | rm lf | lf | lf | * | *
2025-02-13 | 44 | 1836 | ALW | rm lf | rm lf | lf | lf | * | *
2025-02-14 | 45 | 1530 | ALW | rm lf | rm lf | lf | hf | * | *
2025-02-15 | 46 | - | ALW | - | - | - | - | * | *
2025-02-16 | 47 | 1634 | ALW | rm lf | rm lf | lf | lf | - | - | GML page down
2025-02-17 | 48 | 1413 | AR | rm lf | rm lf | lf | hf | * | * | Planned power outages for turnover 1201-1204Z and 1822-1824Z. Instruments supported by UPS.
2025-02-18 | 49 | 1401 | AR | rm lf | rm lf | lf | hf | * | * | RM Young frozen in place and not spinning
2025-02-19 | 50 | 1835 | AR | rm lf | rm lf | lf | hf | * | * | RM Young frozen in place and not spinning
2025-02-20 | 51 | 1415 | AR | rm lf | rm lf | lf | hf | * | * | Power outage today from 0046-1147Z, enclosure temperature droped to -30.75C overnight before power was restored
2025-02-21 | 52 | 1416 | AR | rm lf | rm lf | lf | lf | * | * | Tower climb for cleaning at 1609-1836Z
2025-02-22 | 53 | 1207 | ALW | rm lf | rm lf | lf | lf | * | *
2025-02-23 | 54 | 1327 | AR | rm lf | rm lf | lf | lf | * | *
2025-02-24 | 55 | 1705 | AR | rm lf | rm lf | lf | lf | * | *
2025-02-25 | 56 | 1913 | AR | rm lf | rm lf | lf | lf | * | *
2025-02-26 | 57 | 1347 | AR | rm lf | rm lf | lf | lf | * | *
2025-02-27 | 58 | 1313 | AR | rm lf | rm lf | lf | lf | * | *
2025-02-28 | 59 | 1144 | AR | - | - | - | - | * | * | Con 1- did not go to AWO
2025-03-01 | 60 | 1310 | AR | rm lf | rm lf | lf | lf | * | *
2025-03-02 | 61 | 1559 | AR | rm lf | rm lf | lf | lf | * | *
2025-03-03 | 62 | 1542 | AR | rm lf | rm lf | lf | lf | * | *
2025-03-04 | 63 | 1453 | AR | rm lf | rm lf | lf | lf | * | *
2025-03-05 | 64 | 1751 | AR | - | - | - | - | * | * | ICEsat- did not go to AWO
2025-03-06 | 65 | 1348 | AR | rm lf | rm lf | lf | lf | * | *
2025-03-07 | 66 | 1732 | AR | rm lf | rm lf | lf | lf | * | *
2025-03-08 | 67 | 1155 | AR | rm lf | rm lf | lf | lf | * | *
2025-03-09 | 68 | 1752 | AR | - | - | - | - | * | * | Con 1- did not go to AWO
2025-03-10 | 69 | 1253 | AR | rm lf | rm lf | lf | lf | * | *
2025-03-11 | 70 | 1310 | AR | rm lf | rm lf | lf | lf | * | * | Power outage today at 2150Z (less than 1 min). Instruments supported by UPS.
2025-03-12 | 71 | 1657 | AR | rm lf | rm lf | lf | lf | * | *
2025-03-13 | 72 | 1334 | AR | rm lf | rm lf | lf | lf | * | *
2025-03-14 | 73 | 1225 | AR | rm lf | rm lf | lf | lf | * | *
2025-03-15 | 74 | 1547 | AR | rm lf | rm lf | lf | lf | * | *
2025-03-16 | 75 | 1356 | AR | rm lf | rm lf | lf | lf | * | *
2025-03-17 | 76 | 1600 | AR | rm lf | rm lf | lf | lf | * | *
2025-03-18 | 77 | 1208 | AR | rm lf | rm lf | lf | lf | * | *
2025-03-19 | 78 | 1746 | AR | rm lf | rm lf | lf | lf | * | *
2025-03-20 | 79 | 1139 | AR | rm lf | rm lf | lf | lf | * | *
2025-03-21 | 80 | 1555 | AR | rm lf | rm lf | lf | lf | * | *
2025-03-22 | 81 | 1106 | AR | rm lf | rm lf | lf | lf | * | *
2025-03-23 | 82 | 1324 | AR | rm lf | rm lf | lf | lf | * | *
2025-03-24 | 83 | 1642 | AR | rm lf | rm lf | lf | lf | * | *
2025-03-25 | 84 | 1458 | AR | rm lf | rm lf | lf | lf | * | *
2025-03-26 | 85 | 1721 | AR | rm lf | rm lf | lf | hf | * | *
2025-03-27 | 86 | 1659 | AR | rm lf | rm lf | lf | hf | * | *
2025-03-28 | 87 | 1112 | AR | rm lf | rm lf | lf | hf | * | *
2025-03-29 | 88 | 1307 | AR | rm lf | rm lf | lf | hf | * | *
2025-03-30 | 89 | 1247 | AR | rm lf | rm lf | lf | lf | * | *
2025-03-31 | 90 | 1524 | AR | rm lf | rm lf | lf | lf | * | *
2025-04-01 | 91 | 1107 | AR | rm lf | rm lf | lf | lf | * | *
2025-04-02 | 92 | - | AR | - | - | - | - | * | * | ICEsat- did not go to AWO
2025-04-03 | 93 | 1540 | AR | rm lf | rm lf | lf | lf | * | *
2025-04-04 | 94 | 1442 | AR | rm lf | rm lf | lf | lf | * | *
2025-04-05 | 95 | 1417 | AR | rm lf | rm lf | lf | lf | * | *
2025-04-06 | 96 | 1240 | AR | rm lf | rm lf | lf | lf | * | *
2025-04-07 | 97 | 950 | AR | rm lf | rm lf | lf | lf | * | *
2025-04-08 | 98 | 1221 | AR | rm lf | rm lf | lf | lf | * | *
2025-04-09 | 99 | 1012 | AR | rm lf | rm lf | lf | lf | - | - | GML page down
2025-04-10 | 100 | 1452 | AR | rm lf | rm lf | lf | lf | - | - | GML page down
2025-04-11 | 101 | 1059 | AR | rm lf | rm lf | lf | lf | - | - | GML page down
2025-04-12 | 102 | 1235 | AR | rm lf | rm lf | lf | lf | - | - | GML page down
2025-04-13 | 103 | 1651 | AR | rm lf | rm lf | lf | lf | - | - | GML page down
2025-04-14 | 104 | 1557 | AR | rm lf | rm lf | lf | lf | * | *
2025-04-15 | 105 | 1018 | AR | rm lf | rm lf | lf | lf | * | *
2025-04-16 | 106 | 1442 | AR | rm lf | rm lf | lf | lf | * | *
2025-04-17 | 107 | 1002 | AR | rm lf | rm lf | lf | lf | * | * | MET tower power outage from 1832-1939Z, Aspirated humidity disconnected from 1640-1743Z and again at 1835Z
2025-04-18 | 108 | 1219 | AR | clr | - | clr | clr | * | *
2025-04-19 | 109 | 1119 | AR | clr | - | clr | clr | * | *
2025-04-20 | 110 | 1318 | AR | rm lf | - | lf | lf | * | *
2025-04-21 | 111 | 1441 | AR | rm lf | clr | lf | lf | * | * | Reconnected aspirated humidity at 1639Z
2025-04-22 | 112 | 1427 | AR | rm lf | rm lf | lf | lf | * | *
2025-04-23 | 113 | 943 | AR | rm lf | rm lf | lf | lf | * | * | Tower climb for ICECAPS HMP maintenance from 1623-1647Z
2025-04-24 | 114 | 1707 | AR | rm lf | rm lf | lf | lf | * | *
2025-04-25 | 115 | 1109 | AR | rm lf | rm lf | lf | lf | * | *
2025-04-26 | 116 | 1458 | AR | rm lf | rm lf | lf | lf | * | *
2025-04-27 | 117 | 1244 | AR | rm lf | rm lf | lf | lf | * | *
2025-04-28 | 118 | 1639 | AR | rm lf | rm lf | lf | lf | * | *
2025-04-29 | 119 | 1145 | AR | rm lf | rm lf | lf | lf | * | *
2025-04-30 | 120 | 1432 | AR | rm lf | rm lf | clr | clr | * | *
2025-05-01 | 121 | 1146 | AR | rm lf | rm lf | clr | clr | * | *
2025-05-02 | 122 | 1216 | AR | rm lf | rm lf | lf | lf | * | *
2025-05-03 | 123 | 1103 | AR | rm lf | rm lf | clr | clr | * | *
2025-05-04 | 124 | 1938 | AR | rm lf | rm lf | clr | clr | * | *
2025-05-05 | 125 | 1209 | AR | clr | rm lf | clr | clr | * | *
2025-05-06 | 126 | 1433 | AR | clr | clr | clr | clr | * | *
2025-05-07 | 127 | 1502 | AR | rm lf | rm lf | clr | clr | * | *
2025-05-08 | 128 | 1247 | AR | rm lf | rm lf | lf | lf | * | * | Turned off met enclosure heater at 1532Z
2025-05-09 | 129 | 1159 | AR | rm lf | rm lf | lf | lf | * | * | Enclosure temp below freezing (dropped to -9C overnight, curently ~-2C), Heater turned back on at 1716Z
2025-05-10 | 130 | 1227 | AR | rm lf | rm lf | lf | lf | * | * | Planned station-wide power outage from 1008-1014Z. Instruments supported by UPS.
2025-05-11 | 131 | 1553 | AR | rm lf | rm lf | lf | lf | * | * | Unplanned station-wide power outage from 1455-1524Z. Instruments supported by UPS.
2025-05-12 | 132 | 1134 | AR | rm lf | rm lf | lf | lf | * | *
2025-05-13 | 133 | 1809 | AR | - | - | - | - | * | * | Con 1- did not go to AWO. Unplanned station power outage from 2031-2040Z. Controlled station power outage from 2105-2110Z. Instruments supported by UPS.
2025-05-14 | 134 | 953 | AR | rm lf | rm lf | lf | lf | * | *
2025-05-15 | 135 | 1510 | AR | clr | clr | clr | clr | * | *
2025-05-16 | 136 | 1021 | AR | clr | clr | clr | clr | * | *
2025-05-17 | 137 | 935 | AR | clr | clr | clr | clr | * | * | Planned station power outage from 1015-1019Z. Instruments supported by UPS.
2025-05-18 | 138 | 1203 | AR | clr | clr | clr | clr | * | *
2025-05-19 | 139 | 1605 | AR | clr | clr | clr | clr | * | *
2025-05-20 | 140 | 1033 | AR | rm lf | rm lf | lf | lf | * | *
2025-05-21 | 141 | 1424 | AR | clr | clr | clr | clr | * | *
2025-05-22 | 142 | 958 | AR | rm lf | rm lf | clr | clr | * | *
2025-05-23 | 143 | 1223 | AR | rm lf | rm lf | clr | clr | * | *
2025-05-24 | 144 | 1033 | AR | clr | clr | clr | clr | * | * | Unplanned station-wide power outage from 1806-1808Z. Instruments supported by UPS.
2025-05-25 | 145 | 1232 | AR | rm lf | rm lf | clr | clr | * | *
2025-05-26 | 146 | 1011 | AR | rm lf | rm lf | clr | clr | * | *
2025-05-27 | 147 | 1430 | AR | clr | clr | clr | clr | * | *
2025-05-28 | 148 | 1148 | AR | rm lf | rm lf | lf | lf | * | *
2025-05-29 | 149 | 1506 | AR | rm lf | rm lf | clr | clr | * | *
2025-05-30 | 150 | 1155 | AR | clr | rm lf | clr | clr | * | *
2025-05-31 | 151 | 1535 | AR | clr | clr | clr | clr | * | *
2025-06-01 | 152 | 1749 | AR | clr | clr | clr | clr | * | *
2025-06-02 | 153 | 1547 | AR | rm lf | rm lf | clr | clr | * | *
2025-06-03 | 154 | 1107 | AR | clr | clr | clr | clr | * | *
2025-06-04 | 155 | - | AR | - | - | - | - | * | * | ICEsat- did not go to AWO
2025-06-05 | 156 | 1014 | AR | rm lf | rm lf | clr | clr | * | *
2025-06-06 | 157 | 1534 | AR | clr | clr | clr | clr | * | *
2025-06-07 | 158 | 1105 | AR | rm lf | rm lf | lf | lf | * | *
2025-06-08 | 159 | 1347 | AR | clr | clr | clr | clr | * | *
2025-06-09 | 160 | 1113 | AR | clr | clr | clr | clr | * | * | Planned power outage 1001-1005Z. Instruments supported by UPS.
2025-06-10 | 161 | 1732 | AR | clr | clr | clr | clr | * | * | Planned power outage 1704-1707Z. Instruments supported by UPS.
2025-06-11 | 162 | 1512 | AR | rm lf | rm lf | clr | clr | * | *
2025-06-12 | 163 | 1019 | PC | clr | clr | clr | clr | * | *
2025-06-13 | 164 | 1057 | PC | clr | clr | clr | clr | * | * | Planned power outage 1037-1040Z. Second planned power outage 1412-1416Z. Instruments supported by UPS.
2025-06-14 | 165 | 1618 | PC | clr | clr | clr | clr | * | *
2025-06-15 | 166 | 1146 | PC | clr | clr | clr | clr | * | *
2025-06-16 | 167 | 1657 | PC | clr | clr | clr | clr | * | *
2025-06-17 | 168 | 1158 | PC | clr | clr | clr | clr | * | * | 1602-1630Z: Riggers taking photos of base of tower, measuring flux enclosure, measuring face of tower, and measurements in anticipation of the rail system installation.
2025-06-18 | 169 | 1239 | PC | clr | clr | clr | clr | * | * | 939-1237Z: Annual tower inspection: digging up anchor points to expose supporting dead end grips to check soundness. Clean snow sector preserved as much as possible, photos to come via email. 1425-1555Z: Adjusting bottom hardware on innermost guys, finishing layout for new deadmen anchors for apex guys.
2025-06-19 | 170 | 1154 | PC | rm lf | rm lf | lf | lf | * | * | 1211 to 1217Z Riggers operating around tower to drop off guide materials for anchor points.
2025-06-20 | 171 | 1508 | PC | clr | clr | clr | clr | * | * | Unplanned power loss 1059-1417Z to tower during rigger maintenance. All instruments operating normally.
2025-06-21 | 172 | 1653 | PC | clr | clr | clr | clr | * | * | 0937-1250Z, 1550-1720Z: Riggers climbing tower to measure guy lines; Riggers working around base of tower on anchor holes.
2025-06-22 | 173 | 1548 | PC | rm lf | rm lf | lf | lf | * | *
2025-06-23 | 174 | 1141 | PC | clr | clr | clr | clr | * | * | 1023-1209Z: Riggers climbing tower to secure new guy lines; 1423-1739Z: Riggers around base of tower to fill in anchor holes.
2025-06-24 | 175 | 915 | PC | rm lf | rm lf | lf | lf | * | * | Plane flew a big low contrail directly over station 0905. Riggers at base of tower 1035-1115Z to plan rail system.
2025-06-25 | 176 | 1137 | PC | rm lf | rm lf | lf | lf | * | * | Low winds equipment operations: 11:42-12:07Z- D6, 953, skidsteer, snowmo, welder (machines operating 11:42-11:49, welder missed radio call and operated 11:42-1207Z).
2025-06-26 | 177 | 950 | PC | rm lf | rm lf | lf | lf | * | *
2025-06-27 | 178 | 1429 | PC | clr | clr | clr | clr | * | *
2025-06-28 | 179 | 929 | PC | clr | clr | clr | clr | * | * | 1005-1006Z: Planned power outage for mechanics training. No data impacted for instruments supported by UPS. 1010-1012Z: second planned power outage to switch back to main generator. 1635-1641Z: Pepper at tower with small piece of electrical tape to repair small split in LiCOR cable casing.
2025-06-29 | 180 | 1531 | PC | clr | clr | clr | clr | * | *
2025-06-30 | 181 | 1404 | PC | clr | clr | clr | clr | * | *
2025-07-01 | 182 | 958 | PC | clr | clr | clr | clr | * | * | Planned power outage due to generator issues last night, power swap from 1133-1135Z, again from 1224-1225Z. And again 1602-1604Z. Instruments supported by UPS.
2025-07-02 | 183 | 1021 | PC | rm lf | rm lf | lf | lf | * | *
2025-07-03 | 184 | 1144 | PC | clr | clr | clr | clr | * | * | Several planned power outages to specific instruments during final cable tidy: 2m T Temp and Humidity powered down 0947-0955Z, 9m SONIC and RM Young intended to be powered down 1004-1012Z, however confusion around the ports caused the SONIC to be powered down from 1004-1254Z. Riggers at tower from 0924-1205Z, and again from 1240-1724Z.
2025-07-04 | 185 | 1101 | PC | clr | clr | clr | clr | * | * | Riggers at base of tower to finalize anchor points.
2025-07-05 | 186 | 1638 | PC | clr | clr | clr | clr | * | * | Riggers at base of tower to check plumb. Northwinds operations: 1716-1735Z D6 Snow machine 953 953
2025-07-06 | 187 | 956 | PC | clr | clr | clr | clr | * | *
2025-07-07 | 188 | 1130 | PC | clr | clr | clr | clr | * | * | 1602-1603Z: power swap back to main station generator. Instruments supported by UPS.
2025-07-08 | 189 | 1017 | PC | clr | clr | clr | clr | * | *
2025-07-09 | 190 | 1057 | PC | rm lf | rm lf | lf | lf | * | * | 1315-2008Z: OneWeb Summit internet experienced a loss of connectivity affecting the NOAA Met Tower instruments ability to transmit data. Back up and running normally again after 2008Z. LOW WINDS OPERATIONS 1015-1018: D6 HEAT CANNON, 5K
2025-07-10 | 191 | 1014 | PC | clr | clr | clr | clr | * | *
2025-07-11 | 192 | 1110 | PC | rm lf | rm lf | lf | lf | * | *
2025-07-12 | 193 | 1028 | PC | rm lf | rm lf | lf | lf | * | *
2025-07-13 | 194 | 1244 | PC | rm ice | rm ice | light ice | light ice | * | *
2025-07-14 | 195 | 1044 | PC | clr | clr | clr | clr | * | *
2025-07-15 | 196 | 1043 | PC | rm ice | rm ice | light ice | light ice | * | * | Northwind ops: x2 variances (welding, medical to heat isolation bay during day) from 0900-1400
2025-07-16 | 197 | 1047 | PC | clr | clr | clr | clr | * | *
2025-07-17 | 198 | 1148 | PC | clr | clr | clr | clr | * | * | 1633Z: planned loss of power for 5 seconds as per electrician's request to troubleshoot an ICECAPS instrument. Another planned loss of power for GFCI replacement from 1609-1742Z.
2025-07-18 | 199 | 1009 | PC | clr | clr | clr | clr | * | *
2025-07-19 | 200 | 1013 | PC | rm ice | light ice | light ice | light ice | * | *
2025-07-20 | 201 | 1122 | PC | clr | clr | clr | clr | * | *
2025-07-21 | 202 | 1040 | PC | clr | clr | clr | clr | * | *
2025-07-22 | 203 | 1041 | PC | clr | clr | clr | clr | * | *
2025-07-23 | 204 | 1020 | PC | clr | clr | clr | clr | * | * | Low winds from 0943-1226Z, 1803-1908Z, North winds 1206-1217Z, 1851-1942Z, 2052-2150Z
2025-07-24 | 205 | 1030 | PC | clr | clr | clr | clr | * | * | Low winds from 0823-1429Z, 1643-1643Z. North winds from 0051-0402Z, 0932-1230Z
2025-07-25 | 206 | 1018 | PC | clr | clr | clr | clr | * | *
2025-07-26 | 207 | 1211 | PC | clr | clr | clr | clr | * | * | Low winds from 0223-0224Z. North winds from 2314-0054Z
2025-07-27 | 208 | 1902 | PC | clr | clr | clr | clr | * | * | Low winds from 0520-0520Z, 0722-0741Z, 2147-2357Z. North winds from 0239-0520Z, 2202-2340Z
2025-07-28 | 209 | 1042 | PC | clr | clr | clr | clr | * | * | Low winds 02:06Z to 02:22Z, north winds 00:47Z to 00:47Z
2025-07-29 | 210 | 1042 | PC | clr | clr | clr | clr | * | * | Low winds 09:17Z to 09:56Z
2025-07-30 | 211 | 1040 | PC | clr | clr | clr | clr | * | * | Low winds 02:10Z to 02:11Z, north winds 09:19Z to 09:25Z 11:47Z to 13:28Z 19:54Z to 21:04Z
2025-07-31 | 212 | 1520 | PC | clr | clr | clr | clr | * | *
2025-08-01 | 213 | 1206 | PC | clr | clr | clr | clr | * | *
2025-08-02 | 214 | 1456 | PC | clr | clr | clr | clr | * | *
2025-08-03 | 215 | 1246 | PC | clr | clr | clr | clr | * | *
2025-08-04 | 216 | 1502 | PC | clr | clr | clr | clr | * | *
2025-08-05 | 217 | 1009 | PC | clr | clr | clr | clr | * | *
2025-08-06 | 218 | 1127 | PC | clr | clr | clr | clr | * | *
2025-08-07 | 219 | 1134 | PC | clr | clr | clr | clr | * | *
2025-08-08 | 220 | 1114 | PC | rm lf | rm lf | lf | lf | * | *
2025-08-09 | 221 | 1232 | PC | clr | clr | clr | clr | * | *
2025-08-10 | 222 | 1401 | PC | clr | clr | clr | clr | * | *
2025-08-11 | 223 | 1113 | PC | clr | clr | clr | clr | * | *
2025-08-12 | 224 | 1506 | PC | clr | clr | clr | clr | * | *
2025-08-13 | 225 | 1154 | PC | rm lf | lf | lf | lf | * | *
2025-08-14 | 226 | 1027 | PC | rm lf | lf | lf | lf | * | *
2025-08-15 | 227 | 913 | PC | clr | clr | clr | clr | * | *
2025-08-16 | 228 | 1044 | PC | clr | clr | clr | clr | * | *
2025-08-17 | 229 | 1419 | PC | clr | clr | clr | clr | * | *
2025-08-18 | 230 | 1003 | PC | clr | clr | clr | clr | * | *
2025-08-19 | 231 | 1130 | PC | rm lf | lf | lf | lf | * | *
2025-08-20 | 232 | 1244 | PC | rm lf | lf | lf | lf | * | *
2025-08-21 | 233 | 1046 | PC | rm lf | lf | lf | lf | * | *
2025-08-22 | 234 | 1022 | PC | clr | clr | clr | clr | * | *
2025-08-23 | 235 | 1634 | PC | clr | clr | clr | clr | * | *
2025-08-24 | 236 | 1017 | PC | clr | clr | clr | clr | * | *
2025-08-25 | 237 | 1140 | PC | clr | clr | clr | clr | * | * | Planned power swap, instruments supported by ups 1002-1004Z
2025-08-26 | 238 | 1032 | PC | rm lf | lf | lf | lf | * | * | Planned power swap, instruments supported by ups 1101-1102Z
2025-08-27 | 239 | 959 | PC | clr | clr | clr | clr | * | *
2025-08-28 | 240 | 929 | PC | clr | clr | clr | clr | * | *
2025-08-29 | 241 | 1144 | PC | rm lf | lf | lf | lf | * | *
2025-08-30 | 242 | 1154 | PC | clr | clr | clr | clr | * | *
2025-08-31 | 243 | 1317 | PC | clr | clr | clr | clr | * | *
2025-09-01 | 244 | 1108 | PC | RM LF | RM LF | LF | LF | * | * | Made longer stick to clean instruments at 2m
2025-09-02 | 245 | 1104 | PC | RM LF | RM LF | LF | LF | * | *
2025-09-03 | 246 | 1026 | PC | CLR | CLR | CLR | CLR | * | *
2025-09-04 | 247 | 1055 | PC | CLR | CLR | CLR | CLR | * | *
2025-09-05 | 248 | 1229 | PC | CLR | CLR | CLR | CLR | * | *
2025-09-06 | 249 | 1431 | PC | RM LF | RM LF | LF | LF | * | *
2025-09-07 | 250 | 1441 | PC | RM LF | RM LF | LF | LF | * | *
2025-09-08 | 251 | 1208 | PC | CLR | CLR | CLR | CLR | * | *
2025-09-09 | 252 | 1044 | PC | CLR | CLR | CLR | CLR | * | *
2025-09-10 | 253 | 1456 | PC | CLR | CLR | CLR | CLR | * | *
2025-09-11 | 254 | 1039 | PC | CLR | CLR | CLR | CLR | * | *
2025-09-12 | 255 | 1132 | PC | RM LF | RM LF | LF | LF | * | *
2025-09-13 | 256 | 1058 | PC | CLR | CLR | CLR | CLR | * | *
2025-09-14 | 257 | 1524 | PC | CLR | CLR | CLR | CLR | * | *
2025-09-15 | 258 | 1103 | PC | RM LF | RM LF | LF | LF | * | *
2025-09-16 | 259 | 1138 | PC | RM LF | RM LF | LF | LF | * | *
2025-09-17 | 260 | 1444 | PC | CLR | CLR | CLR | CLR | * | *
2025-09-18 | 261 | 958 | PC | RM LF | RM LF | LF | LF | * | *
2025-09-19 | 262 | 1122 | PC | RM LF | RM LF | LF | LF | * | *
2025-09-20 | 263 | 1122 | PC | RM LF | RM LF | LF | LF | * | *
2025-09-21 | 264 | 1137 | PC | CLR | CLR | CLR | CLR | * | * | 1311-1321Z Under tower to plug in NOAA enclosure heater for winter
2025-09-22 | 265 | 1223 | PC | RM LF | RM LF | LF | LF | * | *
2025-09-23 | 266 | 1132 | PC | RM LF | RM LF | LF | LF | * | *
2025-09-24 | 267 | 1144 | PC | RM LF | RM LF | LF | LF | * | *
2025-09-25 | 268 | 1126 | PC | RM LF | RM LF | LF | LF | * | *
2025-09-26 | 269 | 1142 | PC | RM LF | RM LF | LF | LF | * | *
2025-09-27 | 270 | 1218 | PC | RM LF | RM LF | LF | LF | * | *
2025-09-28 | 271 | 1455 | PC | RM LF | RM LF | LF | LF | * | *
2025-09-29 | 272 | 1301 | PC | RM LF | RM LF | LF | LF | * | *
2025-09-30 | 273 | 1424 | PC | CLR | CLR | CLR | CLR | * | *
2025-10-01 | 274 | 1446 | PC | CLR | CLR | CLR | CLR | * | *
2025-10-02 | 275 | 1205 | PC | RM LF | RM LF | LF | LF | * | *
2025-10-03 | 276 | 1124 | PC | RM LF | RM LF | LF | LF | * | *
2025-10-04 | 277 | 1153 | PC | RM LF | RM LF | LF | LF | * | *
2025-10-05 | 278 | 1623 | PC | RM LF | RM LF | LF | LF | * | *
2025-10-06 | 279 | 1225 | PC | RM LF | RM LF | LF | LF | * | *
2025-10-07 | 280 | 1146 | PC | RM LF | RM LF | LF | LF | * | *
2025-10-08 | 281 | 1155 | PC | - | - | - | - | * | * | CON-1 did not go to AWO
2025-10-09 | 282 | 1449 | PC | RM LF | RM LF | LF | LF | * | *
2025-10-10 | 283 | 1109 | PC | RM LF | RM LF | LF | LF | * | *
2025-10-11 | 284 | 1203 | PC | RM LF | RM LF | LF | LF | * | *
2025-10-12 | 285 | 1254 | GM, PC | CLR | CLR | CLR | CLR | * | *
2025-10-13 | 286 | 1519 | GM | RM LF | RM LF | LF | LF | * | *
2025-10-14 | 287 | 1343 | PC | CLR | CLR | CLR | CLR | * | *
2025-10-15 | 288 | 1615 | GM | RM LF | RM LF | CLR | CLR | * | *
2025-10-16 | 289 | 1019 | GM | CLR | CLR | CLR | CLR | * | *
2025-10-17 | 290 | 1116 | GM | RM LF | RM LF | CLR | CLR | * | *
2025-10-18 | 291 | 1244 | GM | rm lf | rm lf | clr | lf | * | *
2025-10-19 | 292 | 1336 | GM | rm lf | rm lf | clr | lf | * | *
2025-10-20 | 293 | 1251 | GM | rm lf | rm lf | clr | lf | * | *
2025-10-21 | 294 | 1333 | GM | rm lf | rm lf | clr | lf | * | *
2025-10-22 | 295 | 1715 | GM | rm lf | rm lf | clr | lf | * | *
2025-10-23 | 296 | 1233 | GM | clr | clr | clr | lf | * | *
2025-10-24 | 297 | 1641 | GM | rm lf | rm lf | clr | lf | * | *
2025-10-25 | 298 | 1039 | GM | clr | clr | clr | lf | * | *
2025-10-26 | 299 | 1307 | GM | rm lf | rm lf | clr | lf | * | *
2025-10-27 | 300 | 1630 | GM | rm lf | rm lf | clr | hf | * | *
2025-10-28 | 301 | 1523 | GM | rm lf | rm lf | clr | hf | * | * | enclosure dropped to ~2 C with colder outside temps of ~-45C last night
2025-10-29 | 302 | 1312 | GM | rm lf | rm lf | clr | hf | - | * | With colder outside temps, the enclosure is dropping below 12C
2025-10-30 | 303 | 1514 | GM | rm lf | rm lf | lf | hf | * | *
2025-10-31 | 304 | 1621 | GM | rm lf | rm lf | lf | hf | * | *
2025-11-01 | 305 | 1327 | GM | clr | clr | clr | hf | - | * | With colder outside temps, the enclosure is dropping below 12C
2025-11-02 | 306 | 1415 | GM | clr | clr | clr | hf | * | *
2025-11-03 | 307 | - | GM | - | - | - | - | * | * | No in person checks due to ICESat
2025-11-04 | 308 | 1222 | GM | rm lf | rm lf | clr | hf | * | *
2025-11-05 | 309 | 1413 | HB | rm lf | rm lb | clr | hf | * | *
2025-11-06 | 310 | 1344 | GM | rm lf | rm lf | clr | hf | * | *
2025-11-07 | 311 | 1529 | GM | rm lf | rm lf | rm lf | rm hf | * | * | Tower climb for general cleaning performed 1305Z-1454Z
2025-11-08 | 312 | 1223 | GM | rm lf | rm lf | clr | clr | * | *
2025-11-09 | 313 | 1415 | HB | rm lf | rm lf | lf | lf | * | *
2025-11-10 | 314 | 1315 | GM | rm lf | rm lf | clr | lf | * | *
2025-11-11 | 315 | 1810 | GM | - | - | - | - | - | - | Did not go out to AWO due to ICESat, GML webpage has not updated since 0345Z
2025-11-12 | 316 | 1605 | GM | rm lf | rm lf | clr | hf | - | - | GML webpage has not updated since 0345Z
2025-11-13 | 317 | 1245 | GM | clr | clr | clr | clr | * | * | GML page updating again
2025-11-14 | 318 | 1347 | GM | rm lf | rm lf | clr | lf | * | *
2025-11-15 | 319 | 1240 | GM | rm lf | rm lf | clr | clr | * | *
2025-11-16 | 320 | 1321 | GM | rm lf | rm lf | clr | clr | * | *
2025-11-17 | 321 | 1637 | GM | rm lf | rm lf | clr | lf | * | *
2025-11-18 | 322 | 1247 | GM | rm lf | rm lf | clr | hf | * | * | OneWeb outage from 0857Z-0955Z
2025-11-19 | 323 | 1330 | GM | rm lf | rm lf | clr | lf | * | *
2025-11-20 | 324 | 1336 | GM | rm hf | rm hf | clr | hf | * | *
2025-11-21 | 325 | 1513 | GM | rm hf | rm hf | clr | hf | * | *
2025-11-22 | 326 | 1211 | GM | rm lf | rm lf | clr | hf | * | *
2025-11-23 | 327 | 1257 | GM | rm lf | rm lf | clr | hf | * | *
2025-11-24 | 328 | 1427 | GM | rm lf | rm lf | clr | hf | * | *
2025-11-25 | 329 | 1302 | GM | rm lf | rm lf | clr | lf | * | *
2025-11-26 | 330 | 1709 | GM | rm lf | rm lf | clr | lf | * | * | Was very close to 2m T to tighten a loose bolt on enclosure. Unplugged and took down RH HMP-155 to swap in a new one from 1847Z-1914Z. Continued to work around tower base afterward, entire time for work at base was from 1843Z-1941Z.
2025-11-27 | 331 | [6 empty] | * | * | Did not go out to AWO
2025-11-28 | 332 | 1353 | GM | rm lf | rm lf | clr | lf | * | *
2025-11-29 | 333 | 1211 | GM | clr | clr | clr | lf | * | *
2025-11-30 | 334 | 1254 | GM | rm lf | rm lf | clr | lf | * | *
2025-12-01 | 335 | 1739 | GM | rm lf | rm lf | clr | lf | * | *
2025-12-02 | 336 | 1831 | GM | - | - | - | - | * | * | Did not go out to AWO due to ICESat
2025-12-03 | 337 | 1710 | GM | rm lf | rm lf | clr | lf | * | *
2025-12-04 | 338 | 1421 | GM | rm lf | rm lf | clr | hf | * | *
2025-12-05 | 339 | 1703 | GM | rm hf | rm hf | clr | rm hf | * | * | Performed tower climb for cleaning from 1208Z - 1340Z
2025-12-06 | 340 | 1249 | GM | rm lf | rm lf | clr | hf | * | *
2025-12-07 | 341 | 1344 | GM | rm lf | rm lf | clr | hf | * | *
2025-12-08 | 342 | 1704 | GM | clr | clr | clr | hf | * | *
2025-12-09 | 343 | 1336 | HB | clr | clr | clr | lf | * | *
2025-12-10 | 344 | 1714 | GM | rm lf | rm lf | clr | lf | * | *
2025-12-11 | 345 | 1323 | GM | rm lf | rm lf | clr | lf | * | *
2025-12-12 | 346 | 1630 | GM | rm lf | rm lf | clr | clr | * | *
2025-12-13 | 347 | 1235 | GM | rm lf | rm lf | clr | clr | * | *
2025-12-14 | 348 | 1232 | HB | rm lf | rm lf | clr | lf | * | *
2025-12-15 | 349 | 1636 | GM | rm lf | rm lf | clr | hf | * | *
2025-12-16 | 350 | 1328 | GM | rm lf | rm lf | clr | hf | * | *
2025-12-17 | 351 | 1713 | GM | rm lf | rm lf | clr | hf | * | *
2025-12-18 | 352 | 1338 | GM | rm lf | rm lf | clr | hf | * | * | GML page not updating since 2025-12-17 2045Z
2025-12-19 | 353 | 1256 | GM | rm lf | rm lf | clr | hf | * | * | GML page not updating since 2025-12-17 2045Z
2025-12-20 | 354 | 1218 | GM | rm lf | rm lf | clr | hf | * | * | GML page not updating since 2025-12-17 2045Z
2025-12-21 | 355 | - | GM | - | - | - | - | - | - | Did not go to AWO due to Con 1. GML Page down and showing nothing.
2025-12-22 | 356 | 1450 | GM | rm lf | rm lf | clr | clr | * | * | GML page back up as of 1413Z.
2025-12-23 | 357 | 1322 | GM | rm lf | rm lf | clr | clr | * | *
2025-12-24 | 358 | 1230 | GM | rm lf | rm lf | clr | clr | * | *
2025-12-25 | 359 | - | GM | - | - | - | - | * | * | Did not go out to AWO
2025-12-26 | 360 | 1638 | GM | rm lf | rm lf | clr | lf | * | *
2025-12-27 | 361 | 1745 | GM | clr | clr | clr | clr | * | *
2025-12-28 | 362 | 1248 | GM | rm lf | rm lf | clr | clr | * | *
2025-12-29 | 363 | 1434 | GM | rm lf | rm lf | clr | clr | * | *
2025-12-30 | 364 | 1258 | GM | clr | clr | clr | clr | * | *
2025-12-31 | 365 | - | GM | - | - | - | - | * | * | Did not go out to AWO
Note on E1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Shorter enclosure Reply: Aspirated temperature (short)
Note on F1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Longer enclosure Reply: Aspirated humidity (longer)
Note on G1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Ventus 2D Sonic Anemometer (V2) Reply: On NOAA boom with Wingbird
Note on H1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Heated Spinning Anemometer Reply: Secondary anemometer. There to record data if 9m Sonic primary anemometer is under maintenance. Wingbird on NOAA boom.
Note on I1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Refer to NOAA website T Logger Reply: NOAA enclosure (power enclosure on tower)- heater on and working Reply: NOAA enclosure - heater on and working but did not compensate with cooler temperatures very well. PI (Andrew Clarke) says that > -20 should be no cause for concern. Reply: Might have a fan running with a little heat coming out the top.
Note on J1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: NOAA website

## Inlet
Date | Julian Day | Time (UTC) | Observer's Initials | Neph/ Clap inlet | HATS inlet | Prototype inlet deicer | Sfc O3 inlet | Comments: climbs, filter changes, instrument work
2025-01-01 | 1 | 1502 | ALW | rm lf | rm lf | - | replaced filter | Replaced Sfc 03 filter @ 1512Z
2025-01-02 | 2 | 1456 | ALW | rm lf | rm lf | - | rm lf
2025-01-03 | 3 | - | ALW | - | - | - | - | ICEsat- did not go to AWO
2025-01-04 | 4 | 1252 | ALW | rm lf | rm lf | - | rm lf
2025-01-05 | 5 | 1500 | ALW | rm lf | clr | - | rm lf
2025-01-06 | 6 | 1527 | ALW | rm lf | rm f | - | rm lf
2025-01-07 | 7 | 1648 | ALW | rm lf | rm lf | - | rm lf
2025-01-08 | 8 | 1721 | ALW | - | - | - | - | Windy- did not go to roof
2025-01-09 | 9 | 1618 | ALW | rm lf | clr | - | rm lf
2025-01-10 | 10 | - | ALW | - | - | - | - | Con 1- did not go to AWO
2025-01-11 | 11 | 1237 | ALW | rm lf | rm lf | - | rm lf
2025-01-12 | 12 | 1252 | ALW | rm lf | rm lf | - | rm lf
2025-01-13 | 13 | 1612 | ALW | rm lf | rm lf | - | rm lf
2025-01-14 | 14 | 1643 | ALW | rm lf | lf | - | reattached | Reattached orange cap @ 1643Z
2025-01-15 | 15 | 1610 | ALW | rm lf | lf | - | rm lf
2025-01-16 | 16 | - | ALW | - | - | - | - | Con 1- did not go to AWO
2025-01-17 | 17 | - | ALW | - | - | - | - | Con 1- did not go to AWO
2025-01-18 | 18 | 1305 | ALW | rm lf | rm lf | - | rm lf
2025-01-19 | 19 | 1512 | ALW | rm lf | rm lf | - | rm lf
2025-01-20 | 20 | 1520 | ALW | rm lf | rm lf | - | rm lf
2025-01-21 | 21 | - | ALW | - | - | - | - | Too windy to go to roof
2025-01-22 | 22 | 1731 | ALW | rm lf | rm lf | - | rm lf | Found filter blew out of orange cap during high winds. Replaced @ 1857
2025-01-23 | 23 | 1717 | ALW | rm lf | rm lf | - | rm lf
2025-01-24 | 24 | 1320 | ALW | rm lf | rm lf | - | rm lf
2025-01-25 | 25 | 1202 | ALW | rm lf | rm lf | - | rm lf
2025-01-26 | 26 | 1448 | ALW | rm lf | rm lf | - | rm lf
2025-01-27 | 27 | 1520 | ALW | rm lf | rm lf | - | rm lf
2025-01-28 | 28 | 1824 | ALW | rm lf | rm lf | - | rm lf
2025-01-29 | 29 | - | ALW | - | - | - | - | Too windy to go to roof
2025-01-30 | 30 | 1514 | ALW | rm lf | rm lf | - | rm lf
2025-01-31 | 31 | 1318 | ALW | rm lf | rm lf | - | rm lf | Installed SFO candy cane @ 1540Z
2025-02-01 | 32 | - | ALW | - | - | - | - | Con 1- did not go to AWO
2025-02-02 | 33 | 1452 | ALW | rm lf | rm lf | - | rm lf
2025-02-03 | 34 | 1453 | ALW | rm lf | rm lf | - | rm lf
2025-02-04 | 35 | - | ALW | - | - | - | - | ICEsat- did not go to AWO
2025-02-05 | 36 | 1827 | ALW | rm lf | rm lf | - | rm lf
2025-02-06 | 37 | 1219 | ALW | rm lf | rm lf | - | rm lf
2025-02-07 | 38 | - | ALW | - | - | - | - | Con 1- did not go to AWO
2025-02-08 | 39 | 1758 | ALW | rm lf | rm lf | - | rm lf
2025-02-09 | 40 | - | ALW | - | - | - | - | Con 1- did not go to AWO
2025-02-10 | 41 | - | ALW | - | - | - | - | ICEsat- did not go to AWO
2025-02-11 | 42 | 1416 | ALW | rm lf | rm lf | - | rm lf
2025-02-12 | 43 | 1047 | ALW | rm lf | rm lf | - | rm lf
2025-02-13 | 44 | 1659 | ALW | rm lf | rm lf | - | rm lf
2025-02-14 | 45 | 1349 | AR | rm lf | rm lf | - | rm lf
2025-02-15 | 46 | - | - | - | - | - | - | ICEsat- did not go to AWO. Carbon vane pump failure at 1238Z. Pump was restarted at 1704Z after replacing carbon vanes.
2025-02-16 | 47 | 1541 | AR | rm lf | rm lf | - | rm lf
2025-02-17 | 48 | 1225 | AR | rm lf | rm lf | - | rm lf | Planned power outages for turnover 1201-1204Z and 1822-1824Z. Instruments supported by UPS.
2025-02-18 | 49 | 1259 | AR | rm lf | rm lf | - | rm lf
2025-02-19 | 50 | 1749 | AR | rm lf | rm lf | - | rm lf
2025-02-20 | 51 | 1328 | AR | rm lf | rm lf | - | rm lf | Power outage to UPS supported instruments from 0046-1147Z. Caused by carbon vane pump failure. Pump moved to non-UPS circuit.
2025-02-21 | 52 | 1302 | AR | rm lf | rm lf | - | rm lf
2025-02-22 | 53 | 1202 | AR | rm lf | rm lf | - | rm lf | Inlet pump turned off 1107-1343Z. Inlet head removed for cleaning and pump maintanence.
2025-02-23 | 54 | 1246 | AR | rm lf | rm lf | - | rm lf
2025-02-24 | 55 | 1520 | AR | rm lf | rm lf | - | rm lf
2025-02-25 | 56 | 1850 | AR | rm lf | rm lf | - | rm lf | Carbon vane pump failure at 0608Z
2025-02-26 | 57 | 1328 | AR | rm lf | - | - | rm lf | Replaced carbon vane pump with smaller cv pump at 1258Z.
2025-02-27 | 58 | 1241 | AR | rm lf | rm lf | - | rm lf | Replaced cv pump with stronger diaphragm pump at 1633Z.
2025-02-28 | 59 | - | - | - | - | - | - | Con 1- did not go to AWO
2025-03-01 | 60 | 1240 | AR | rm lf | rm lf | - | rm lf
2025-03-02 | 61 | 1537 | AR | rm lf | rm lf | - | rm lf
2025-03-03 | 62 | 1452 | AR | rm lf | rm lf | - | rm lf
2025-03-04 | 63 | 1519 | AR | rm lf | rm lf | - | rm lf
2025-03-05 | 64 | - | AR | - | - | - | - | ICEsat- did not go to AWO
2025-03-06 | 65 | 1307 | AR | rm lf | rm lf | - | rm lf
2025-03-07 | 66 | 1647 | AR | rm lf | rm lf | - | rm lf
2025-03-08 | 67 | 1112 | AR | rm lf | rm lf | - | rm lf
2025-03-09 | 68 | - | AR | - | - | - | - | Con 1- did not go to AWO
2025-03-10 | 69 | 1218 | AR | rm lf | rm lf | - | rm lf
2025-03-11 | 70 | 1242 | AR | rm lf | clr | - | rm lf | Power outage today at 2150Z (less than 1 min). Instruments supported by UPS.
2025-03-12 | 71 | 1609 | AR | rm lf | rm lf | - | rm lf
2025-03-13 | 72 | 1256 | AR | rm lf | rm lf | - | rm lf
2025-03-14 | 73 | 1137 | AR | rm lf | rm lf | - | rm lf
2025-03-15 | 74 | 1233 | AR | rm lf | rm lf | - | rm lf
2025-03-16 | 75 | 1330 | AR | rm lf | clr | - | rm lf
2025-03-17 | 76 | 1505 | AR | rm lf | rm lf | - | rm lf
2025-03-18 | 77 | 1155 | AR | rm lf | clr | - | rm lf
2025-03-19 | 78 | 1714 | AR | rm lf | rm lf | - | rm lf
2025-03-20 | 79 | 1110 | AR | rm lf | rm lf | - | rm lf
2025-03-21 | 80 | 1505 | AR | rm lf | rm lf | - | rm lf
2025-03-22 | 81 | 1208 | AR | rm lf | rm lf | - | rm lf
2025-03-23 | 82 | 1248 | AR | rm lf | rm lf | - | rm lf
2025-03-24 | 83 | 1613 | AR | rm lf | rm lf | - | rm lf
2025-03-25 | 84 | 1219 | AR | rm lf | rm lf | - | rm lf
2025-03-26 | 85 | 1635 | AR | rm lf | rm lf | - | rm lf
2025-03-27 | 86 | 1614 | AR | rm lf | rm lf | - | rm lf
2025-03-28 | 87 | 1153 | AR | rm lf | rm lf | - | rm lf
2025-03-29 | 88 | 1218 | AR | rm lf | rm lf | - | rm lf
2025-03-30 | 89 | 1215 | AR | rm lf | rm lf | - | rm lf
2025-03-31 | 90 | 1450 | AR | rm lf | rm lf | - | rm lf
2025-04-01 | 91 | 946 | AR | rm lf | rm lf | - | rm lf
2025-04-02 | 92 | - | AR | - | - | - | - | ICEsat- did not go to AWO
2025-04-03 | 93 | 1502 | AR | rm lf | rm lf | - | rm lf
2025-04-04 | 94 | 1518 | AR | rm lf | rm lf | - | rm lf
2025-04-05 | 95 | 1047 | AR | rm lf | rm lf | - | rm lf
2025-04-06 | 96 | 1210 | AR | rm lf | clr | - | rm lf
2025-04-07 | 97 | 1123 | AR | rm lf | clr | - | rm lf
2025-04-08 | 98 | 1154 | AR | rm lf | rm lf | - | rm lf
2025-04-09 | 99 | 1038 | AR | rm lf | rm lf | - | rm lf
2025-04-10 | 100 | 1421 | AR | rm lf | clr | - | rm lf
2025-04-11 | 101 | 1025 | AR | rm lf | rm lf | - | rm lf
2025-04-12 | 102 | 1152 | AR | rm lf | lf | - | rm lf
2025-04-13 | 103 | 1602 | AR | rm lf | rm lf | - | rm lf
2025-04-14 | 104 | 1504 | AR | rm lf | clr | - | rm lf
2025-04-15 | 105 | 1054 | AR | rm lf | rm lf | - | rm lf
2025-04-16 | 106 | 1406 | AR | rm lf | rm lf | - | rm lf
2025-04-17 | 107 | 1038 | AR | clr | clr | - | clr
2025-04-18 | 108 | 1201 | AR | clr | clr | - | clr
2025-04-19 | 109 | 1112 | AR | clr | clr | - | clr
2025-04-20 | 110 | 1225 | AR | rm lf | clr | - | rm lf
2025-04-21 | 111 | 1511 | AR | rm lf | rm lf | - | rm lf
2025-04-22 | 112 | 1452 | AR | rm lf | rm lf | - | rm lf
2025-04-23 | 113 | 1013 | AR | rm lf | rm lf | - | rm lf
2025-04-24 | 114 | 1630 | AR | rm lf | clr | - | rm lf
2025-04-25 | 115 | 1127 | AR | rm lf | rm lf | - | rm lf
2025-04-26 | 116 | 1437 | AR | rm lf | rm lf | - | rm lf
2025-04-27 | 117 | 1140 | AR | rm lf | clr | - | rm lf
2025-04-28 | 118 | 1623 | AR | clr | clr | - | rm lf
2025-04-29 | 119 | 1059 | AR | rm lf | clr | - | rm lf
2025-04-30 | 120 | 1425 | AR | clr | clr | - | clr
2025-05-01 | 121 | 1115 | AR | clr | clr | - | rm lf
2025-05-02 | 122 | 1151 | AR | rm lf | rm lf | - | rm lf
2025-05-03 | 123 | 1018 | AR | rm lf | rm lf | - | clr
2025-05-04 | 124 | 1930 | AR | clr | clr | - | clr
2025-05-05 | 125 | 1120 | AR | clr | clr | - | clr
2025-05-06 | 126 | 1255 | AR | rm lf | clr | - | rm lf
2025-05-07 | 127 | 1454 | AR | clr | clr | - | rm lf
2025-05-08 | 128 | 1212 | AR | clr | clr | - | rm lf
2025-05-09 | 129 | 1136 | AR | rm lf | rm lf | - | rm lf
2025-05-10 | 130 | 1026 | AR | rm lf | rm lf | - | rm lf | Planned station-wide power outage from 1008-1014Z. Instruments supported by UPS.
2025-05-11 | 131 | 1544 | AR | clr | clr | - | rm lf | Unplanned station-wide power outage from 1455-1524Z. Instruments supported by UPS.
2025-05-12 | 132 | 1104 | AR | rm lf | clr | - | rm lf
2025-05-13 | 133 | 1811 | AR | - | - | - | - | Con 1- did not go to AWO. Unplanned station power outage from 2031-2040Z. Controlled station power outage from 2105-2110Z. Instruments supported by UPS.
2025-05-14 | 134 | 944 | AR | clr | clr | - | clr
2025-05-15 | 135 | 1504 | AR | clr | clr | - | clr
2025-05-16 | 136 | 1035 | AR | clr | clr | - | clr
2025-05-17 | 137 | 931 | AR | clr | clr | - | clr | Planned station power outage from 1015-1019Z. Instruments supported by UPS.
2025-05-18 | 138 | 1149 | AR | clr | clr | - | clr
2025-05-19 | 139 | 1547 | AR | clr | clr | - | clr
2025-05-20 | 140 | 1006 | AR | rm lf | rm lf | - | rm lf
2025-05-21 | 141 | 1434 | AR | clr | clr | - | clr
2025-05-22 | 142 | 952 | AR | clr | clr | - | clr
2025-05-23 | 143 | 1213 | AR | rm lf | clr | - | rm lf
2025-05-24 | 144 | 1026 | AR | clr | clr | - | clr | Unplanned station-wide power outage from 1806-1808Z. Instruments supported by UPS.
2025-05-25 | 145 | 1300 | AR | clr | clr | - | rm lf
2025-05-26 | 146 | 1005 | AR | clr | rm lf | - | clr
2025-05-27 | 147 | 1454 | AR | clr | clr | - | clr
2025-05-28 | 148 | 1123 | AR | rm lf | lf | - | rm lf
2025-05-29 | 149 | 1444 | AR | clr | clr | - | rm lf
2025-05-30 | 150 | 1137 | AR | clr | clr | - | rm lf
2025-05-31 | 151 | 1550 | AR | clr | clr | - | clr
2025-06-01 | 152 | 1745 | AR | clr | clr | - | clr
2025-06-02 | 153 | 1543 | AR | clr | clr | - | clr
2025-06-03 | 154 | 1048 | AR | clr | clr | - | clr
2025-06-04 | 155 | - | AR | - | - | - | - | Icesat- did not go to AWO
2025-06-05 | 156 | 1006 | AR | clr | rm lf | - | rm lf
2025-06-06 | 157 | 1530 | AR | clr | clr | - | clr
2025-06-07 | 158 | 1045 | AR | rm lf | lf | - | rm lf
2025-06-08 | 159 | 1341 | AR | clr | clr | - | rm lf
2025-06-09 | 160 | 1109 | AR | clr | clr | - | clr | Planned power outage 1001-1005Z. Instruments supported by UPS.
2025-06-10 | 161 | 1428 | AR | rm lf | clr | - | rm lf | Planned power outage 1704-1707Z. Instruments supported by UPS.
2025-06-11 | 162 | 1439 | AR | clr | clr | - | rm lf
2025-06-12 | 163 | 951 | PC | clr | clr | - | rm lf
2025-06-13 | 164 | 951 | PC | clr | clr | - | clr | Planned power outage 1037-1040Z. Second planned power outage 1412-1416Z. Instruments supported by UPS.
2025-06-14 | 165 | 1611 | PC | clr | clr | - | clr
2025-06-15 | 166 | 1139 | PC | clr | clr | - | clr
2025-06-16 | 167 | 1413 | PC | clr | clr | - | clr
2025-06-17 | 168 | 1126 | PC | clr | clr | - | rm lf | 1545-1548Z Riggers near inlet, orientation to small tower on top of AWO roof.
2025-06-18 | 169 | 1224 | PC | clr | clr | - | clr
2025-06-19 | 170 | 1049 | PC | rm lf | rm lf | - | rm lf
2025-06-20 | 171 | 1445 | PC | clr | clr | - | clr
2025-06-21 | 172 | 1627 | PC | clr | clr | - | clr
2025-06-22 | 173 | 1534 | PC | rm lf | rm lf | - | rm lf
2025-06-23 | 174 | 1135 | PC | clr | clr | - | clr
2025-06-24 | 175 | 905 | PC | clr | clr | - | rm lf | Plane flew a big low contrail directly over station 0905.
2025-06-25 | 176 | 1123 | PC | clr | clr | - | rm lf | Low winds equipment operations: 11:42-12:07Z- D6, 953, skidsteer, snowmo, welder (machines operating 11:42-11:49, welder missed radio call and operated 11:42-1207Z).
2025-06-26 | 177 | 940 | PC | clr | clr | - | rm lf
2025-06-27 | 178 | 1422 | PC | clr | clr | - | clr
2025-06-28 | 179 | 924 | PC | clr | clr | - | clr | 1005-1006Z: Planned power outage for mechanics training. No data impacted for instruments supported by UPS. 1010-1012Z: second planned power outage to switch back to main generator.
2025-06-29 | 180 | 1522 | PC | clr | clr | - | clr
2025-06-30 | 181 | 1401 | PC | clr | clr | - | clr
2025-07-01 | 182 | 958 | PC | clr | clr | - | clr | Planned power outage due to generator issues last night, power swap from 1133-1135Z, again from 1224-1225Z. And again 1602-1604Z. Instruments supported by UPS.
2025-07-02 | 183 | 1006 | PC | rm lf | rm lf | - | rm lf
2025-07-03 | 184 | 1128 | PC | clr | clr | - | rm lf
2025-07-04 | 185 | 1056 | PC | clr | clr | - | clr
2025-07-05 | 186 | 1631 | PC | clr | clr | - | clr | Northwinds operations: 1716-1735Z D6 Snow machine 953 953
2025-07-06 | 187 | 946 | PC | clr | clr | - | clr
2025-07-07 | 188 | 1119 | PC | clr | clr | - | clr | 1602-1603Z: power swap back to main station generator. Instruments supported by UPS.
2025-07-08 | 189 | 1011 | PC | clr | clr | - | clr
2025-07-09 | 190 | 1043 | PC | rm lf | rm lf | - | rm lf | LOW WINDS OPERATIONS 1015-1018Z: D6 HEAT CANNON, 5K
2025-07-10 | 191 | 1007 | PC | clr | clr | - | clr
2025-07-11 | 192 | 1101 | PC | rm lf | rm lf | - | rm lf
2025-07-12 | 193 | 1028 | PC | rm lf | rm lf | - | rm lf
2025-07-13 | 194 | 1229 | PC | rm ice | clr | - | rm ice
2025-07-14 | 195 | 1039 | PC | clr | clr | - | clr
2025-07-15 | 196 | 1031 | PC | rm ice | rm ice | - | rm ice | Northwind ops: x2 variances (welding, medical to heat isolation bay during day) from 0900-1400
2025-07-16 | 197 | 1047 | PC | clr | clr | - | rm lf
2025-07-17 | 198 | 1148 | PC | clr | clr | - | rm lf
2025-07-18 | 199 | 1009 | PC | clr | clr | - | clr
2025-07-19 | 200 | 1019 | PC | rm ice | rm ice | - | rm ice
2025-07-20 | 201 | 1139 | PC | clr | clr | - | clr
2025-07-21 | 202 | 1034 | PC | clr | clr | - | clr
2025-07-22 | 203 | 1044 | PC | clr | clr | - | rm lf
2025-07-23 | 204 | 1017 | PC | clr | clr | - | rm lf | Low winds from 0943-1226Z, 1803-1908Z, North winds 1206-1217Z, 1851-1942Z, 2052-2150Z
2025-07-24 | 205 | 1025 | PC | clr | clr | - | rm lf | Low winds from 0823-1429Z, 1643-1643Z. North winds from 0051-0402Z, 0932-1230Z
2025-07-25 | 206 | 1013 | PC | clr | clr | - | clr
2025-07-26 | 207 | 1202 | PC | clr | clr | - | clr | Low winds from 0223-0224Z. North winds from 2314-0054Z
2025-07-27 | 208 | 1857 | PC | clr | clr | - | clr | Low winds from 0520-0520Z, 0722-0741Z, 2147-2357Z. North winds from 0239-0520Z, 2202-2340Z
2025-07-28 | 209 | 1035 | PC | clr | clr | - | rm lf
2025-07-29 | 210 | 1046 | PC | clr | clr | - | clr | Low winds 09:17Z to 09:56Z
2025-07-30 | 211 | 1026 | PC | clr | clr | - | clr | Low winds 02:10Z to 02:11Z, north winds 09:19Z to 09:25Z 11:47Z to 13:28Z 19:54Z to 21:04Z
2025-07-31 | 212 | 1524 | PC | clr | clr | - | clr
2025-08-01 | 213 | 1201 | PC | clr | clr | - | clr
2025-08-02 | 214 | 1452 | PC | clr | clr | - | rm lf
2025-08-03 | 215 | 1125 | PC | clr | clr | - | clr
2025-08-04 | 216 | 1503 | PC | clr | clr | - | clr
2025-08-05 | 217 | 958 | PC | rm lf | rm lf | - | rm lf
2025-08-06 | 218 | 1127 | PC | clr | clr | - | clr
2025-08-07 | 219 | 1058 | PC | clr | clr | - | clr
2025-08-08 | 220 | 1006 | PC | rm lf | rm lf | - | rm lf
2025-08-09 | 221 | 1226 | PC | clr | clr | - | clr | Inlet pump unplugged at 1649Z and moved to GFCI outlet.
2025-08-10 | 222 | 1404 | PC | clr | clr | - | clr
2025-08-11 | 223 | 1108 | PC | clr | clr | - | clr
2025-08-12 | 224 | 1503 | PC | clr | clr | - | rm lf
2025-08-13 | 225 | 102 | PC | rm lf | rm lf | - | rm lf | 1115-1135Z, 1145-1200 ICECAPS team on roof near inlet
2025-08-14 | 226 | 1022 | PC | rm lf | rm lf | - | rm lf
2025-08-15 | 227 | 909 | PC | clr | clr | - | clr
2025-08-16 | 228 | 1032 | PC | clr | clr | - | clr
2025-08-17 | 229 | 1429 | PC | clr | clr | - | clr
2025-08-18 | 230 | 1003 | PC | clr | clr | - | clr
2025-08-19 | 231 | 1126 | PC | rm lf | rm lf | - | rm lf
2025-08-20 | 232 | 1210 | PC | rm lf | rm lf | - | rm lf
2025-08-21 | 233 | 1052 | PC | rm lf | rm lf | - | rm lf
2025-08-22 | 234 | 1056 | PC | clr | clr | - | clr
2025-08-23 | 235 | 1624 | PC | clr | clr | - | clr
2025-08-24 | 236 | 1012 | PC | clr | clr | - | clr
2025-08-25 | 237 | 1300 | PC | clr | clr | - | clr | Planned power swap, instruments supported by ups 1002-1004Z
2025-08-26 | 238 | 1028 | PC | rm lf | rm lf | - | rm lf | Planned power swap, instruments supported by ups 1101-1102Z
2025-08-27 | 239 | 954 | Pc | clr | clr | - | clr
2025-08-28 | 240 | 1052 | PC | clr | clr | - | clr
2025-08-29 | 241 | 1141 | PC | rm lf | rm lf | - | rm lf
2025-08-30 | 242 | 1033 | PC | clr | clr | - | clr
2025-08-31 | 243 | 1313 | PC | clr | clr | - | clr
2025-09-01 | 244 | 1058 | PC | RM LF | RM LF | - | RM LF
2025-09-02 | 245 | 1110 | PC | RM LF | RM LF | - | RM LF
2025-09-03 | 246 | 1023 | PC | CLR | CLR | - | CLR
2025-09-04 | 247 | 1058 | PC | CLR | CLR | - | CLR
2025-09-05 | 248 | 1234 | PC | CLR | CLR | - | CLR
2025-09-06 | 249 | 1037 | PC | CLR | CLR | - | CLR
2025-09-07 | 250 | 1445 | PC | RM LF | RM LF | - | RM LF
2025-09-08 | 251 | 1211 | PC | CLR | CLR | - | CLR
2025-09-09 | 252 | 1047 | PC | CLR | CLR | - | CLR
2025-09-10 | 253 | 1454 | PC | CLR | CLR | - | CLR
2025-09-11 | 254 | 1042 | PC | CLR | CLR | - | CLR
2025-09-12 | 255 | 1128 | PC | RM LF | RM LF | - | RM LF
2025-09-13 | 256 | 1104 | PC | CLR | CLR | - | CLR
2025-09-14 | 257 | 1538 | PC | CLR | CLR | - | CLR
2025-09-15 | 258 | 1109 | PC | RM LF | RM LF | - | RM LF
2025-09-16 | 259 | 1147 | PC | RM LF | RM LF | - | RM LF
2025-09-17 | 260 | 1448 | PC | CLR | CLR | - | CLR
2025-09-18 | 261 | 1001 | PC | RM LF | RM LF | - | RM LF
2025-09-19 | 262 | 1128 | PC | RM LF | RM LF | - | RM LF
2025-09-20 | 263 | 1118 | PC | RM LF | RM LF | - | RM LF
2025-09-21 | 264 | 1143 | PC | CLR | CLR | - | CLR
2025-09-22 | 265 | 1230 | PC | RM LF | RM LF | - | RM LF
2025-09-23 | 266 | 1127 | PC | RM LF | RM LF | - | RM LF
2025-09-24 | 267 | 1138 | PC | RM LF | RM LF | - | RM LF
2025-09-25 | 268 | 1120 | PC | RM LF | RM LF | - | RM LF
2025-09-26 | 269 | 1134 | PC | RM LF | RM LF | - | RM LF
2025-09-27 | 270 | 1212 | PC | RM LF | RM LF | - | RM LF
2025-09-28 | 271 | 1451 | PC | RM LF | RM LF | - | RM LF
2025-09-29 | 272 | 1247 | PC | RM LF | RM LF | - | RM LF
2025-09-30 | 273 | 1417 | PC | CLR | CLR | - | CLR
2025-10-01 | 274 | 1417 | PC | CLR | CLR | - | CLR
2025-10-02 | 275 | 1200 | PC | RM LF | RM LF | - | RM LF
2025-10-03 | 276 | 1119 | PC | RM LF | RM LF | - | RM LF
2025-10-04 | 277 | 1149 | PC | RM LF | RM LF | - | RM LF
2025-10-05 | 278 | 1619 | PC | RM LF | RM LF | - | RM LF
2025-10-06 | 279 | 1222 | PC | RM LF | RM LF | - | RM LF
2025-10-07 | 280 | 1143 | PC | RM LF | RM LF | - | RM LF
2025-10-08 | 281 | 1155 | PC | - | - | - | - | CON-1 did not go to AWO
2025-10-09 | 282 | 1443 | PC | RM LF | RM LF | - | RM LF
2025-10-10 | 283 | 1113 | PC | RM LF | RM LF | - | RM LF
2025-10-11 | 284 | 1150 | PC | RM LF | RM LF | - | RM LF
2025-10-12 | 285 | 1249 | GM, PC | RM LF | CLR | - | CLR
2025-10-13 | 286 | 1518 | PC | RM LF | CLR | - | RM LF
2025-10-14 | 287 | 1238 | GM, PC | - | - | - | - | Did not go on roof due to high winds
2025-10-15 | 288 | 1627 | GM | RM LF | RM LF | - | RM LF
2025-10-16 | 289 | 1017 | GM | CLR | CLR | - | CLR
2025-10-17 | 290 | 1109 | GM | RM LF | CLR | - | CLR
2025-10-18 | 291 | 1236 | GM | rm lf | rm lf | - | rm lf
2025-10-19 | 292 | 1327 | GM | rm lf | rm lf | - | rm lf
2025-10-20 | 293 | 1244 | GM | rm lf | rm lf | - | rm lf
2025-10-21 | 294 | 1258 | GM | rm lf | rm lf | - | rm lf
2025-10-22 | 295 | 1637 | GM | clr | clr | - | clr
2025-10-23 | 296 | 1227 | GM | rm lf | clr | - | rm lf
2025-10-24 | 297 | 1631 | GM | rm lf | rm lf | - | rm lf
2025-10-25 | 298 | 1032 | GM | clr | clr | - | clr
2025-10-26 | 299 | 1258 | GM | rm lf | rm lf | - | rm lf
2025-10-27 | 300 | 1612 | GM | rm lf | rm lf | - | rm lf
2025-10-28 | 301 | 1515 | GM | rm lf | rm lf | - | rm lf
2025-10-29 | 302 | 1302 | GM | rm lf | rm lf | - | rm lf
2025-10-30 | 303 | 1502 | GM | rm lf | rm lf | - | rm lf
2025-10-31 | 304 | 1610 | GM | rm lf | rm lf | - | rm lf
2025-11-01 | 305 | 1317 | GM | rm lf | rm lf | - | rm hf
2025-11-02 | 306 | 1408 | GM | rm lf | clr | - | clr
2025-11-03 | 307 | - | GM | - | - | - | - | no in person checks due to ICESat
2025-11-04 | 308 | 1215 | GM | rm lf | rm lf | - | rm lf
2025-11-05 | 309 | 1404 | HB | rm lf | rm lf | - | rm lf
2025-11-06 | 310 | 1334 | GM | rm lf | clr | - | clr
2025-11-07 | 311 | 1505 | GM | rm lf | rm lf | - | rm lf
2025-11-08 | 312 | 1204 | GM | rm lf | rm lf | - | rm lf
2025-11-09 | 313 | 1408 | HB | rm lf | rm lf | - | lm lf
2025-11-10 | 314 | 1304 | GM | rm lf | rm lf | - | lm lf
2025-11-11 | 315 | 1810 | GM | - | - | - | - | Did not go to AWO due to ICESat
2025-11-12 | 316 | 1557 | GM | rm lf | rm lf | - | rm lf
2025-11-13 | 317 | 1220 | GM | clr | clr | - | clr
2025-11-14 | 318 | - | GM | - | - | - | - | Did not go on roof due to high winds
2025-11-15 | 319 | 1232 | GM | rm lf | rm lf | - | rm lf
2025-11-16 | 320 | 1315 | GM | rm lf | clr | - | rm lf
2025-11-17 | 321 | 1621 | GM | rm lf | rm lf | - | rm lf
2025-11-18 | 322 | 1234 | GM | rm lf | rm lf | - | rm lf
2025-11-19 | 323 | 1310 | GM | rm lf | rm lf | - | rm lf
2025-11-20 | 324 | 1324 | GM | rm lf | rm lf | - | rm lf
2025-11-21 | 325 | 1502 | GM | rm lf | rm lf | - | rm lf
2025-11-22 | 326 | 1158 | GM | rm lf | rm lf | - | rm lf
2025-11-23 | 327 | 1243 | GM | rm lf | rm lf | - | rm lf
2025-11-24 | 328 | 1415 | GM | rm lf | rm lf | - | rm lf
2025-11-25 | 329 | 1252 | GM | rm lf | rm lf | - | rm lf
2025-11-26 | 330 | 1655 | GM | rm lf | rm lf | - | rm lf
2025-11-27 | 331 | [6 empty] | Did not go out the AWO
2025-11-28 | 332 | 1348 | GM | rm lf | rm lf | - | rm lf
2025-11-29 | 333 | 1205 | GM | rm lf | clr | - | rm lf
2025-11-30 | 334 | 1246 | GM | rm lf | clr | - | rm lf
2025-12-01 | 335 | 1723 | GM | rm lf | rm lf | - | rm lf
2025-12-02 | 336 | - | GM | - | - | - | - | Did not go out to AWO due to ICESat
2025-12-03 | 337 | 1655 | GM | rm lf | rm lf | - | rm lf
2025-12-04 | 338 | 1414 | GM | rm lf | rm lf | - | rm lf
2025-12-05 | 339 | 1645 | GM | rm lf | rm lf | - | rm hf
2025-12-06 | 340 | 1235 | GM | rm lf | rm lf | - | rm hf
2025-12-07 | 341 | 1332 | GM | rm lf | rm lf | - | rm hf
2025-12-08 | 342 | 1653 | GM | rm lf | rm lf | - | rm hf
2025-12-09 | 343 | 1330 | HB | rm lf | rm lf | - | rm lf
2025-12-10 | 344 | 1706 | GM | rm lf | rm lf | - | rm lf
2025-12-11 | 345 | 1313 | GM | rm lf | rm lf | - | rm lf
2025-12-12 | 346 | - | GM | - | - | - | - | Did not go on roof due to high winds
2025-12-13 | 347 | 1228 | GM | rm lf | clr | - | clr
2025-12-14 | 348 | 1226 | HB | rm lf | rm lf | - | rm lf
2025-12-15 | 349 | 1616 | GM | rm lf | rm lf | - | rm lf
2025-12-16 | 350 | - | GM | - | - | - | - | Did not go on roof due to high winds
2025-12-17 | 351 | 1645 | GM | rm lf | rm lf | - | rm lf
2025-12-18 | 352 | 1330 | GM | rm lf | rm lf | - | rm lf
2025-12-19 | 353 | 1245 | GM | rm lf | rm lf | - | rm lf
2025-12-20 | 354 | 1207 | GM | rm lf | rm lf | - | rm lf
2025-12-21 | 355 | - | GM | - | - | - | - | Did not go to AWO due to Con 1
2025-12-22 | 356 | 1447 | GM | clr | clr | - | clr
2025-12-23 | 357 | 1316 | GM | rm lf | rm lf | - | clr
2025-12-24 | 358 | 1220 | GM | rm lf | rm lf | - | rm lf
2025-12-25 | 359 | - | GM | - | - | - | - | Did not go out to AWO
2025-12-26 | 360 | 1632 | GM | rm lf | rm lf | - | rm lf
2025-12-27 | 361 | - | GM | - | - | - | - | Did not go on roof due to high winds
2025-12-28 | 362 | 1236 | GM | rm lf | rm lf | - | rm lf
2025-12-29 | 363 | 1428 | GM | clr | rm lf | - | rm lf
2025-12-30 | 364 | 1254 | GM | clr | clr | - | clr | Did not go on roof due to high winds, just a visual inspection from the deck
2025-12-31 | 365 | - | GM | - | - | - | - | Did not go out to AWO
Note on E1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Black and Gray Inlet Reply: Now has the pot over top Reply: aerosol inlet on s side railing Reply: Has mesh covering the intake to help keep rime out. Reach hand up pot and inside intake to remove rime from mesh.
Note on F1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Top Inlet- looks like bells Reply: On left side of mini tower on roof.
Note on G1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: On right side of mini tower on roof. Usually surrounded by rime. Leave be.
Note on H1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Orange Cap Inlet Reply: Has paperish filter inside to keep frost free and be replaced ~January and July 1st

## NEPH & CLAP
Date | Julian Day | Time (UTC) | Observer's Initials | Sample inlet? | Check Daily Aerosol Email | TSI flow lpm <12 | CLAP flow ~0.60 lpm | CLAP green tr>0.7 | CLAP Spot # | MAGIC 210 cm-3 > 0 | MAGIC 210 dewpoint > 10 | MAGIC 250 cm-3 > 0 | MAGIC 250 dewpoint ~-10C | Refilled MAGIC Reservoir | Monthly Span Check | NEPH Lamp On/Current (<7A) | NEPH Flag Status | NEPH Flow (~18 lpm) | Ecotech Flag status | Ecotech Flow (~4.5lpm) | Ecotech flag check on instrument (every other week) | Comments
2025-01-01 | 1 | 1637 | ALW | rm lf | * |  | 0.6 | 0.73806 | 4 |  |  | 37 | -12.2 | n | echotech: 1.6% | 5.6 | none | 18 | none | 4.5
2025-01-02 | 2 | 1551 | ALW | rm lf | * |  | 0.599 | 0.73507 | 4 |  |  | 41 | -12.2 | n |  | 5.6 | none | 18 | none | 4.5
2025-01-03 | 3 | - | ALW | - | * |  | - | - | - |  |  | - | - | - | - | - | - | - | - | - |  | ICEsat- did not go to AWO
2025-01-04 | 4 | 1347 | ALW | rm lf | * |  | 0.602 | 0.73329 | 4 |  |  | 13 | -12.2 | n |  | 5.6 | none | 18 | none | 4.5
2025-01-05 | 5 | 1526 | ALW | rm lf | * |  | 0.601 | 0.73284 | 4 |  |  | 18 | -12.2 | n |  | 5.6 | none | 18 | none | 4.5
2025-01-06 | 6 | 1702 | ALW | rm lf | * |  | 0.597 | 0.7298 | 4 |  |  | 53 | -12.2 | n |  | 5.6 | none | 18 | none | 4.5
2025-01-07 | 7 | 1847 | ALW | rm lf | * |  | 0.603 | 0.72816 | 4 |  |  | 45 | -12.2 | n |  | 5.6 | none | 18 | none | 4.5
2025-01-08 | 8 | 1745 | ALW | - | * |  | 0.591 | 0.7267 | 4 |  |  | 69 | -12.2 | n |  | 5.6 | none | 18 | none | 4.5
2025-01-09 | 9 | 1543 | ALW | rm lf | * |  | 0.597 | 0.72557 | 4 |  |  | 24 | -12.2 | n |  | 5.7 | none | 18 | none | 4.5
2025-01-10 | 10 | - | ALW | - | * |  | - | - | - |  |  | - | - | - | - | - | - | - | - | - |  | Con1- did not go to AWO
2025-01-11 | 11 | 1250 | ALW | rm lf | * |  | 0.599 | 0.71762 | 4 |  |  | 33 | -12.2 | n |  | 5.6 | none | 18 | none | 4.5 | pass
2025-01-12 | 12 | 1301 | ALW | rm lf | * |  | 0.587 | 0.71554 | 4 |  |  | 70 | -12.2 | n |  | 5.6 | none | 18 | none | 4.5
2025-01-13 | 13 | 1556 | ALW | rm lf | * |  | 0.587 | 0.71544 | 4 |  |  | 15 | -12.2 | n |  | 5.6 | none | 18 | none | 4.5
2025-01-14 | 14 | 1446 | ALW | rm lf | * |  | 0.597 | 0.7134 | 4 |  |  | 49 | -12.2 | y |  | 5.6 | none | 18 | none | 4.5
2025-01-15 | 15 | 1622 | ALW | rm lf | * |  | 0.593 | 0.71162 | 4 |  |  | 83 | -12.2 | n |  | 5.6 | none | 18 | none | 4.5
2025-01-16 | 16 | - | ALW | - | * |  | - | - | - |  |  |  | - | - | - | - | - | - | - |  |  | Con 1- did not go to AWO
2025-01-17 | 17 | 0 | ALW | - | * |  | 0.608 | 0.70735 | 4 |  |  | - | -12.2 | n |  | 5.6 | - | - | - | -
2025-01-18 | 18 | 1243 | ALW | rm lf | * |  | 0.593 | 0.70374 | 4 |  |  | 6 | -12.2 | n |  | 5.5 | none | 18 | none | 4.5
2025-01-19 | 19 | 1527 | ALW | rm lf | * |  | 0.603 | 0.99891 | 5 |  |  | 68 | -12.2 | n |  | 5.6 | none | 18 | none | 4.5
2025-01-20 | 20 | 1500 | ALW | rm lf | * |  | 0.599 | 0.99341 | 5 |  |  | 9 | -12.2 | n |  | 5.6 | none | 18 | none | 4.5
2025-01-21 | 21 | 1238 | ALW | - | * |  | 0.588 | 0.99226 | 5 |  |  | 16 | -12.2 | y |  | 5.6 | none | 18 | none | 4.5 | pass
2025-01-22 | 22 | 1538 | ALW | rm lf | * |  | 0.604 | 0.98831 | 5 |  |  | 16 | -12.2 | n |  | 5.6 | none | 18 | none | 4.5
2025-01-23 | 23 | 1603 | ALW | rm lf | * |  | 0.602 | 0.98774 | 5 |  |  | 8 | -12.2 | n |  | 5.6 | none | 18 | none | 4.5
2025-01-24 | 24 | 1304 | ALW | rm lf | * |  | 0.596 | 0.98473 | 5 |  |  | 36 | -12.2 | n |  | 5.7 | none | 18 | none | 4.5
2025-01-25 | 25 | 1141 | ALW | rm lf |  |  | 0.602 | 0.98279 | 5 |  |  | 20 | -12.2 | n |  | 5.5 | none | 18 | none | 4.5
2025-01-26 | 26 | 1412 | ALW | rm lf |  |  | 0.601 | 0.9788 | 5 |  |  | 21 | -12.2 | n |  | 5.5 | none | 18 | none | 4.5
2025-01-27 | 27 | 1452 | ALW | rm lf | * |  | 0.601 | 0.95949 | 5 |  |  | 32 | -12.2 | n |  | 5.5 | none | 18 | none | 4.5
2025-01-28 | 28 | 1510 | ALW | rm lf | * |  | 0.597 | 0.95581 | 5 |  |  | 18 | -12.2 | y |  | 5.5 | none | 18 | none | 4.5
2025-01-29 | 29 | 1621 | ALW | - |  |  | 0.605 | 0.95234 | 5 |  |  | 30 | -12.2 | n |  | 5.5 | none | 18 | none | 4.5 |  | Too windy to go to roof
2025-01-30 | 30 | 1302 | ALW | rm lf |  |  | 0.598 | 0.95047 | 5 |  |  | 24 | -12.2 | n |  | 5.5 | none | 18 | none | 4.5 | pass
2025-01-31 | 31 | 1329 | ALW | rm lf | * |  | 0.597 | 0.94861 | 5 |  |  | 21 | -12.2 | n |  | 5.5 | none | 18 | none | 4.5
2025-02-01 | 32 | - | ALW | - |  |  | - | - | - |  |  | - | - | - | - | - | - | - | - | - | - | Con 1- did not go to AWO
2025-02-02 | 33 | 1512 | ALW | rm lf | * |  | 0.587 | 0.94299 | 5 |  |  | 26 | -12.2 | n |  | 5.7 | none | 18 | none | 4.5
2025-02-03 | 34 | 1416 | ALW | rm lf | * |  | 0.589 | 0.94058 | 5 |  |  | 20 | -12.2 | y | TSI 0.8 % error | 5.8 | none | 18 | none | 4.5
2025-02-04 | 35 | - | ALW | - | * |  | - | - | - |  |  | - | - | - | - | - | - | - | - | - |  | ICEsat- did not go to AWO
2025-02-05 | 36 | 1623 | ALW | rm lf | * |  | 0.602 | 0.93739 | 5 |  |  | 9 | -12.2 | n |  | 6.1 | none | 18 | none | 4.5
2025-02-06 | 37 | 1143 | ALW | rm lf | * |  | 0.609 | 0.93679 | 5 |  |  | 9 | -12.2 | n |  | 6.1 | none | 18 | none | 4.5
2025-02-07 | 38 | - | ALW | - | * | 5.01 hPa | - | - | - |  |  | - | - | - | - | - | - | - | - | - | - | Con 1- did not go to AWO
2025-02-08 | 39 | 1806 | ALW | rm lf | * | 10.9 | 0.607 | 0.93156 | 5 |  |  | 8 | -12.2 | n | - | 6.2 | none | 18 | none | 4.5 | -
2025-02-09 | 40 | - | ALW | - | * | - | - | - | - |  |  | - | - | - | - | - | - | - | - | - |  | Con1- did not go to AWO
2025-02-10 | 41 | - | ALW | - | * | - | - | - | - |  |  | - | - | - | - | - | - | - | - | - |  | ICEsat- did not go to AWO
2025-02-11 | 42 | 1351 | ALW | rm lf | * | 11.2 | 0.603 | 0.92683 | 5 |  |  | 4 | -12.2 | y |  | 6.3 | none | 18 | none | 4.5 | pass
2025-02-12 | 43 | 1059 | ALW | rm lf | * | 11.3 | 0.61 | 0.92505 | 5 |  |  | 14 | -12.2 | n |  | 6.3 | none | 18 | none | 4.5
2025-02-13 | 44 | 1734 | ALW | rm lf | * | 11.3 | 0.608 | 0.92328 | 5 |  |  | 34 | -12.2 | n |  | 6.2 | none | 18 | none | 4.5
2025-02-14 | 45 | 1422 | AR | rm lf | * | 11.3 | 0.608 | 0.91338 | 5 |  |  | 22 | -12.2 | n |  | 6.2 | none | 18 | none | 4.5 | pass
2025-02-15 | 46 | - | - | - | * | - | - | - | - |  |  | - | - | - | - | - | - | - | - | - | - | ICEsat- did not go to AWO. Carbon vane pump failure at 1238Z. Pump was restarted at 1704Z after replacing carbon vanes.
2025-02-16 | 47 | 1606 | AR | rm lf | * | 9.497 | 0.601 | 0.90877 | 5 |  |  | 36 | -12.2 | n |  | 6.2 | none | 18 | none | 4.5 | pass
2025-02-17 | 48 | 1345 | AR | rm lf | * | 9.398 | 0.595 | 0.90374 | 5 |  |  | 92 | -12.2 | y | - | 6.2 | none | 18 | none | 4.5 | - | Planned power outage for turnover from 1201-1204Z and 1822-824Z. Instruments supported by UPS.
2025-02-18 | 49 | 1334 | AR | rm lf | * | 9.262 | 0.585 | 0.89212 | 5 |  |  | 88 | -12.2 | n | - | 6.2 | none | 18 | none | 4.5
2025-02-19 | 50 | 1812 | AR | rm lf | * | 9.006 | 0.566 | 0.8012 | 5 |  |  | 99 | -12.2 | n | - | 6.2 | none | 18 | none | 4.5
2025-02-20 | 51 | 1910 | ALW | rm lf | * | 8.89 | 0.603 | 0.79659 | 5 |  |  | 68 | -12.2 | n |  | 6.2 | none | 18 | none | 4.5 |  | Power outage to UPS supported instruments from 0046-1147Z. Forge acquisition restarted at 1530Z. Caused by carbon vane pump failure. Pump moved to non-UPS circuit.
2025-02-21 | 52 | 1448 | AR | rm lf | * | 8.843 | 0.603 | 0.79354 | 5 |  |  | 44 | -12.2 | n | - | 6.2 | none | 18 | none | 4.5 | -
2025-02-22 | 53 | 1414 | AR | rm lf | * | 7.555 | 0.6 | 0.79111 | 5 |  |  | 20 | -12.2 | n | - | 6.5 | none | 16 | none | 4.2 |  | Inlet head removed and cleaned from 1117-1343Z. Pump was unpluged, instruments still adjusting
2025-02-23 | 54 | 1308 | AR | rm lf | * | 8.247 | 0.599 | 0.78977 | 5 |  |  | 20 | -12.2 | n | - | 6.5 | none | 16 | none | 4.2 | -
2025-02-24 | 55 | 1549 | AR | rm lf | * | 8.285 | 0.612 | 0.78797 | 5 |  |  | 21 | -12.2 | y | - | 6.5 | none | 16 | none | 4.4 | pass
2025-02-25 | 56 | - | - | - | * | - | - | - | - |  |  | - | - | - | - | - | - | - | - | - | - | Carbon vane pump failure 0608Z
2025-02-26 | 57 | 1526 | AR | rm lf | * | 6.289 | 0.461 | 0.7865 | 5 |  |  | 0 | -12.2 | n | - | 6 | none | 10 | none | 2.2 | - | Spare pump (carbon vane, Gast model 0211-VI03A-G8C , 30 vlpm) plugged in at 1258Z
2025-02-27 | 58 | 1258 | AR | rm lf |  | 6.281 | 0.467 | 0.78611 | 5 |  |  | 0 | -12.2 | n | - | 6 | none | 10 | none | 2.2 | - | Swapped to spare diaphragm pump at 1633Z due to suboptimal flows from spare cv pump
2025-02-28 | 59 | 1146 | AR | - |  | 9.154 | 0.599 | 0.7843 | 5 |  |  | 5.3 | -12.2 | - | - | 6 | none | - | none | - | - | Con 1- did not go to AWO
2025-03-01 | 60 | 1253 | AR | rm lf |  | 9.13 | 0.603 | 0.78069 | 5 |  |  | 4 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-03-02 | 61 | 1548 | AR | rm lf | * | 9.094 | 0.601 | 0.77806 | 5 |  |  | 1 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | - | MagicCPC wick wetting and reprepping humidifier 1632-1702Z
2025-03-03 | 62 | 1513 | AR | rm lf | * | 9.103 | 0.607 | 0.77601 | 5 |  |  | 9 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-03-04 | 63 | 1535 | AR | rm lf | * | 9.075 | 0.603 | 0.77389 | 5 |  |  | 18 | -12.2 | y |  | 6 | none | 18 | none | 4.5 | pass
2025-03-05 | 64 | 1755 | AR | - | * | 9.19 | 0.61 | 0.77227 | 5 |  |  | 23 | -12.2 | n | - | 6 | none | - | none | - | - | ICEsat- did not go to AWO
2025-03-06 | 65 | 1327 | AR | rm lf | * | 9.244 | 0.602 | 0.77065 | 5 |  |  | 25 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-03-07 | 66 | 1713 | AR | rm lf | * | 9.328 | 0.606 | 0.76036 | 5 |  |  | 22 | -12.2 | n | Ecotech: 2.3% | 6 | none | 18 | none | 4.5 | -
2025-03-08 | 67 | 1137 | AR | rm lf | * | 9.365 | 0.609 | 0.75527 | 5 |  |  | 38 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-03-09 | 68 | 1728 | AR | - | * | 9.454 | 0.614 | 0.75279 | 5 |  |  | 30.4 | -11.4 | n | - | 6 | none | - | none | - | - | Con 1- did not go to AWO
2025-03-10 | 69 | 1230 | AR | rm lf | * | 9.499 | 0.601 | 0.75051 | 5 |  |  | 79 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-03-11 | 70 | 1252 | AR | rm lf | * | 9.554 | 0.605 | 0.74041 | 5 |  |  | 95 | -7.1 | y | - | 6 | none | 18 | none | 4.5 | pass | Power outage today at 2150Z (less than 1 min). Instruments supported by UPS.
2025-03-12 | 71 | 1632 | AR | rm lf | * | 9.478 | 0.601 | 0.73165 | 5 |  |  | 143 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-03-13 | 72 | 1316 | AR | rm lf | * | 9.494 | 0.6 | 0.72665 | 5 |  |  | 113 | -12 | n | - | 6 | none | 18 | none | 4.5 | -
2025-03-14 | 73 | 1157 | AR | rm lf | * | 9.281 | 0.603 | 0.71808 | 5 |  |  | 52 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-03-15 | 74 | 1252 | AR | rm lf | * | 9.325 | 0.605 | 0.70058 | 5 |  |  | 47 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-03-16 | 75 | 1345 | AR | rm lf | * | 9.36 | 0.598 | 0.99127 | 6 |  |  | 67 | -11.5 | n | - | 6 | none | 18 | none | 4.5 | -
2025-03-17 | 76 | 1520 | AR | rm lf | * | 9.527 | 0.596 | 0.98634 | 6 |  |  | 66 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-03-18 | 77 | 1221 | AR | rm lf | * | 9.364 | 0.597 | 0.9802 | 6 |  |  | 128 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-03-19 | 78 | 1730 | AR | rm lf | * | 9.239 | 0.591 | 0.97754 | 6 |  |  | 13 | -12.2 | y | - | 6 | none | 18 | none | 4.5 | pass
2025-03-20 | 79 | 1126 | AR | rm lf | * | 9.161 | 0.601 | 0.97571 | 6 |  |  | 31 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-03-21 | 80 | 1518 | AR | rm lf | * | 9.095 | 0.598 | 0.97218 | 6 |  |  | 40 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-03-22 | 81 | 1225 | AR | rm lf | * | 9.111 | 0.597 | 0.96625 | 6 |  |  | 43 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-03-23 | 82 | 1306 | AR | rm lf | * | 9.188 | 0.605 | 0.95006 | 6 |  |  | 84 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-03-24 | 83 | 1628 | AR | rm lf | * | 9.169 | 0.601 | 0.9338 | 6 |  |  | 209 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-03-25 | 84 | 1234 | AR | rm lf |  | 9.183 | 0.602 | 0.92021 | 6 |  |  | 81 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | pass
2025-03-26 | 85 | 1700 | AR | rm lf | * | 9.164 | 0.601 | 0.90949 | 6 |  |  | 152 | -12.2 | y | - | 6 | none | 18 | none | 4.5 | -
2025-03-27 | 86 | 1631 | AR | rm lf | * | 9.182 | 0.601 | 0.89749 | 6 |  |  | 297 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-03-28 | 87 | 1234 | AR | rm lf | * | 9.199 | 0.602 | 0.88654 | 6 |  |  | 257 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-03-29 | 88 | 1241 | AR | rm lf | * | 9.193 | 0.6 | 0.86936 | 6 |  |  | 163 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-03-30 | 89 | 1232 | AR | rm lf | * | 9.221 | 0.601 | 0.85778 | 6 |  |  | 93 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-03-31 | 90 | 1502 | AR | rm lf | * | 9.132 | 0.601 | 0.84897 | 6 |  |  | 709 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-04-01 | 91 | 959 | AR | rm lf | * | 9.229 | 0.601 | 0.84652 | 6 |  |  | 205 | -12.2 | y | - | 6 | none | 18 | none | 4.5 | pass
2025-04-02 | 92 | 1004 | AR | - | * | 9.296 | 0.605 | 0.83839 | 6 |  |  | 114.5 | -12.2 | n | - | 6 | none | - | none | - | - | ICEsat- did not go to AWO
2025-04-03 | 93 | 1516 | AR | rm lf | * | 9.434 | 0.604 | 0.83076 | 6 |  |  | 65 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-04-04 | 94 | 1527 | AR | rm lf | * | 9.467 | 0.607 | 0.82743 | 6 |  |  | 16 | -10.8 | n | tsi: 2.1% | 6 | none | 18 | none | 4.5 | -
2025-04-05 | 95 | 1058 | AR | rm lf | * | 9.506 | 0.608 | 0.81915 | 6 |  |  | 98 | -10.8 | n | - | 6 | none | 18 | none | 4.5 | -
2025-04-06 | 96 | 1225 | AR | rm lf | * | 9.468 | 0.604 | 0.8102 | 6 |  |  | 54 | -11.1 | n | - | 6 | none | 18 | none | 4.5 | -
2025-04-07 | 97 | 1133 | AR | rm lf | * | 9.461 | 0.604 | 0.80469 | 6 |  |  | 237 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-04-08 | 98 | 1204 | AR | rm lf | * | 9.615 | 0.612 | 0.78796 | 6 |  |  | 183 | -11.4 | y | - | 6 | none | 18 | none | 4.5 | pass
2025-04-09 | 99 | 1047 | AR | rm lf | * | 9.512 | 0.607 | 0.78306 | 6 |  |  | 28 | -9.6 | n | - | 6 | none | 18 | none | 4.5 | -
2025-04-10 | 100 | 1432 | AR | rm lf | * | 9.331 | 0.6 | 0.78169 | 6 |  |  | 28 | -11.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-04-11 | 101 | 1040 | AR | rm lf | * | 9.18 | 0.602 | 0.78073 | 6 |  |  | 46 | -10.8 | n | - | 6 | none | 18 | none | 4.5 | -
2025-04-12 | 102 | 1205 | AR | rm lf | * | 9.228 | 0.609 | 0.77426 | 6 |  |  | 576 | -12 | n | - | 6 | none | 18 | none | 4.5 | -
2025-04-13 | 103 | 1618 | AR | rm lf | * | 9.298 | 0.602 | 0.76123 | 6 |  |  | 138 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-04-14 | 104 | 1522 | AR | rm lf | * | 9.35 | 0.603 | 0.75396 | 6 |  |  | 382 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | - | No Communications error with TSI MFM box (Q11) which usually shows inlet flow, resolved itself by 1640Z
2025-04-15 | 105 | 1106 | AR | rm lf | * | 9.445 | 0.609 | 0.74541 | 6 |  |  | 283 | -12.2 | y | - | 6 | none | 18 | none | 4.5 | pass
2025-04-16 | 106 | 1424 | AR | rm lf | * |  | 0.601 | 0.72959 | 6 |  |  | 411 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | - | No Communications error with TSI MFM box (Q11)
2025-04-17 | 107 | 1050 | AR | clr | * |  | 0.601 | 0.71965 | 6 |  |  | 209 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | - | No Communications error with TSI MFM box (Q11)
2025-04-18 | 108 | 1205 | AR | clr | * | 9.524 | 0.6 | 0.99383 | 7 |  |  | 213 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-04-19 | 109 | 1130 | AR | clr | * | 9.529 | 0.598 | 0.94884 | 7 |  |  | 227 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-04-20 | 110 | 1236 | AR | rm lf | * | 9.51 | 0.596 | 0.92346 | 7 |  |  | 266 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-04-21 | 111 | 1519 | AR | rm lf | * | 9.546 | 0.6 | 0.90213 | 7 |  |  | 91 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-04-22 | 112 | 1511 | AR | rm lf | * | 9.513 | 0.6 | 0.89508 | 7 |  |  | 113 | -11.9 | y | - | 6 | none | 18 | none | 4.5 | pass
2025-04-23 | 113 | 1028 | AR | rm lf | * | 9.586 | 0.598 | 0.88843 | 7 |  |  | 125 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-04-24 | 114 | 1636 | AR | rm lf | * | 9.56 | 0.597 | 0.86264 | 7 |  |  | 172 | -11.3 | n | - | 6 | none | 18 | none | 4.5 | -
2025-04-25 | 115 | 1134 | AR | rm lf | * | 9.523 | 0.595 | 0.84297 | 7 |  |  | 232 | -11.6 | n | - | 6 | none | 18 | none | 4.5 | - | MagicCPC wick wetting and reprepping humidifier 1157-1218Z
2025-04-26 | 116 | 1449 | AR | rm lf | * | 9.385 | 0.592 | 0.82444 | 7 |  |  | 238 | -10.9 | n | - | 6 | none | 18 | none | 4.5 | -
2025-04-27 | 117 | 1151 | AR | rm lf | * | 9.433 | 0.591 | 0.81675 | 7 |  |  | 192 | -11.5 | n | - | 6 | none | 18 | none | 4.5 | -
2025-04-28 | 118 | 1629 | AR | rm lf | * | 9.482 | 0.592 | 0.81154 | 7 |  |  | 245 | -11.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-04-29 | 119 | 1107 | AR | rm lf | * | 9.468 | 0.59 | 0.809 | 7 |  |  | 225 | -11.5 | y | - | 6 | none | 18 | none | 4.5 | pass
2025-04-30 | 120 | 1443 | AR | clr | * | 9.38 | 0.598 | 0.80077 | 7 |  |  | 364 | -11.5 | n | - | 6 | none | 18 | none | 4.5 | -
2025-05-01 | 121 | 1123 | AR | clr | * | 9.378 | 0.596 | 0.79174 | 7 |  |  | 254 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-05-02 | 122 | 1206 | AR | rm lf | * | 9.56 | 0.607 | 0.74109 | 7 |  |  | 284 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-05-03 | 123 | 1025 | AR | rm lf | * | 9.552 | 0.604 | 0.73066 | 7 |  |  | 165 | -10.9 | n | ecotech: 2.5% | 6 | none | 18 | none | 4.5 | -
2025-05-04 | 124 | 1954 | AR | clr | * | 9.529 | 0.602 | 0.71229 | 7 |  |  | 96 | -9.3 | n | - | 6 | none | 18 | none | 4.5 | -
2025-05-05 | 125 | 1128 | AR | clr | * | 9.462 | 0.6 | 0.70561 | 7 |  |  | 135 | -10.9 | y | - | 6 | none | 18 | none | 4.5 | pass
2025-05-06 | 126 | 1306 | AR | rm lf | * | 9.307 | 0.592 | 0.70364 | 7 |  |  | 276 | -11.4 | n | - | 6 | none | 18 | none | 4.5 | -
2025-05-07 | 127 | 1511 | AR | clr | * | 9.151 | 0.598 | 0.99426 | 8 |  |  | 241 | -11.7 | n | - | 6 | none | 18 | none | 4.5 | -
2025-05-08 | 128 | 1236 | AR | clr | * | 9.226 | 0.602 | 0.98759 | 8 |  |  | 564 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | - | MagicCPC wick wetting and reprepping humidifier 1443-1504Z
2025-05-09 | 129 | 1146 | AR | rm lf | * | 9.201 | 0.601 | 0.87045 | 8 |  |  | 188 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-05-10 | 130 | 1044 | AR | rm lf | * | 9.277 | 0.604 | 0.85984 | 8 |  |  | 271 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | - | Planned station-wide power outage from 1008-1014Z. Instruments supported by UPS.
2025-05-11 | 131 | 1549 | AR | rm lf | * | 9.239 | 0.602 | 0.85275 | 8 |  |  | 113 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | - | Unplanned station-wide power outage from 1455-1524Z. Instruments supported by UPS.
2025-05-12 | 132 | 1112 | AR | rm lf | * | 9.328 | 0.61 | 0.8369 | 8 |  |  | 91 | -12.2 | y | - | 6 | none | 18 | none | 4.5 | pass
2025-05-13 | 133 | 1813 | AR | - | * | 9.396 | 0.614 | 0.81618 | 8 |  |  | 50.2 | -10 | n | - | 6 | - | - | - | - | - | Con 1- did not go to AWO. Unplanned station power outage from 2031-2040Z. Controlled station power outage from 2105-2110Z. Instruments supported by UPS.
2025-05-14 | 134 | 1005 | AR | clr | * | 9.48 | 0.603 | 0.81515 | 8 |  |  | 163 | -10.5 | n | - | 6 | none | 18 | none | 4.5 | -
2025-05-15 | 135 | 1521 | AR | clr | * | 9.613 | 0.6 | 0.80329 | 8 |  |  | 199 | -10.4 | n | - | 6 | none | 18 | none | 4.5 | -
2025-05-16 | 136 | 1040 | AR | clr | * | 9.695 | 0.604 | 0.7858 | 8 |  |  | 162 | -9.8 | n | - | 6 | none | 18 | none | 4.5 | -
2025-05-17 | 137 | 946 | AR | clr | * | 9.706 | 0.604 | 0.77195 | 8 |  |  | 492 | -9 | n | - | 6 | none | 18 | none | 4.5 | - | Planned station power outage from 1015-1019Z. Instruments supported by UPS. Changed CLAP fliter 1523-1541Z.
2025-05-18 | 138 | 1153 | AR | clr | * | 9.753 | 0.608 | 0.9899 | 1 |  |  | 326 | -9.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-05-19 | 139 | 1552 | AR | clr | * | 9.857 | 0.6 | 0.97439 | 1 |  |  | 247 | -7.1 | n | - | 6 | none | 18 | none | 4.5 | -
2025-05-20 | 140 | 1019 | AR | rm lf | * | 9.871 | 0.6 | 0.95456 | 1 |  |  | 228 | -9.5 | y | - | 6 | none | 18 | none | 4.5 | -
2025-05-21 | 141 | 1439 | AR | clr | * | 9.591 | 0.598 | 0.93058 | 1 |  |  | 233 | -10.9 | n | - | 6 | none | 18 | none | 4.5 | pass
2025-05-22 | 142 | 1007 | AR | clr | * | 9.489 | 0.597 | 0.90788 | 1 |  |  | 199 | -10.7 | n | - | 6 | none | 18 | none | 4.5 | -
2025-05-23 | 143 | 1233 | AR | rm lf | * | 9.383 | 0.591 | 0.8836 | 1 |  |  | 193 | -10.5 | n | - | 6 | none | 18 | none | 4.5 | -
2025-05-24 | 144 | 1042 | AR | clr | * | 9.297 | 0.601 | 0.87626 | 1 |  |  | 161 | -11.9 | n | - | 6 | none | 18 | none | 4.5 | - | Unplanned station-wide power outage from 1806-1808Z. Instruments supported by UPS.
2025-05-25 | 145 | 1307 | AR | clr | * | 9.399 | 0.607 | 0.85278 | 1 |  |  | 171 | -12.2 | n | - | 6 | none | 18 | none | 4.5 | -
2025-05-26 | 146 | 1027 | AR | clr | * | 9.374 | 0.604 | 0.84013 | 1 |  |  | 675 | -11.1 | n | - | 6 | none | 18 | none | 4.5 | -
2025-05-27 | 147 | 1500 | AR | clr | * | 9.326 | 0.604 | 0.83066 | 1 |  |  | 659 | -11.8 | y | - | 6 | none | 18 | none | 4.5 | pass
2025-05-28 | 148 | 1133 | AR | rm lf | * | 9.396 | 0.605 | 0.82484 | 1 |  |  | 299 | -11.6 | n | - | 6 | none | 18 | none | 4.5 | -
2025-05-29 | 149 | 1454 | AR | clr | * | 9.383 | 0.605 | 0.81583 | 1 |  |  | 190 | -11.4 | n | - | 6 | none | 18 | none | 4.5 | -
2025-05-30 | 150 | 1206 | AR | clr | * | 9.42 | 0.605 | 0.80727 | 1 |  |  | 175 | -11.5 | n | - | 6 | none | 18 | none | 4.5 | -
2025-05-31 | 151 | 1553 | AR | clr | * | 9.439 | 0.607 | 0.78626 | 1 |  |  | 121 | -11.1 | n | - | 6 | none | 18 | none | 4.5 | -
2025-06-01 | 152 | 1800 | AR | clr | * | 9.363 | 0.605 | 0.77439 | 1 |  |  | 137 | -10.5 | n | - | 6 | none | 18 | none | 4.5 | -
2025-06-02 | 153 | 1555 | AR | clr | * | 9.395 | 0.609 | 0.76782 | 1 |  |  | 310 | -11.6 | n | - | 6 | none | 18 | none | 4.5 | -
2025-06-03 | 154 | 1055 | AR | clr | * | 9.466 | 0.6 | 0.76029 | 1 |  |  | 152 | -11.6 | y | - | 6 | none | 18 | none | 4.5 | pass
2025-06-04 | 155 | 1612 | AR | - | * | 9.362 | 0.593 | 0.75046 | 1 |  |  | 98.2 | -11.6 | n | - | 5.9 | none | - | none | - | - | Icesat-did not go to AWO
2025-06-05 | 156 | 1036 | AR | clr | * | 9.33 | 0.59 | 0.7447 | 1 |  |  | 187 | -11.6 | n | - | 5.9 | none | 18 | none | 4.5 | - | MagicCPC wick wetting and humidifier re-prepping 1639-1700Z
2025-06-06 | 157 | 1542 | AR | clr | * | 9.316 | 0.591 | 0.73698 | 1 |  |  | 174 | -11.2 | n | - | 5.9 | none | 18 | none | 4.5 | -
2025-06-07 | 158 | 1055 | AR | rm lf | * | 9.393 | 0.594 | 0.73002 | 1 |  |  | 166 | -12 | n | - | 5.9 | none | 18 | none | 4.5 | -
2025-06-08 | 159 | 1354 | AR | clr | * | 9.438 | 0.597 | 0.71971 | 1 |  |  | 125 | -11.9 | n | - | 5.9 | none | 18 | none | 4.5 | -
2025-06-09 | 160 | 1122 | AR | clr | * | 9.515 | 0.603 | 0.71104 | 1 |  |  | 152 | -11.6 | y | - | 5.9 | none | 18 | none | 4.5 | - | Planned power outage 1001-1005Z. Instruments supported by UPS.
2025-06-10 | 161 | 1439 | AR | rm lf | * | 9.477 | 0.601 | 0.70081 | 1 |  |  | 194 | -11.4 | n | - | 5.9 | none | 18 | none | 4.5 | - | Planned power outage 1704-1707Z. Instruments supported by UPS.
2025-06-11 | 162 | 1454 | AR | clr | * | 9.441 | 0.593 | 0.9361 | 2 |  |  | 202 | -11.4 | n | - | 5.9 | none | 18 | none | 4.5 | pass
2025-06-12 | 163 | 1003 | PC | clr | * | 9.449 | 0.601 | 0.84964 | 2 |  |  | 188 | -11.7 | n | - | 5.9 | none | 18 | none | 4.5 | -
2025-06-13 | 164 | 1004 | PC | clr | * | 9.479 | 0.6 | 0.77572 | 2 |  |  | 182 | -10.2 | n | TSI: 1.2% | 5.9 | none | 18 | none | 4.5 | - | Planned power outage 1037-1040Z. Second planned power outage 1412-1416Z. Instruments supported by UPS.
2025-06-14 | 165 | 1555 | PC | clr | * | 9.572 | 0.606 | 0.75529 | 2 |  |  | 177 | -7.6 | n | - | 5.9 | none | 18 | none | 4.5 | -
2025-06-15 | 166 | 1121 | PC | clr | * | 9.639 | 0.607 | 0.73011 | 2 |  |  | 175 | -8.2 | y | - | 6 | none | 18 | none | 4.5 | -
2025-06-16 | 167 |  | PC | clr | * |  |  |  | 2 |  |  |  |  | n | - | 5.9 | none | 18 | none | 4.5 | - | Issued with spreadsheet- deletes last day or two of data no matter how many times I save. Filled in what I remembered from yesterday.
2025-06-17 | 168 | 1651 | PC | rm lf | * | 9.495 | 0.622 | 0.97644 | 3 |  |  | 164 | -8.1 | n | - | 5.9 | none | 18 | none | 4.5 | -
2025-06-18 | 169 | 1224 | PC | clr | * | 9.518 | 0.623 | 0.91951 | 3 |  |  | 117 | -9.6 | y | - | 5.9 | none | 18 | none | 4.5 | - | Adjusted CLAP flow to 0.60.
2025-06-19 | 170 | 1049 | PC | rm lf | * | 9.628 | 0.559 | 0.86857 | 3 |  |  | 131 | -9.8 | n | - | 5.9 | none | 18 | none | 4.5 | - | Wick wetting procedure as recommended in Forge email 1611-1641Z. Magic CPC consumed 0.5mL water. Prepared humidifier 1641Z and verified no remaining air bubbles.
2025-06-20 | 171 | 1445 | PC | clr | * | 9.494 | 0.559 | 0.83662 | 3 |  |  | 188 | -9.3 | n | - | 5.9 | none | 18 | none | 4.5 | -
2025-06-21 | 172 | 1627 | PC | clr | * | 9.479 | 0.555 | 0.78592 | 3 |  |  | 162 | -9 | n | - | 5.9 | none | 18 | none | 4.5 | - | Adjusted CLAP flow from 0.59 to 0.61.
2025-06-22 | 173 | 1534 | PC | rm lf | * | 9.398 | 0.558 | 0.74657 | 3 |  |  | 101 | -10.4 | y | - | 6 | none | 18 | none | 4.5 | -
2025-06-23 | 174 | 1135 | PC | clr | * | 9.491 | 0.559 | 0.7231 | 3 |  |  | 133 | -10.3 | n | - | 5.9 | none | 18 | none | 4.5 | - | Adjusted CLAP flow from 0.60 to 0.61.
2025-06-24 | 175 | 905 | PC | rm lf | * | 9.514 | 0.568 | 0.70998 | 3 |  |  | 109 | -10.5 | n | - | 5.9 | none | 18 | none | 4.5 | - | Plane flew a big low contrail directly over station 0905.
2025-06-25 | 176 | 1123 | PC | rm lf | * | 9.459 | 0.564 | 0.94803 | 4 |  |  | 151 | -10.4 | y | - | 5.9 | none | 18 | none | 4.5 | - | Low winds equipment operations: 11:42-12:07Z- D6, 953, skidsteer, snowmo, welder (machines operating 11:42-11:49, welder missed radio call and operated 11:42-1207Z).
2025-06-26 | 177 | 940 | PC | rm lf | * | 9.368 | 0.555 | 0.84201 | 4 |  |  | 126 | -10.5 | n | - | 5.9 | none | 18 | none | 4.5 | -
2025-06-27 | 178 | 1422 | PC | clr | * | 9.349 | 0.564 | 0.82554 | 4 |  |  | 24 | -9.8 | n | - | 5.9 | none | 18 | none | 4.5 | -
2025-06-28 | 179 | 924 | PC | clr | * | 9.384 | 0.565 | 0.81564 | 4 |  |  | 57 | -9.9 | y | - | 5.9 | none | 18 | none | 4.5 | - | 1005-1006Z: Planned power outage for mechanics training. No data impacted for instruments supported by UPS. 1010-1012Z: second planned power outage to switch back to main generator.
2025-06-29 | 180 | 1522 | PC | clr | * | 9.366 | 0.567 | 0.80616 | 4 |  |  | 1556 | -10.2 | n | - | 5.9 | none | 18 | none | 4.5 | -
2025-06-30 | 181 | 1401 | PC | clr | * | 9.41 | 0.567 | 0.7776 | 4 |  |  | 274 | -10.7 | n | - | 5.9 | none | 18 | none | 4.5 | pass
2025-07-01 | 182 | 958 | PC | clr | * | 9.348 | 0.563 | 0.75999 | 4 |  |  | 132 | -10.4 | y | - | 5.9 | none | 18 | none | 4.5 | - | Planned power outage due to generator issues last night, power swap from 1133-1135Z, again from 1224-1225Z. And again 1602-1604Z. Instruments supported by UPS.
2025-07-02 | 183 | 1006 | PC | rm lf | * | 9.308 | 0.563 | 0.7476 | 4 |  |  | 131 | -9.2 | n | - | 5.9 | none | 18 | none | 4.5 | -
2025-07-03 | 184 | 1128 | PC | clr | * | 9.345 | 0.571 | 0.73679 | 4 |  |  | 157 | -10.5 | n | - | 5.9 | none | 18 | none | 4.5 | -
2025-07-04 | 185 | 1056 | PC | clr | * | 9.326 | 0.569 | 0.72215 | 4 |  |  | 167 | -9.3 | y | - | 5.9 | none | 18 | none | 4.5 | -
2025-07-05 | 186 | 1631 | PC | clr | * | 9.355 | 0.57 | 0.70743 | 4 |  |  | 162 | -11.6 | n | - | 5.9 | none | 18 | none | 4.5 | - | Northwinds operations: 1716-1735Z D6 Snow machine 953 953
2025-07-06 | 187 | 947 | PC | clr | * | 9.33 | 0.569 | 0.70358 | 4 |  |  | 133 | -12 | n | - | 5.9 | none | 18 | none | 4.5 | -
2025-07-07 | 188 | 1119 | PC | clr | * | 9.342 | 0.588 | 0.9877 | 5 |  |  | 163 | -8.2 | y | - | 5.9 | none | 18 | none | 4.5 | -
2025-07-08 | 189 | 1011 | PC | clr | * | 9.335 | 0.585 | 0.92966 | 5 |  |  | 273 | -8.1 | n | Ecotech: 1.8% | 5.9 | none | 18 | none | 4.5 | pass | 953
2025-07-09 | 190 | 1043 | PC | rm lf | * | 9.275 | 0.565 | 0.90368 | 5 |  |  | 148 | -9.2 | n | - | 5.9 | none | 18 | none | 4.5 | - | LOW WINDS OPERATIONS 1015-1018Z: D6 HEAT CANNON, 5K
2025-07-10 | 191 | 1007 | PC | clr | * | 9.253 | 0.563 | 0.88678 | 5 |  |  | 314 | -9.1 | y | - | 5.9 | none | 18 | none | 4.5 | -
2025-07-11 | 192 | 1101 | PC | rm lf | * | 9.423 | 0.572 | 0.81116 | 5 |  |  | 73 | -11.3 | n | - | 5.9 | none | 18 | none | 4.5 | -
2025-07-12 | 193 | 1028 | PC | rm lf | * | 9.471 | 0.574 | 0.80338 | 5 |  |  | 157 | -9 | n | - | 5.9 | none | 18 | none | 4.5 | -
2025-07-13 | 194 | 1229 | PC | rm ice | * | 9.506 | 0.562 | 0.75706 | 5 |  |  | 83 | -6.2 | y | - | 5.9 | none | 18 | none | 4.5 | -
2025-07-14 | 195 | 1039 | PC | clr | * | 9.628 | 0.565 | 0.7352 | 5 |  |  | 270 | -6.9 | n | - | 5.9 | none | 18 | none | 4.5 | -
2025-07-15 | 196 | 1031 | PC | rm ice | * | 9.649 | 0.566 | 0.70804 | 5 |  |  | 268 | -6.8 | n | - | 5.9 | none | 18 | none | 4.5 | pass | Northwind ops: x2 variances (welding, medical to heat isolation bay during day) from 0900-1400
2025-07-16 | 197 | 1047 | PC | rm lf | * | 9.604 | 0.558 | 0.96907 | 6 |  |  | 154 | -8.2 | y | - | 5.8 | none | 18 | none | 4.5 | -
2025-07-17 | 198 | 1148 | PC | rm lf | * | 9.592 | 0.556 | 0.93373 | 6 |  |  | 170 | -8.9 | n | - | 5.9 | none | 18 | none | 4.5 | -
2025-07-18 | 199 | 1009 | PC | clr | * | 9.614 | 0.554 | 0.89748 | 6 |  |  | 151 | -6 | y | - | 5.8 | none | 18 | none | 4.5 | -
2025-07-19 | 200 | 1019 | PC | rm ice | * | 10.469 | 0.553 | 0.83893 | 6 |  |  | 3020 | -5.6 | n | - | 5.9 | none | 18 | none | 4.5 | -
2025-07-20 | 201 | 1139 | PC | clr | * | 10.974 | 0.559 | 0.74304 | 6 |  |  | 567 | -7 | n | - | 5.9 | none | 18 | none | 4.5 | - | Adjusted Ecotech flow from 4.2 to 4.5
2025-07-21 | 202 | 1034 | PC | clr | * | 11.028 | 0.572 | 0.98076 | 7 |  |  | 602 | -7.6 | y | - | 5.8 | none | 18 | none | 4.5 | -
2025-07-22 | 203 | 1044 | PC | rm lf | * | 11.328 | 0.574 | 0.92892 | 7 |  |  | 399 | -7 | n | - | 5.8 | none | 18 | none | 4.5 | pass
2025-07-23 | 204 | 1017 | PC | clr | * | 11.408 | 0.573 | 0.89673 | 7 |  |  | 387 | -6.6 | n | - | 5.9 | none | 18 | none | 4.5 | - | Low winds from 0943-1226Z, 1803-1908Z, North winds 1206-1217Z, 1851-1942Z, 2052-2150Z
2025-07-24 | 205 | 1025 | PC | clr | * | 11.279 | 0.565 | 0.84914 | 7 |  |  | 250 | -7.7 | n | - | 5.9 | none | 18 | none | 4.5 | - | Low winds from 0823-1429Z, 1643-1643Z. North winds from 0051-0402Z, 0932-1230Z
2025-07-25 | 206 | 1013 | PC | clr | * | 11.245 | 0.563 | 0.80429 | 7 |  |  | 166 | -8.8 | y | - | 5.8 | none | 18 | none | 4.5 | -
2025-07-26 | 207 | 1202 | PC | clr | * | 11.292 | 0.561 | 0.76984 | 7 |  |  | 214 | -7.8 | n | - | 5.8 | none | 18 | none | 4.5 | - | Low winds from 0223-0224Z. North winds from 2314-0054Z
2025-07-27 | 208 | 1857 | PC | clr | * | 11.341 | 0.564 | 0.73609 | 7 |  |  | 118 | -8.2 | n | - | 5.8 | none | 18 | none | 4.5 | - | Low winds from 0520-0520Z, 0722-0741Z, 2147-2357Z. North winds from 0239-0520Z, 2202-2340Z
2025-07-28 | 209 | 1035 | PC | clr | * | 11.303 | 0.564 | 0.72039 | 7 |  |  | 41 | -8.1 | y | - | 5.8 | none | 18 | none | 4.5 | -
2025-07-29 | 210 | 1046 | PC | clr | * | 11.259 | 0.551 | 0.98071 | 8 |  |  | 225 | -8.4 | n | - | 5.8 | none | 18 | none | 4.5 | - | Low winds 09:17Z to 09:56Z
2025-07-30 | 211 | 1026 | PC | clr | * | 11.296 | 0.552 | 0.93229 | 8 |  |  | 179 | -8.8 | n | - | 5.8 | none | 18 | none | 4.5 | pass | Low winds 02:10Z to 02:11Z, north winds 09:19Z to 09:25Z 11:47Z to 13:28Z 19:54Z to 21:04Z
2025-07-31 | 212 | 1524 | PC | clr | * | 11.334 | 0.563 | 0.89894 | 8 |  |  | 181 | -9.7 | y | - | 5.8 | none | 18 | none | 4.5 | -
2025-08-01 | 213 | 1201 | PC | clr | * | 11.251 | 0.557 | 0.88172 | 8 |  |  | 123 | -10.4 | n | - | 5.8 | none | 18 | none | 4.5 | -
2025-08-02 | 214 | 1452 | PC | clr | * | 11.161 | 0.552 | 0.87136 | 8 |  |  | 406 | -9 | n | - | 5.8 | none | 18 | none | 4.5 | -
2025-08-03 | 215 | 1125 | PC | clr | * | 11.08 | 0.548 | 0.86312 | 8 |  |  | 67 | -9.6 | n | - | 5.8 | none | 18 | none | 4.5 | -
2025-08-04 | 216 | 1503 | PC | clr | * | 11.157 | 0.554 | 0.83961 | 8 |  |  | 301 | -10.3 | y | - | 5.8 | none | 18 | none | 4.5 | -
2025-08-05 | 217 | 958 | PC | rm lf | * | 11.17 | 0.554 | 0.83444 | 8 |  |  | 664 | -9.8 | n | - | 5.8 | none | 18 | none | 4.5 | -
2025-08-06 | 218 | 1127 | PC | clr | * | 11.161 | 0.553 | 0.82006 | 8 |  |  | 390 | -10.1 | n | - | 5.8 | none | 18 | none | 4.5 | -
2025-08-07 | 219 | 1058 | PC | clr | * | 11.173 | 0.554 | 0.75016 | 8 |  |  | 217 | -12.1 | N | - | 5.8 | none | 18 | none | 4.5 | - | Changed CLAP filter 1154Z. Adjusted flow to 0.608 lpm.
2025-08-08 | 220 | 1006 | PC | RM LF | * | 11.187 | 0.608 | 0.98536 | 1 |  |  | 274 | -9.4 | Y | TSI 44% | 5.8 | none | 18 | none | 4.5 | - | Failed spancheck twice, emailed Erin who advised to tighten CO2 valve and try again.
2025-08-09 | 221 | 1226 | PC | clr | * | 11.157 | 0.611 | 0.95084 | 1 |  |  | 4630 | -10.1 | N | TSI 0.7% | 5.8 | none | 18 | none | 4.5 | - | Tightened CO2 valve, spancheck passed
2025-08-10 | 222 | 1404 | PC | clr | * | 11.206 | 0.61 | 0.84849 | 1 |  |  | 46 | -9.01 | N | - | 5.8 | none | 18 | none | 4.5 | -
2025-08-11 | 223 | 1108 | PC | clr | * | 11.236 | 0.602 | 0.84227 | 1 |  |  | 47 | -10.1 | y | - | 5.8 | none | 18 | none | 4.5 | -
2025-08-12 | 224 | 1503 | PC | clr | * | 11.179 | 0.601 | 0.83143 | 1 |  |  | 172 | -10.3 | n | - | 5.8 | none | 18 | none | 4.5 | pass | Adjusted CLAP flow from 0.598 to 0.601
2025-08-13 | 225 | 1202 | PC | rm lf | * | 11.193 | 0.602 | 0.82458 | 1 |  |  | 71 | -10.6 | n | - | 5.8 | none | 18 | none | 4.5 | -
2025-08-14 | 226 | 1022 | PC | rm lf | * | 11.269 | 0.601 | 0.81971 | 1 |  |  | 104 | -10.9 | n | - | 5.8 | none | 18 | none | 4.5 | -
2025-08-15 | 227 | 909 | PC | clr | * | 11.206 | 0.602 | 0.81342 | 1 |  |  | 105 | -11.5 | y | - | 5.8 | none | 18 | none | 4.5 | -
2025-08-16 | 228 | 1032 | PC | clr | * | 11.182 | 0.601 | 0.80564 | 1 |  |  | 137 | -10.3 | n | - | 5.8 | none | 18 | none | 4.5 | -
2025-08-17 | 229 | 1429 | PC | clr | * | 11.395 | 0.601 | 0.79974 | 1 |  |  | 118 | -11.3 | n | - | 5.8 | none | 18 | none | 4.5 | -
2025-08-18 | 230 | 1007 | PC | clr | * | 11.546 | 0.602 | 0.79348 | 1 |  |  | 204 | -9.2 | n | - | 5.8 | none | 18 | none | 4.5 | -
2025-08-19 | 231 | 1126 | PC | rm lf | * | 11.741 | 0.601 | 0.78718 | 1 |  |  | 38 | -6.6 | y | - | 5.7 | none | 18 | none | 4.5 | -
2025-08-20 | 232 | 1210 | PC | rm lf | * | 11.719 | 0.603 | 0.78249 | 1 |  |  | 287 | -8.9 | n | - | 5.7 | none | 18 | none | 4.5 | - | Magic CPC wick rewetting completed at 1508Z. Humidifier prepped, several rounds of bubbles removed.
2025-08-21 | 233 | 1052 | PC | rm lf | * | 11.644 | 0.602 | 0.87322 | 2 |  |  | 225 | -6.3 | n | - | 5.7 | none | 18 | none | 4.5 | -
2025-08-22 | 234 | 1056 | PC | clr | * | 11.481 | 0.603 | 0.76262 | 2 |  |  | 154 | -7.5 | n | - | 5.7 | none | 18 | none | 4.5 | -
2025-08-23 | 235 | 1624 | PC | clr | * | 11.335 | 0.6 | 0.70375 | 2 |  |  | 231 | -7.8 | n | - | 5.7 | none | 18 | none | 4.5 | -
2025-08-24 | 236 | 1017 | PC | clr | * | 11.294 | 0.601 | 0.94925 | 3 |  |  | 75 | -7.7 | n | - | 5.7 | none | 18 | none | 4.5 | -
2025-08-25 | 237 | 1300 | PC | clr | * | 11.191 | 0.6 | 0.89557 | 3 |  |  | 450 | -9.2 | y | - | 5.7 | none | 18 | none | 4.5 | - | Planned power swap, instruments supported by ups 1002-1004Z
2025-08-26 | 238 | 1028 | PC | rm lf | * | 11.326 | 0.602 | 0.86549 | 3 |  |  | 87 | -9.6 | n | - | 5.7 | none | 18 | none | 4.5 | - | Planned power swap, instruments supported by ups 1101-1102Z
2025-08-27 | 239 | 959 | pc | clr | * | 11.336 | 0.602 | 0.84733 | 3 |  |  | 85 | -8.7 | n | - | 5.7 | none | 18 | none | 4.5 | -
2025-08-28 | 240 | 1052 | PC | clr | * | 11.205 | 0.601 | 0.84297 | 3 |  |  | 46 | -9 | n | - | 5.7 | none | 18 | none | 4.5 | - | Magic wick re-wet, humidifier prepped, and nozzle cleaned 1115-1204Z
2025-08-29 | 241 | 1141 | PC | rm lf | * | 11.23 | 0.599 | 0.83649 | 3 |  |  | 86 | -12 | n | - | 5.9 | none | 18 | none | 4.5 | -
2025-08-30 | 242 | 1033 | PC | clr | * | 11.246 | 0.599 | 0.82742 | 3 |  |  | 244 | -10 | n | - | 5.9 | none | 18 | none | 4.5 | -
2025-08-31 | 243 | 1313 | PC | clr | * | 11.344 | 0.602 | 0.91727 | 3 |  |  | 100 | -8.8 | n | - | 6 | none | 18 | none | 4.5 | -
2025-09-01 | 244 | 1058 | PC | RM LF | * | 11.329 | 0.604 | 0.81389 | 3 |  |  | 124 | -10.1 | n | - | 5.9 | none | 18 | none | 4.5 | pass
2025-09-02 | 245 | 1110 | PC | RM LF | * | 11.257 | 0.6 | 0.80716 | 3 |  |  | 85 | -10.7 | y | - | 6.2 | none | 18 | none | 4.5 | -
2025-09-03 | 246 | 1023 | PC | CLR | * | 11.223 | 0.598 | 0.80177 | 3 |  |  | 150 | -11.2 | n | - | 6.2 | none | 18 | none | 4.5 | -
2025-09-04 | 247 | 1058 | PC | CLR | * | 11.278 | 0.601 | 0.79439 | 3 |  |  | 213 | -10.9 | N | - | 6.2 | none | 18 | none | 4.5 | -
2025-09-05 | 248 | 1234 | PC | CLR | * | 11.352 | 0.601 | 0.78396 | 3 |  |  | 176 | -10.3 | N | Ecotech 1.9% | 6.2 | none | 18 | none | 4.5 | - | Ecotech spancheck complete, error 1.9%
2025-09-06 | 249 | 1037 | PC | CLR | * | 11.368 | 0.602 | 0.77825 | 3 |  |  | 198 | -9.8 | N | - | 6.2 | none | 18 | none | 4.5 | -
2025-09-07 | 250 | 1445 | PC | RM LF | * | 11.284 | 0.601 | 0.76858 | 3 |  |  | 141 | -11.7 | Y | - | 6.2 | none | 18 | none | 4.5 | -
2025-09-08 | 251 | 1208 | PC | CLR | * | 11.267 | 0.601 | 0.76152 | 3 |  |  | 144 | -12 | N | - | 6.2 | none | 18 | none | 4.5 | -
2025-09-09 | 252 | 1047 | PC | CLR | * | 11.101 | 0.6 | 0.7554 | 3 |  |  | 52 | -10.2 | N | - | 6.2 | none | 18 | none | 4.5 | -
2025-09-10 | 253 | 1454 | PC | CLR | * | 11.189 | 0.601 | 0.75404 | 3 |  |  | 51 | -9.9 | Y | - | 6.2 | none | 18 | none | 4.5 | -
2025-09-11 | 254 | 1042 | PC | CLR | * | 11.318 | 0.601 | 0.75373 | 3 |  |  | 876 | -10.3 | N | - | 6.2 | none | 18 | none | 4.5 | -
2025-09-12 | 255 | 1128 | PC | RM LF | * | 11.393 | 0.6 | 0.75324 | 3 |  |  | 73 | -10 | N | - | 6.2 | none | 18 | none | 4.5 | -
2025-09-13 | 256 | 1104 | PC | CLR | * | 11.358 | 0.601 | 0.75116 | 3 |  |  | 88 | -9 | N | - | 6.2 | none | 18 | none | 4.5 | -
2025-09-14 | 257 | 1527 | PC | CLR | * | 11.318 | 0.599 | 0.74734 | 3 |  |  | 75 | -8.9 | n | - | 6.2 | none | 18 | none | 4.5 | -
2025-09-15 | 258 | 1109 | PC | RM LF | * | 11.373 | 0.602 | 0.74315 | 3 |  |  | 95 | -10.1 | y | - | 6.1 | none | 18 | none | 4.5 | -
2025-09-16 | 259 | 1147 | PC | RM LF | * | 11.412 | 0.601 | 0.7401 | 3 |  |  | 93 | -10.8 | n | - | 6.1 | none | 18 | none | 4.5 | -
2025-09-17 | 260 | 1448 | PC | CLR | * | 11.457 | 0.601 | 0.7386 | 3 |  |  | 127 | -10.4 | n | - | 6.1 | none | 18 | none | 4.5 | -
2025-09-18 | 261 | 1001 | PC | RM LF | * | 11.447 | 0.6 | 0.73675 | 3 |  |  | 133 | -10.6 | Y | - | 6.1 | none | 18 | none | 4.5 | PASS
2025-09-19 | 262 | 1128 | PC | RM LF | * | 11.459 | 0.601 | 0.72812 | 3 |  |  | 101 | -11.4 | N | - | 6.1 | none | 18 | none | 4.5 | -
2025-09-20 | 263 | 1118 | PC | RM LF | * | 11.377 | 0.601 | 0.72293 | 3 |  |  | 119 | -11.5 | Y | - | 6.1 | none | 18 | none | 4.5 | -
2025-09-21 | 264 | 1143 | PC | CLR | * | 11.271 | 0.601 | 0.72062 | 3 |  |  | 88 | -11.7 | N | - | 6.1 | none | 18 | none | 4.5 | -
2025-09-22 | 265 | 1230 | PC | rm lf | * | 11.264 | 0.6 | 0.71871 | 3 |  |  | 114 | -8.5 | n | - | 6.1 | none | 18 | none | 4.5 | -
2025-09-23 | 266 | 1127 | PC | rm lf | * | 11.238 | 0.598 | 0.71649 | 3 |  |  | 87 | -9.7 | y | - | 6.1 | none | 18 | none | 4.5 | -
2025-09-24 | 267 | 1138 | PC | rm lf | * | 11.26 | 0.602 | 0.71617 | 3 |  |  | 17 | -10.7 | N | - | 6.1 | none | 18 | none | 4.5 | -
2025-09-25 | 268 | 1120 | PC | RM LF | * | 11.085 | 0.601 | 0.70766 | 3 |  |  | 61 | -11.4 | N | - | 6.1 | none | 18 | none | 4.5 | -
2025-09-26 | 269 | 1134 | PC | RM LF | * | 10.974 | 0.599 | 0.70542 | 3 |  |  | 42 | -12.2 | Y | - | 6.3 | none | 18 | none | 4.5 | - | Magic wick re-wet, humidifier prepped, and nozzle cleaned 1203-1227Z
2025-09-27 | 270 | 1212 | PC | RM LF | * | 11.01 | 0.601 | 0.98102 | 4 |  |  | 110 | -12.2 | N | - | 6.3 | none | 18 | none | 4.5 | -
2025-09-28 | 271 | 1451 | PC | RM LF | * | 11.076 | 0.602 | 0.97521 | 4 |  |  | 99 | -12.2 | N | - | 6.3 | none | 18 | none | 4.5 | -
2025-09-29 | 272 | 1247 | PC | RM LF | * | 11.135 | 0.601 | 0.97357 | 4 |  |  | 31 | -12.2 | Y | - | 6.6 | none | 18 | none | 4.5 | -
2025-09-30 | 273 | 1417 | PC | CLR | * | 11.068 | 0.6 | 0.97013 | 4 |  |  | 120 | -11.8 | n | - | 6.7 | none | 18 | none | 4.5 | -
2025-10-01 | 274 | 1417 | PC | CLR | * | 11.002 | 0.601 | 0.95843 | 4 |  |  | 47 | -12.2 | N | - | 6.7 | none | 18 | none | 4.5 | PASS
2025-10-02 | 275 | 1200 | PC | RM LF | * | 10.948 | 0.601 | 0.95781 | 4 |  |  | 67 | -11.1 | N | - | 6.7 | none | 18 | none | 4.5 | -
2025-10-03 | 276 | 1119 | PC | RM LF | * | 10.914 | 0.601 | 0.95689 | 4 |  |  | 76 | -12.2 | y | - | 6.6 | none | 18 | none | 4.5 | -
2025-10-04 | 277 | 1149 | PC | RM LF | * | 10.942 | 0.601 | 0.95485 | 4 |  |  | 66 | -11.6 | N | TSI 2.8% | 6.7 | none | 18 | none | 4.5 | -
2025-10-05 | 278 | 1619 | PC | RM LF | * | 10.921 | 0.601 | 0.92771 | 4 |  |  | 41 | -12.2 | n | - | 6.7 | none | 18 | none | 4.5 | -
2025-10-06 | 279 | 1222 | PC | RM LF | * | 10.908 | 0.601 | 0.8813 | 4 |  |  | 411 | -12.2 | y | - | 6.6 | none | 18 | none | 4.5 | -
2025-10-07 | 280 | 1143 | PC | RM LF | * | 10.928 | 0.602 | 0.88673 | 4 |  |  | 49 | -12.2 | n | - | 6.6 | none | 18 | none | 4.5 | - | Magic wick re-wet and humidifier prepped 1524-1537Z
2025-10-08 | 281 | 1155 | PC | - | * | 10.917 | 0.597 | 0.88489 | 4 |  |  | 17.8 | -12.2 | - | - | 6.6 | none | - | none | - | - | CON-1 did not go to AWO
2025-10-09 | 282 | 1443 | PC | RM LF | * | 11.058 | 0.601 | 0.88178 | 4 |  |  | 34.2 | -12.2 | y | - | 6.8 | none | 18 | none | 4.5 | -
2025-10-10 | 283 | 1113 | PC | RM LF | * | 11.122 | 0.601 | 0.87414 | 4 |  |  | 22 | -12.2 | n | - | 6.8 | none | 18 | none | 4.5 | -
2025-10-11 | 284 | 1150 | PC | RM LF | * | 11.25 | 0.6 | 0.87205 | 4 |  |  | 69 | -12.2 | y | - | 6.7 | none | 18 | none | 4.5 | -
2025-10-12 | 285 | 1231 | GM, PC | RM LF | * | 11.212 | 0.598 | 0.86505 | 4 |  |  | 18 | -12.2 | n | - | 6.5 | none | 18 | none | 4.5 | PASS | Adjusted CLAP flow from .598 to .600
2025-10-13 | 286 | 1004 | GM, PC | RM LF | * | 11.237 | 0.603 | 0.86443 | 4 |  |  | 16.5 | -12.2 | n | - | 6.7 | none | 18 | none | 4.5 | - | Adjusted TSI flow closer to 18 after discussion, Ecotech flow changed as a result of TSI change and adjusted back to 4.5. Also changed TSI lamp bulb at 1600Z. Adjusted CLAP flow from 0.633 to 0.599
2025-10-14 | 287 | 1024 | GM, PC | - | * | 8.145 | 0.592 | 0.85998 | 4 |  |  | 19.6 | -11.6 | n | - | 6 | none | 18 | none | 4.5 | - | Too windy to go to roof
2025-10-15 | 288 | 1627 | GM, PC | CLR | * | 8.375 | 0.6 | 0.85467 | 4 |  |  | 63 | -12.2 | y | - | 6 | none | 18 | none | 4.5 | - | CLAP flow changed from .608 to .600
2025-10-16 | 289 | 1017 | GM | CLR | * | 8.446 | 0.6 | 0.85058 | 4 |  |  | 95 | -10.2 | n | - | 6 | none | 18 | none | 4.5 | - | CLAP flow adjusted from .605 to .600
2025-10-17 | 290 | 1109 | GM | RM LF | * | 8.379 | 0.593 | 0.84539 | 4 |  |  | 122 | -10.7 | n | - | 6 | none | 18 | none | 4.5 | - | Adjusted CLAP flow from .593 to .600
2025-10-18 | 291 | 1236 | GM | rm lf | * | 8.392 | 0.599 | 0.84427 | 4 |  |  | 50 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | -
2025-10-19 | 292 | 1353 | GM | rm lf | * | 8.368 | 0.595 | 0.84189 | 4 |  |  | 79 | -12.2 | y | - | 6 | RH out of range | 18 | none | 4.5 | - | Adjusted CLAP flow from .595 to .600
2025-10-20 | 293 | 1131 | GM | rm lf | * | 8.382 | 0.599 | 0.84003 | 4 |  |  | 54 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | -
2025-10-21 | 294 | 1136 | GM | rm lf | * | 8.348 | 0.601 | 0.83062 | 4 |  |  | 142 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | -
2025-10-22 | 295 | 1648 | GM | clr | * | 8.295 | 0.597 | 0.82533 | 4 |  |  | 76 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | - | Adjusted CLAP flow from 0.597 to 0.600
2025-10-23 | 296 | 1246 | GM | rm lf | * | 8.331 | 0.6 | 0.82298 | 4 |  |  | 5 | -12.2 | y | - | 6 | RH out of range | 18 | none | 4.5 | - | Magic CPC wick re-wetting and humidifier prep performed 1149Z - 1218Z. Some particles noticed inside humidifier, could not get them out during humidifier prep.
2025-10-24 | 297 | 1101 | GM | rm lf | * | 8.38 | 0.603 | 0.82207 | 4 |  |  | 54.9 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | - | Set up Magic 210 for inter-comparison, had Magic 205 turned off 1224Z-1233Z, rebooted acquisition software and laptop several times between 1235Z-1257Z.
2025-10-25 | 298 | 1056 | GM | clr | * | 8.273 | 0.593 | 0.8187 | 4 |  |  | 88 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | PASS | Adjusted CLAP flow from 0.593 to 0.600
2025-10-26 | 299 | 1326 | GM | rm lf | * | 8.302 | 0.6 | 0.81592 | 4 |  |  | 11 | -12 | n | - | 6 | RH out of range | 18 | none | 4.5 | -
2025-10-27 | 300 | 1546 | GM | rm lf | * | 8.319 | 0.6 | 0.81418 | 4 |  |  | 64 | -12.2 | y | - | 6 | RH out of range | 18 | none | 4.5 | -
2025-10-28 | 301 | 1542 | GM | rm lf | * | 8.332 | 0.601 | 0.81118 | 4 |  |  | 2 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | - | Magic 250 wick re-wetting, humidifier prep, and nozzle cleaning performed 1552Z - 1633Z
2025-10-29 | 302 | 1342 | GM | rm lf | * | 8.322 | 0.6 | 0.80873 | 4 |  |  | 56 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | -
2025-10-30 | 303 | 1450 | GM | rm lf |  | 8.375 | 0.604 | 0.80635 | 4 |  |  | 7 | -12.2 | y | - | 6 | RH out of range | 18 | none | 4.5 | -
2025-10-31 | 304 | 1644 | GM | rm lf |  | 8.367 | 0.601 | 0.80253 | 4 |  |  | 62 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | -
2025-11-01 | 305 | 1345 | GM | rm lf |  | 8.336 | 0.599 | 0.80159 | 4 |  |  | 124 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | -
2025-11-02 | 306 | 1440 | GM | rm lf |  | 8.268 | 0.592 | 0.7999 | 4 |  |  | 24 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | - | Adjusted CLAP flow from 0.592 to 0.600
2025-11-03 | 307 | 1941 | GM | - |  | 8.212 | 0.595 | 0.79798 | 4 |  |  | 39.3 | -12.2 | n | - | 6 | RH out of range | - | none | - | - | No in person checks due to ICESat
2025-11-04 | 308 | 1250 | GM | rm lf |  | 8.247 | 0.598 | 0.77696 | 4 |  |  | 31 | -12.2 | y | - | 6 | RH out of range | 18 | none | 4.5 | -
2025-11-05 | 309 | 1750 | GM | rm lf |  | 8.385 | 0.609 | 0.79547 | 4 |  |  | 9.3 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | -
2025-11-06 | 310 | 1146 | HB | rm lf | * | 8.44 | 0.613 | 0.79383 | 4 |  |  | 32.9 | -11.9 | n | - | 6 | RH out of range | 18 | none | 4.5 | - | Adjusted CLAP flow from 0.613 to 0.600
2025-11-07 | 311 | 1621 | GM | rm lf | * | 8.487 | 0.603 | 0.79225 | 4 |  |  | 20 | -11.5 | y | Ecotech: 2.3% | 6 | none | 18 | none | 4.5 | PASS
2025-11-08 | 312 | 1307 | GM | rm lf | * | 8.534 | 0.606 | 0.79228 | 4 |  |  | 43 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | - | Adjusted CLAP fow from 0.606 to 0.600
2025-11-09 | 313 | 1447 | HB | rm lf | * | [8 empty] | n | - | 6 |  | 18 |  | 4.5 | -
2025-11-10 | 314 | 1247 | GM | rm lf | * | 8.519 | 0.6 | 0.7874 | 4 |  |  | 31 | -10.7 | y | - | 6 | RH out of range | 18 | none | 4.5 | -
2025-11-11 | 315 | 1812 | GM | - | * | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | Did not go to AWO due to ICESat and could not login to FORGE acquisition webpage
2025-11-12 | 316 | 1712 | GM | rm lf | * | 8.533 | 0.599 | 0.78488 | 4 |  |  | 24 | -11.1 | n | - | 6 | none | 18 | none | 4.5 | -
2025-11-13 | 317 | 1309 | GM | clr | * | 8.567 | 0.601 | 0.78271 | 4 |  |  | 73 | -10.1 | y | - | 6 | none | 18 | none | 4.5 | -
2025-11-14 | 318 | 1357 | GM | lf | * | 8.524 | 0.596 | 0.78065 | 4 |  |  | 58 | -11 | n | - | 6 | none | 18 | none | 4.4 | - | Adjusted CLAP flow from 0.596 to 0.600
2025-11-15 | 319 | 1251 | GM | rm lf | * | 8.612 | 0.606 | 0.7792 | 4 |  |  | 68 | -12 | n | - | 6 | RH out of range | 18 | none | 4.5 | - | Adjusted CLAP flow from 0.606 to 0.600
2025-11-16 | 320 | 1339 | GM | rm lf | * | 8.587 | 0.597 | 0.7777 | 4 |  |  | 30 | -12.2 | y | - | 6 | RH out of range | 18 | none | 4.5 | - | Adjusted CLAP flow from 0.597 to 0.600
2025-11-17 | 321 | 1701 | GM | rm lf | * | 8.661 | 0.605 | 0.77531 | 4 |  |  | 80 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | - | Adjusted CLAP flow from 0.605 to 0.600
2025-11-18 | 322 | 1308 | GM | rm lf | * | 8.678 | 0.599 | 0.76962 | 4 |  |  | 103 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | -
2025-11-19 | 323 | 1401 | GM | rm lf | * | 8.571 | 0.591 | 0.76126 | 4 |  |  | 47 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | - | Adjusted CLAP flow from 0.591 to 0.600
2025-11-20 | 324 | 1409 | GM | rm lf | * | 8.463 | 0.591 | 0.75989 | 4 |  |  | 37 | -12.1 | n | - | 6 | RH out of range | 18 | none | 4.5 | - | Adjusted CLAP flow from 0.591 to 0.600
2025-11-21 | 325 | 1609 | GM | rm lf | * | 8.393 | 0.593 | 0.75954 | 4 |  |  | 19 | -12.2 | y | - | 6 | RH out of range | 18 | none | 4.5 | PASS | Adjusted CLAP flow from 0.593 to 0.600, Magic 250 wick re-wetting, humidifier prep, and nozzle cleaning from 1617Z-1642Z. Got Magic 210 onine for comparisson with 250.
2025-11-22 | 326 | 1234 | GM | rm lf | * | 8.39 | 0.598 | 0.75874 | 4 |  |  | 31 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | -
2025-11-23 | 327 | 1317 | GM | rm lf | * | 8.433 | 0.601 | 0.73969 | 4 |  |  | 58 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | -
2025-11-24 | 328 | 1537 | GM | rm lf | * | 8.346 | 0.594 | 0.73762 | 4 |  |  | 9 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | - | Adjusted CLAP flow from 0.594 to 0.600
2025-11-25 | 329 | 1325 | GM | rm lf | * | 8.272 | 0.593 | 0.73495 | 4 |  |  | 17 | -11.9 | n | - | 6 | none | 18 | none | 4.5 | - | Adjusted CLAP flow from 0.593 to 0.600
2025-11-26 | 330 | 1747 | GM | rm lf | * | 8.279 | 0.598 | 0.73438 | 4 |  |  | 29 | -12.2 | y | - | 6 | RH out of range | 18 | none | 4.5 | - | Unplugged Magic250 and performed wick re-wetting, humidifier prep, and nozzle cleaning from 1808Z-1835Z
2025-11-27 | 331 | 1558 | GM |  | * |  | 0.595 | 0.73355 | 4 |  |  | 34 |  | n | - | 6 |  |  |  |  | - | Did not go out to AWO and could not get all values from remote website
2025-11-28 | 332 | 1326 | GM | rm lf | * | 8.277 | 0.598 | 0.73315 | 4 |  |  | 14 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | - | Magic250 wick saturation is already down to 67% and a pulse height of 415mV. Down from 99% and 1173mV right after wick re-wetting suite performed on 2025-11-26.
2025-11-29 | 333 | 1305 | GM | rm lf | * | 8.315 | 0.601 | 0.73279 | 4 |  |  | 8 | -12.2 | y | - | 6 | RH out of range | 18 | none | 4.5 | - | Unplugged Magic250 and performed wick re-wetting, humidifier prep, and nozzle cleaning from 1311Z-1342Z
2025-11-30 | 334 | 1306 | GM | rm lf | * | 8.289 | 0.598 | 0.73199 | 4 |  |  | 5 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | - | Magic 250 already back at wick saturation 68% and pulse height 434mV.
2025-12-01 | 335 | 1659 | GM | rm lf | * | 8.324 | 0.601 | 0.73169 | 4 |  |  | 1 | -12.2 | y | - | 6 | RH out of range | 18 | none | 4.5 | - | Unplugged Magic250 and performed wick re-wetting, humidifier prep, and nozzle cleaning from 1850Z-1919Z
2025-12-02 | 336 | 1826 | GM | - | * | 8.373 | 0.604 | 0.72971 | 4 |  |  | 19.5 | -12.2 | n | - | 6 | RH out of range | - | none | - | - | Did not go out to AWO due to ICESat
2025-12-03 | 337 | 1737 | GM | rm lf | * | 8.448 | 0.61 | 0.72868 | 4 |  |  | 37 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | - | Adjusted CLAP flow from .610 to .600
2025-12-04 | 338 | 1456 | GM | rm lf | * | 8.46 | 0.601 | 0.72712 | 4 |  |  | 64 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | -
2025-12-05 | 339 | 1856 | GM | rm lf | * | 8.516 | 0.608 | 0.72553 | 4 |  |  | 8 | -12.2 | n | TSI: 5.7%, 1.7% | 6 | RH out of range | 18 | none | 4.5 | PASS | Spancheck 1823Z-1855Z & 1906-1937
2025-12-06 | 340 | 1303 | GM | rm lf | * | 8.569 | 0.601 | 0.72576 | 4 |  |  | 1 | -12.2 | y | - | 6 | RH out of range | 18 | none | 4.5 | - | Magic250 wick re-wetting suite: 1151-1217Z
2025-12-07 | 341 | 1409 | GM | rm lf | * | 8.554 | 0.598 | 0.72474 | 4 |  |  | 26 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | - | Magic250 showing Errors "Wick dry, wick recovery mode active, Thermistor fault detected, check connections (temperatures suspect), Dewpoint calculation error (contact manufacturer)". Magic250 turned off for wick re-wetting and humidifier prep and inspection from 1420-1430Z, Errors have gone away for now, will be checking on it periodically.
2025-12-08 | 342 | 1728 | GM | rm lf | * | 8.53 | 0.597 | 0.71067 | 4 |  |  | 107 | -12.2 | y | - | 6 | RH out of range | 18 | none | 4.5 | -
2025-12-09 | 343 | 1312 | HB | rm lf | * | 8.508 | 0.594 | 0.70249 | 4 |  |  | 26 | -12.2 | y | - | 6 | RH out of range | 18 | none | 4.5 | - | Adjusted CLAP flow from 0.594 to 0.600. Magic250 wick re-wetting suite: 1503-1537Z
2025-12-10 | 344 | 1739 | GM | rm lf | * | 8.443 | 0.594 | 0.70106 | 4 |  |  | 45 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | - | Adjusted CLAP flow from 0.594 to 0.600.
2025-12-11 | 345 | 1342 | GM | rm lf | * | 8.471 | 0.621 | 1.0004 | 5 |  |  | 28 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | - | Adjusted CLAP flow from 0.622 to 0.600
2025-12-12 | 346 | 1556 | GM | - | * | 8.332 | 0.589 | 0.99824 | 5 |  |  | 39 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | - | Adjusted CLAP flow from 0.589 to 0.600
2025-12-13 | 347 | 1300 | GM | rm lf | * | 8.347 | 0.596 | 0.99716 | 5 |  |  | 7 | -12.2 | y | - | 6 | RH out of range | 18 | none | 4.5 | -
2025-12-14 | 348 | 1251 | HB | rm lf | * | 8.351 | 0.6 | 0.99705 | 5 |  |  | 1 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | -
2025-12-15 | 349 | 1727 | GM | rm lf | * | 8.358 | 0.6 | 0.99705 | 5 |  |  | 0 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | - | Magic 250 showing moderator temp is out of range (>1.25C) and wick is dry and in wick recovery mode. Magic210 is measuring 10/cc.
2025-12-16 | 350 | 1423 | GM | - | * | 8.342 | 0.599 | 0.99656 | 5 |  |  | 0 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | - | Magic 250 showing conditioner and moderator temp is out of range (>1.25C) and wick is dry and in wick recovery mode. Magic210 is measuring 5/cc.
2025-12-17 | 351 | 1626 | GM | rm lf | * | 8.395 | 0.604 | 0.99609 | 5 |  |  | 11 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | -
2025-12-18 | 352 | 1317 | GM | rm lf | * | 8.477 | 0.606 | 0.99495 | 5 |  |  | 0 | -12.2 | y | - | 6 | RH out of range | 18 | none | 4.5 | -
2025-12-19 | 353 | 1316 | GM | rm lf | * | 8.46 | 0.605 | 0.99428 | 5 |  |  | 4 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | - | Magic 250 showing conditioner and moderator temp is out of range (>1.25C) and wick is dry and in wick recovery mode
2025-12-20 | 354 | 1237 | GM | rm lf | * | 8.543 | 0.612 | 0.9934 | 5 |  |  | 8 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | PASS | Adjusted CLAP flow from 0.612 to 0.600.
2025-12-21 | 355 | 1223 | GM | - | * | 8.628 | 0.598 | 0.99155 | 5 |  |  | 32.7 | -10.1 | n | - | 6 | none | - | none | - | - | Did not go to AWO due to Con 1
2025-12-22 | 356 | 1517 | GM | clr | * | 8.645 | 0.6 | 0.99005 | 5 |  |  | 13 | -11 | n | - | 6 | none | 18 | none | 4.5 | -
2025-12-23 | 357 | 1341 | GM | rm lf | * | 8.548 | 0.593 | 0.98911 | 5 |  |  | 70 | -11.9 | n | - | 6 | RH out of range | 18 | none | 4.5 | - | Adjusted CLAP flow from 0.593 to 0.600
2025-12-24 | 358 | 1244 | GM | rm lf | * | 8.548 | 0.598 | 0.98611 | 5 |  |  | 16 | -12.2 | y | - | 6 | RH out of range | 18 | none | 4.5 | - | Magic250 in wick recovery mode
2025-12-25 | 359 | 1812 | GM | - | * | 8.622 | 0.605 | 0.98133 | 5 |  |  | 46.2 | -12.2 | n | - | 6 | RH out of range | - | none | - | - | Did not go out to AWO
2025-12-26 | 360 | 1713 | GM | rm lf | * | 8.814 | 0.619 | 0.9748 | 5 |  |  | 3 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | - | Magic 250 showing conditioner and moderator temp is out of range (>1.25C) and wick is dry and in wick recovery mode. Magic210 is measuring ~13/cc. Adjusted CLAP flow from 0.619 to 0.600.
2025-12-27 | 361 | 1242 | GM | - | * | 8.699 | 0.593 | 0.97248 | 5 |  |  | 4.1 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | PASS | Ecotech not communicating as of 0715Z. Magic250 showing moderator temp out of range and wick is dry and in recovery mode. Restarted acquisition software at 1723Z to attempt reconnection of ecotech. Power cycled Ecotech at 1730Z, this brought it back online. Errors "RH sensor fault detected" and "Failure detected during instrument warmup" appeared at first boot-up but went away within first minute.
2025-12-28 | 362 | 1303 | GM | rm lf | * | 8.9 | 0.609 | 0.97228 | 5 |  |  | 0 | -12.2 | n | - | 6 | RH out of range | 18 | none | 4.5 | - | Magic250 modirator temp out of range and wick in recovery mode. Magic210 reading ~6.0/cm3. Adjusted CLAP flow from 0.609 to .600
2025-12-29 | 363 | 1452 | GM | rm lf | * | 8.916 | 0.602 | 0.96954 | 5 |  |  | 56 | -12.2 | y | - | 6 | RH out of range | 18 | none | 4.5 | -
2025-12-30 | 364 | 1314 | GM | clr | * | 8.818 | 0.595 | 0.96696 | 5 |  |  | 11 | -11.1 | n | - | 6 | none | 18 | none | 4.5 | -
2025-12-31 | 365 | 1220 | GM | - | * | 8.943 | 0.604 | 0.96673 | 5 |  |  | 37.6 | -11.9 | N | - | 6 | none | - | none | - | - | Did not go out to AWO
Note on G1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: If >12, inlet is likely clogged. Remove pot and add to flux master log.
Note on H1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: A11 GMD CLAP Box: Flow (Between 0.59-0.61 lpm)
Note on I1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Green Transmittance: (When it drops below 0.7 it should change to a new spot) Reply: Has a total of 8 spots. Filter should be changed when spot 8 is nearing 0.7.
Note on J1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Spot
Note on K1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Concentration
Note on L1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Dewpoint
Note on M1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Concentration Reply: Per Erin/PI: above 0 is normal. Concentration seems to trend lower in the winter. If hits 0, ask Erin and consider opening the lines to room air (search concentration in email and consult Erin)
Note on N1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Inlet Dewpoint
Note on P1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Alternate between Nephelometers so that each neph gets a span check every other month but you complete a span check every month. Per protocol
Note on Q1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: S11 TSI Box: Lamp (A)
Note on R1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Flags
Note on S1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: On physical Neph stucture Reply: Per protocol. 18 is marked with sharpie.
Note on T1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Check the flag status on Forge (computer)
Note on U1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Check rotameter on Ecotech. Okay if it changes 5> flow >4. Per protocol.
Note on V1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Check sys status on ecotech screen, verify everything is “pass”

## Surface Ozone
Date | Julian Day | Time (UTC) | Observer's Initials | Inlet | Panel Display (30-70 ppb) | Bench Temp (20-40C) | Pressure (533-1040 hPa) | Flow rate (LPM) (0.4-0.9) | Flow rate (LPM) (0.4-0.9) | Intensity Cell A (65,000- 120,000Hz) | Intensity Cell B (65,000- 120,000Hz) | Comments (Alarms, Calibration Levels...)
2025-01-01 | 1 | 1602 | ALW | rm lf | 42.52 | 32.5 | 676.2 | 0.472 | 0.506 | 78098 | 93352 | Replaced filter @ 1512Z
2025-01-02 | 2 | 1553 | ALW | rm lf | 47.19 | 34 | 676.7 | 0.471 | 0.506 | 78721 | 94075
2025-01-03 | 3 | - | ALW | - | - | - | - | - | - | - | - | Icesat- did not go to AWO
2025-01-04 | 4 | 1348 | ALW | rm lf | 35.82 | 34 | 677.4 | 0.472 | 0.507 | 78638 | 93970
2025-01-05 | 5 | 1527 | ALW | rm lf | 41.47 | 34.2 | 676.9 | 0.472 | 0.506 | 78726 | 94058
2025-01-06 | 6 | 1702 | ALW | rm lf | 50.51 | 34.6 | 675.7 | 0.471 | 0.505 | 78859 | 94176
2025-01-07 | 7 | 1848 | ALW | rm lf | 44.16 | 33.8 | 677.4 | 0.472 | 0.507 | 78663 | 93966 | Surface Ozone Calibration: 1508-1610Z. Z: 0.5, L1: 169.2, L2: 42.3.
2025-01-08 | 8 | 1746 | ALW | - | 47.52 | 34.2 | 671.3 | 0.469 | 0.504 | 78623 | 93899
2025-01-09 | 9 | 1545 | ALW | rm lf | 42.31 | 33.8 | 660.9 | 0.464 | 0.5 | 78562 | 93850
2025-01-10 | 10 | - | ALW | - | - | - | - | - | - | - | - | Con 1- did not go to AWO
2025-01-11 | 11 | 1253 | ALW | rm lf | 46.28 | 31.8 | 664.5 | 0.467 | 0.501 | 77971 | 93155
2025-01-12 | 12 | 1300 | ALW | rm lf | 48.9 | 34.7 | 657.3 | 0.462 | 0.497 | 78835 | 94069
2025-01-13 | 13 | 1557 | ALW | rm lf | 41.77 | 34.2 | 655.4 | 0.462 | 0.497 | 78595 | 93811
2025-01-14 | 14 | 1448 | ALW | rm lf | 47.17 | 34.2 | 653.7 | 0.461 | 0.496 | 78652 | 93844 | Orange cap found on the ground 1345Z, windy-did not go to roof
2025-01-15 | 15 | 1623 | ALW | rm lf | 47.07 | 34 | 649.3 | 0.459 | 0.494 | 78627 | 93786
2025-01-16 | 16 | - | ALW | - | - | - | - | - | - | - | - | Con 1- did not go to AWO
2025-01-17 | 17 | - | ALW | - | - | - | - | - | - | - | - | Con 1- did not go to AWO
2025-01-18 | 18 | 1245 | ALW | rm lf | 39.75 | 28.8 | 669.5 | 0.47 | 0.505 | 77121 | 92023
2025-01-19 | 19 | 1526 | ALW | rm lf | 50.24 | 33.8 | 676.5 | 0.471 | 0.506 | 78397 | 93434
2025-01-20 | 20 | 1501 | ALW | rm lf | 43.85 | 33.8 | 675.4 | 0.471 | 0.507 | 78387 | 93461
2025-01-21 | 21 | 1247 | ALW | rm lf | 43.83 | 33.6 | 669.7 | 0.469 | 0.503 | 78364 | 93392 | Surface zone Calibration 1330-1415Z. Z: 0.2, L1: 169.6, L2: 42.9.
2025-01-22 | 22 | 1540 | ALW | rm lf | 45.96 | 34.3 | 665.9 | 0.467 | 0.503 | 78149 | 93400 | Replaced filter at 1821Z due to high wids the previous day.
2025-01-23 | 23 | 1605 | ALW | rm lf | 43.54 | 34.5 | 659.7 | 0.464 | 0.499 | 78579 | 93592
2025-01-24 | 24 | 1305 | ALW | rm lf | 46.33 | 34.8 | 657.9 | 0.463 | 0.498 | 78730 | 93745
2025-01-25 | 25 | 1142 | ALW | rm lf | 48.54 | 35.1 | 659.8 | 0.464 | 0.499 | 78848 | 93870
2025-01-26 | 26 | 1416 | ALW | rm lf | 46.59 | 34.7 | 660.7 | 0.465 | 0.5 | 78527 | 93554
2025-01-27 | 27 | 1453 | ALW | rm lf | 40.72 | 31.9 | 658.3 | 0.464 | 0.499 | 77818 | 92732
2025-01-28 | 28 | 1520 | ALW | rm lf | 33.27 | 34.7 | 659.1 | 0.464 | 0.499 | 78566 | 93521 | Surface Ozone calibration 1604-1650Z. Z: 1.4, L1: 170.4, L2: 42.9.
2025-01-29 | 29 | 1621 | ALW | - | 37.02 | 34.1 | 663 | 0.466 | 0.501 | 78262 | 93133 | Too windy to go to roof
2025-01-30 | 30 | 1304 | ALW | rm lf | 43.73 | 34.5 | 662.1 | 0.465 | 0.501 | 78577 | 93509
2025-01-31 | 31 | 1330 | ALW | rm lf | 39.95 | 34.3 | 660.5 | 0.464 | 0.5 | 78395 | 93287 | Candy cane installed 1540Z
2025-02-01 | 32 | - | ALW | - | - | - | - | - | - | - | - | Con 1- did not go to AWO
2025-02-02 | 33 | 1513 | ALW | rm lf | 42.95 | 34.6 | 653.1 | 0.461 | 0.496 | 78529 | 93397
2025-02-03 | 34 | 1417 | ALW | rm lf | 43.17 | 34.1 | 654.7 | 0.461 | 0.497 | 78310 | 93159 | Surface Ozone Calibration 1510-1556Z. Z: 0.0, L1: 169.2, L2: 42.5.
2025-02-04 | 35 | - | ALW | - | - | - | - | - | - | - | - | ICEsat- did not go to AWO
2025-02-05 | 36 | 1624 | ALW | rm lf | 42.61 | 35.6 | 639.4 | 0.454 | 0.489 | 78724 | 93543
2025-02-06 | 37 | 1142 | ALW | rm lf | 42.8 | 34.2 | 642.3 | 0.456 | 0.491 | 78386 | 93187
2025-02-07 | 38 | - | ALW | - | - | - | - | - | - | - | - | Con 1- did not go to AWO
2025-02-08 | 39 | 1809 | ALW | rm lf | 30.72 | 31.2 | 663.4 | 0.466 | 0.501 | 77507 | 92186
2025-02-09 | 40 | - | ALW | - | - | - | - | - | - | - | - | Con 1- did not go to AWO
2025-02-10 | 41 | 1606 | ALW | - | - | - | - | - | - | - | - | Icesat- did not go to AWO
2025-02-11 | 42 | 1441 | ALW | rm lf | 40.98 | 34.7 | 681.8 | 0.475 | 0.51 | 78243 | 92918 | Surface Ozone Calibration: 11312-1359Z. Z: 0.4, L1: 170.0, L2: 42.7.
2025-02-12 | 43 | 1107 | ALW | rm lf | 45.35 | 34.9 | 691.4 | 0.478 | 0.513 | 78483 | 93143
2025-02-13 | 44 | 1752 | ALW | rm lf | 48.75 | 34.6 | 690.6 | 0.478 | 0.513 | 78423 | 93122
2025-02-14 | 45 | 1428 | AR | rm lf | 44.27 | 34.7 | 692.1 | 0.48 | 0.514 | 78406 | 93070
2025-02-15 | 46 | - | - | - | - | - | - | - | - | - | - | Icesat- did not go to AWO
2025-02-16 | 47 | 1603 | AR | rm lf | 45.1 | 34.6 | 689.7 | 0.479 | 0.513 | 78300 | 92903
2025-02-17 | 48 | 1353 | AR | rm lf | 48.03 | 34.4 | 685.5 | 0.477 | 0.511 | 78208 | 92817 | Surface Ozone Calibration 1826-1911Z. Z: 0.1, L1: 170.6, L2: 43.7. Planned power outages for turnover 1201-1204Z and 1822-1824Z. Instruments supported by UPS.
2025-02-18 | 49 | 1337 | AR | rm lf | 51.15 | 34.9 | 676.6 | 0.472 | 0.507 | 78372 | 92958
2025-02-19 | 50 | 1816 | AR | rm lf | 52.11 | 35.1 | 667.7 | 0.468 | 0.503 | 78456 | 93028
2025-02-20 | 51 | 1913 | ALW | rm lf | 51.54 | 34.7 | 658.9 | 0.462 | 0.497 | 78297 | 92754 | Power outage today from 0046-1147Z. Forge acquisition restarted at 1530Z.
2025-02-21 | 52 | 1448 | AR | rm lf | 50.73 | 32.2 | 660.5 | 0.464 | 0.499 | 77838 | 92245
2025-02-22 | 53 | 1413 | AR | rm lf | 48.92 | 34.7 | 649.7 | 0.458 | 0.493 | 78120 | 92536
2025-02-23 | 54 | 1313 | AR | rm lf | 48.4 | 35.1 | 652.3 | 0.459 | 0.494 | 78226 | 92657
2025-02-24 | 55 | 1513 | AR | rm lf | 47.17 | 35.2 | 653.5 | 0.46 | 0.495 | 78344 | 92760 | Surface Ozone Calibration 1547-1634Z Z: 0.1, L1:169.8, L2:42.9
2025-02-25 | 56 | - | AR | rm lf | - | - | - | - | - | - | -
2025-02-26 | 57 | 1533 | AR | rm lf | 43.14 | 35.1 | 659 | 0.463 | 0.498 | 78221 | 92630
2025-02-27 | 58 | 1300 | AR | rm lf | 42.47 | 34.9 | 657 | 0.462 | 0.498 | 78164 | 92577
2025-02-28 | 59 | 1149 | AR | - | 43.35 | 31.9 | 654.9 | 0.462 | 0.497 | 77253 | 91521 | Con 1- did not go to AWO
2025-03-01 | 60 | 1256 | AR | rm lf | 44.51 | 34.9 | 659.9 | 0.463 | 0.499 | 78159 | 92501
2025-03-02 | 61 | 1550 | AR | rm lf | 43.66 | 34.8 | 656.5 | 0.461 | 0.497 | 78057 | 92370
2025-03-03 | 62 | 1532 | AR | rm lf | 46.48 | 34.9 | 660.1 | 0.463 | 0.498 | 78252 | 92534
2025-03-04 | 63 | 1534 | AR | rm lf | 44.54 | 35.2 | 656.3 | 0.461 | 0.497 | 78275 | 92591 | Surface Ozone Calibration: 1543-1629Z. Z:0.0, L1:170.0, L2:43.2.
2025-03-05 | 64 | 1757 | AR | - | 43.43 | 34.9 | 662.9 | 0.465 | 0.5 | 78136 | 92401 | ICEsat- did not go to AWO
2025-03-06 | 65 | 1330 | AR | rm lf | 43.34 | 34.9 | 666.1 | 0.466 | 0.502 | 78205 | 92459
2025-03-07 | 66 | 1715 | AR | rm lf | 41.38 | 34.9 | 670.3 | 0.468 | 0.503 | 78084 | 92288
2025-03-08 | 67 | 1139 | AR | rm lf | 43.27 | 35.1 | 672.5 | 0.469 | 0.504 | 77981 | 92188
2025-03-09 | 68 | 1730 | AR | - | 45.19 | 34.8 | 675.9 | 0.471 | 0.506 | 77923 | 92082 | Con 1- did not go to AWO
2025-03-10 | 69 | 1233 | AR | rm lf | 49.25 | 34.8 | 677.7 | 0.472 | 0.507 | 77964 | 92126
2025-03-11 | 70 | 1258 | AR | rm lf | 49.08 | 35.2 | 681.5 | 0.474 | 0.508 | 78064 | 92190 | Surface Ozone Calibration 1333-1420Z Z:-0.3, L1:170.0, L2:43.8; Power outage today at 2150Z (less than 1 min). Instruments supported by UPS.
2025-03-12 | 71 | 1634 | AR | rm lf | 49.73 | 35.2 | 678.3 | 0.472 | 0.507 | 77876 | 91953
2025-03-13 | 72 | 1318 | AR | rm lf | 50.42 | 34.8 | 677.4 | 0.472 | 0.507 | 77936 | 92011
2025-03-14 | 73 | 1159 | AR | rm lf | 48.29 | 35.2 | 668.2 | 0.468 | 0.503 | 77870 | 91942
2025-03-15 | 74 | 1254 | AR | rm lf | 40.76 | 35 | 669.9 | 0.468 | 0.503 | 78194 | 92281
2025-03-16 | 75 | 1347 | AR | rm lf | 47.72 | 34.7 | 668.5 | 0.468 | 0.503 | 77876 | 91892
2025-03-17 | 76 | 1522 | AR | rm lf | 41.78 | 34.7 | 668.5 | 0.468 | 0.502 | 77932 | 91908
2025-03-18 | 77 | 1221 | AR | rm lf | 49.41 | 33.9 | 670.1 | 0.468 | 0.504 | 77604 | 91575
2025-03-19 | 78 | 1736 | AR | rm lf | 47.44 | 35.4 | 665.4 | 0.466 | 0.502 | 78032 | 92003 | Surface Ozone Calibration 1603-1649 Z Z:0.0, L1:169.8, L2:43.0
2025-03-20 | 79 | 1129 | AR | rm lf | 47.37 | 35 | 660.3 | 0.464 | 0.499 | 77938 | 91913
2025-03-21 | 80 | 1521 | AR | rm lf | 46.95 | 34.9 | 657.8 | 0.462 | 0.498 | 77934 | 91826
2025-03-22 | 81 | 1227 | AR | rm lf | 43.91 | 34.9 | 658.1 | 0.463 | 0.499 | 77953 | 91870
2025-03-23 | 82 | 1308 | AR | rm lf | 46.54 | 34.6 | 663.7 | 0.465 | 0.501 | 77842 | 91751
2025-03-24 | 83 | 1631 | AR | rm lf | 50.99 | 35.1 | 661.5 | 0.464 | 0.5 | 77848 | 91690
2025-03-25 | 84 | 1236 | AR | rm lf | 46.86 | 35 | 663.5 | 0.465 | 0.501 | 77913 | 91767 | Surface Ozone Calibration 1606-1652Z Z:0.4, L1:171.1, L2:43.5
2025-03-26 | 85 | 1704 | AR | rm lf | 47.6 | 35 | 663.5 | 0.465 | 0.5 | 77779 | 91566
2025-03-27 | 86 | 1634 | AR | rm lf | 49.33 | 34.9 | 663.7 | 0.465 | 0.501 | 77876 | 91677
2025-03-28 | 87 | 1235 | AR | rm lf | 47.9 | 34.8 | 664.3 | 0.466 | 0.501 | 77870 | 91654
2025-03-29 | 88 | 1243 | AR | rm lf | 49.29 | 34.8 | 664.1 | 0.465 | 0.501 | 77938 | 91735
2025-03-30 | 89 | 1233 | AR | rm lf | 51.68 | 34.7 | 664.6 | 0.466 | 0.501 | 77789 | 91552
2025-03-31 | 90 | 1506 | AR | rm lf | 43.27 | 35.2 | 664.3 | 0.466 | 0.501 | 77968 | 91703
2025-04-01 | 91 | 1003 | AR | rm lf | 44.47 | 34.6 | 665.3 | 0.466 | 0.502 | 77423 | 91057 | Surface Ozone Calibration 1008-1054Z Z:-0.0 , L1:171.0 , L2:43.2
2025-04-02 | 92 | 1005 | AR | - | 40.49 | 34.8 | 668.5 | 0.468 | 0.503 | 77504 | 91129 | Icesat- did not go to AWO
2025-04-03 | 93 | 1519 | AR | rm lf | 46.05 | 34.2 | 674.9 | 0.471 | 0.506 | 77411 | 91030
2025-04-04 | 94 | 1528 | AR | rm lf | 45.04 | 34.8 | 677.7 | 0.472 | 0.507 | 77487 | 91076
2025-04-05 | 95 | 1100 | AR | rm lf | 51.13 | 34.7 | 679 | 0.473 | 0.507 | 77557 | 91164
2025-04-06 | 96 | 1228 | AR | rm lf | 51.16 | 34.3 | 672.9 | 0.471 | 0.505 | 77366 | 90933
2025-04-07 | 97 | 1135 | AR | rm lf | 46.04 | 34.3 | 676.1 | 0.472 | 0.507 | 77231 | 90754
2025-04-08 | 98 | 1207 | AR | rm lf | 54.33 | 34.5 | 681.5 | 0.474 | 0.509 | 77420 | 90912 | Surface Ozone Calibration 1245-1331Z Z:0.2 , L1:169.9 , L2:42.6
2025-04-09 | 99 | 1048 | AR | rm lf | 46.76 | 34.9 | 677.8 | 0.472 | 0.507 | 77413 | 90885
2025-04-10 | 100 | 1433 | AR | rm lf | 48.55 | 35 | 672.3 | 0.47 | 0.504 | 77533 | 90994
2025-04-11 | 101 | 1043 | AR | rm lf | 50.81 | 35.2 | 661.9 | 0.465 | 0.5 | 77624 | 91120
2025-04-12 | 102 | 1207 | AR | rm lf | 45.97 | 35.1 | 663.3 | 0.465 | 0.501 | 77694 | 91177
2025-04-13 | 103 | 1621 | AR | rm lf | 51.11 | 35 | 666.3 | 0.466 | 0.502 | 77600 | 91013
2025-04-14 | 104 | 1526 | AR | rm lf | 48.96 | 34.8 | 666.7 | 0.467 | 0.502 | 77655 | 91092
2025-04-15 | 105 | 1110 | AR | rm lf | 55.38 | 34.6 | 671.5 | 0.469 | 0.504 | 77402 | 90730 | Surface Ozone Calibration 1113-1159Z Z:0.2 , L1:170.2 , L2:42.8
2025-04-16 | 106 | 1427 | AR | rm lf | 49.72 | 34.5 | 678.1 | 0.472 | 0.507 | 77455 | 90825
2025-04-17 | 107 | 1053 | AR | clr | 56.7 | 34.7 | 679.4 | 0.473 | 0.507 | 77308 | 90587
2025-04-18 | 108 | 1208 | AR | clr | 56.15 | 34.7 | 676.7 | 0.472 | 0.506 | 77477 | 90798
2025-04-19 | 109 | 1132 | AR | clr | 60.86 | 34.5 | 676.3 | 0.471 | 0.507 | 77416 | 90703
2025-04-20 | 110 | 1238 | AR | rm lf | 51.48 | 34.9 | 676.1 | 0.417 | 0.508 | 77242 | 90466
2025-04-21 | 111 | 1521 | AR | rm lf | 49.56 | 35 | 680.1 | 0.473 | 0.508 | 77563 | 90828
2025-04-22 | 112 | 1515 | AR | rm lf | 46.87 | 34.9 | 678.9 | 0.473 | 0.508 | 77282 | 90417 | Surface Ozone Calibration 1614-1700Z Z:0.2 , L1:169.6 , L2:42.8
2025-04-23 | 113 | 1030 | AR | rm lf | 52.13 | 34.6 | 679.1 | 0.473 | 0.508 | 77199 | 90405
2025-04-24 | 114 | 1638 | AR | rm lf | 55.79 | 34.6 | 678.3 | 0.473 | 0.508 | 77245 | 90426
2025-04-25 | 115 | 1135 | AR | rm lf | 54.01 | 34.9 | 677.3 | 0.471 | 0.506 | 77348 | 90491
2025-04-26 | 116 | 1450 | AR | rm lf | 53.26 | 35 | 673.5 | 0.47 | 0.505 | 77272 | 90406
2025-04-27 | 117 | 1153 | AR | rm lf | 49.57 | 34.6 | 673.1 | 0.47 | 0.505 | 77436 | 90573
2025-04-28 | 118 | 1632 | AR | rm lf | 43.33 | 34.4 | 676.2 | 0.471 | 0.506 | 77246 | 90321
2025-04-29 | 119 | 1110 | AR | rm lf | 45.17 | 34.9 | 674.9 | 0.471 | 0.506 | 77208 | 90290 | Surface Ozone Calibration 1154-1240Z. Z:0.5, L1:170.5, L2:41.7
2025-04-30 | 120 | 1446 | AR | clr | 48.45 | 34.9 | 671.4 | 0.469 | 0.505 | 77484 | 90557
2025-05-01 | 121 | 1125 | AR | clr | 45.75 | 35 | 671.1 | 0.469 | 0.504 | 77245 | 90254
2025-05-02 | 122 | 1207 | AR | rm lf | 40.71 | 34.3 | 679.3 | 0.473 | 0.508 | 77282 | 90328
2025-05-03 | 123 | 1026 | AR | rm lf | 50 | 34.5 | 677.8 | 0.473 | 0.508 | 76957 | 89902
2025-05-04 | 124 | 1956 | AR | clr | 47.42 | 34.8 | 677.9 | 0.473 | 0.507 | 77899 | 90030
2025-05-05 | 125 | 1131 | AR | clr | 50.69 | 34.6 | 674.9 | 0.471 | 0.506 | 77109 | 90052
2025-05-06 | 126 | 1309 | AR | rm lf | 47.83 | 34.7 | 668.7 | 0.468 | 0.503 | 77125 | 90027 | Surface Ozone Calibration 1457-1544Z. Z:0.1, L1:169.4, L2:42.9
2025-05-07 | 127 | 1513 | AR | rm lf | 47.88 | 34.9 | 660.2 | 0.464 | 0.5 | 77210 | 90107
2025-05-08 | 128 | 1238 | AR | rm lf | 33.75 | 34.7 | 662.9 | 0.466 | 0.501 | 76978 | 89822
2025-05-09 | 129 | 1148 | AR | rm lf | 46.19 | 34.9 | 662.1 | 0.465 | 0.5 | 76974 | 89821
2025-05-10 | 130 | 1045 | AR | rm lf | 40.24 | 34.7 | 665 | 0.466 | 0.502 | 77071 | 89908 | Planned station-wide power outage from 1008-1014Z. Instruments supported by UPS.
2025-05-11 | 131 | 1549 | AR | rm lf | 46.87 | 34.8 | 663.8 | 0.466 | 0.501 | 77087 | 89918 | Unplanned station-wide power outage from 1455-1524Z. Instruments supported by UPS.
2025-05-12 | 132 | 1115 | AR | rm lf | 22.37 | 34.3 | 668.5 | 0.468 | 0.503 | 77025 | 89812
2025-05-13 | 133 | 1815 | AR | - | 48.53 | 35.3 | 673.1 | 0.47 | 0.505 | 77081 | 89837 | Con 1- did not go to AWO. Unplanned station power outage from 2031-2040Z. Controlled station power outage from 2105-2110Z. Instruments supported by UPS.
2025-05-14 | 134 | 1008 | AR | clr | 45.5 | 35.2 | 677.4 | 0.472 | 0.507 | 77214 | 89986 | Surface Ozone Calibration 1032-1118Z. Z:1.5, L1:164.8, L2:41.2
2025-05-15 | 135 | 1524 | AR | clr | 58.51 | 34.8 | 684.9 | 0.476 | 0.511 | 77171 | 89923
2025-05-16 | 136 | 1042 | AR | clr | 49.33 | 34.7 | 688.3 | 0.478 | 0.512 | 77048 | 89720
2025-05-17 | 137 | 948 | AR | clr | 49.15 | 34.6 | 687.8 | 0.478 | 0.512 | 76922 | 89598 | Planned station power outage from 1015-1019Z. Instruments supported by UPS.
2025-05-18 | 138 | 1154 | AR | clr | 48.4 | 34.4 | 690.7 | 0.479 | 0.513 | 76926 | 89579
2025-05-19 | 139 | 1555 | AR | clr | 52.3 | 34.2 | 696.3 | 0.482 | 0.516 | 76874 | 89496
2025-05-20 | 140 | 1023 | AR | rm lf | 57.64 | 34.7 | 696.7 | 0.481 | 0.516 | 76871 | 89469 | Surface Ozone Calibration 1135-1221Z. Z:0.4, L1:168.4, L2:42.7
2025-05-21 | 141 | 1442 | AR | clr | 60.86 | 34.7 | 684.2 | 0.476 | 0.51 | 76831 | 89410
2025-05-22 | 142 | 1009 | AR | clr | 58.58 | 34.9 | 677.5 | 0.472 | 0.508 | 76864 | 89433
2025-05-23 | 143 | 1235 | AR | rm lf | 54.84 | 35.1 | 673.1 | 0.47 | 0.505 | 76911 | 89449
2025-05-24 | 144 | 1045 | AR | clr | 35.24 | 35.1 | 668.1 | 0.468 | 0.503 | 76805 | 89342 | Unplanned station-wide power outage from 1806-1808Z. Instruments supported by UPS.
2025-05-25 | 145 | 1308 | AR | rm lf | 47.68 | 34.8 | 673.7 | 0.471 | 0.506 | 76625 | 89098
2025-05-26 | 146 | 1029 | AR | rm lf | 40.35 | 35.1 | 672.5 | 0.47 | 0.505 | 77061 | 89592
2025-05-27 | 147 | 1503 | AR | clr | 49.77 | 35 | 672.3 | 0.47 | 0.505 | 76765 | 89249 | Surface Ozone Calibration 1547-1633Z, Zero:0.1, L1:168.9, L2:42.6
2025-05-28 | 148 | 1135 | AR | rm lf | 46.99 | 35 | 673.1 | 0.47 | 0.506 | 76806 | 89243
2025-05-29 | 149 | 1456 | AR | rm lf | 47.48 | 35 | 673.8 | 0.471 | 0.506 | 76940 | 89401
2025-05-30 | 150 | 1208 | AR | rm lf | 46.42 | 34.1 | 674.5 | 0.472 | 0.507 | 76580 | 88962
2025-05-31 | 151 | 1555 | AR | clr | 50.4 | 34.7 | 676.9 | 0.472 | 0.507 | 76815 | 89216
2025-06-01 | 152 | 1801 | AR | clr | 46.2 | 36.5 | 676.5 | 0.47 | 0.585 | 77221 | 89618
2025-06-02 | 153 | 1557 | AR | clr | 43.83 | 37.2 | 678.5 | 0.471 | 0.507 | 77404 | 89796
2025-06-03 | 154 | 1059 | AR | clr | 52.03 | 34.9 | 677.4 | 0.472 | 0.507 | 76810 | 89186 | Surface Ozone Calibration 1143-1229Z, Zero: 0.2, L1: 169.1, L2: 42.3
2025-06-04 | 155 | 1613 | AR | - | 44.25 | 34.8 | 672.4 | 0.47 | 0.505 | 76674 | 88959 | Icesat- did not go to AWO
2025-06-05 | 156 | 1038 | AR | rm lf | 44.58 | 34.6 | 670.5 | 0.469 | 0.504 | 76584 | 88832
2025-06-06 | 157 | 1545 | AR | clr | 45.13 | 34.6 | 670.2 | 0.469 | 0.504 | 76694 | 88953
2025-06-07 | 158 | 1057 | AR | rm lf | 43.8 | 34.4 | 673.3 | 0.47 | 0.505 | 76505 | 88731
2025-06-08 | 159 | 1404 | AR | rm lf | 42.99 | 34.7 | 677.4 | 0.472 | 0.507 | 76720 | 88935
2025-06-09 | 160 | 1124 | AR | clr | 44.84 | 34.5 | 680.1 | 0.474 | 0.509 | 76588 | 88781 | Planned power outage 1001-1005Z. Instrument supported by UPS
2025-06-10 | 161 | 1641 | AR | clr | 48.64 | 35.2 | 678.1 | 0.472 | 0.507 | 76852 | 89043 | Surface Ozone Calibration 1545-1630Z, Zero: 0.3, L1: 165.7, L2: 41.6. Planned power outage 1704-1707Z. Instrument supported by UPS.
2025-06-11 | 162 | 1506 | AR | rm lf | 51.93 | 34.9 | 677.9 | 0.473 | 0.507 | 76597 | 88704
2025-06-12 | 163 | 1011 | PC | rm lf | 49.49 | 34.9 | 676.1 | 0.472 | 0.507 | 76477 | 88613
2025-06-13 | 164 | 1009 | PC | clr | 48.57 | 34.9 | 677.9 | 0.472 | 0.507 | 76457 | 88531 | Planned power outage 1037-1040Z. Second planned power outage 1412-1416Z. Instruments supported by UPS.
2025-06-14 | 165 | 1601 | PC | clr | 50.78 | 35.5 | 684.6 | 0.475 | 0.51 | 76651 | 88727
2025-06-15 | 166 | 1126 | PC | clr | 55.62 | 34.7 | 686.5 | 0.476 | 0.511 | 76485 | 88498
2025-06-16 | 167 | 1636 | PC | clr | 52.02 | 35.8 | 682.6 | 0.474 | 0.509 | 76883 | 88941
2025-06-17 | 168 | 1126 | PC | CLR | 47.86 | 35.5 | 682.7 | 0.474 | 0.509 | 76695 | 88720 | Surface ozone calibration 1500Z-1546Z. Zero: 0.4, L1: 168.6. L2: 42.3.
2025-06-18 | 169 | 1224 | PC | clr | 52.89 | 36.2 | 683.1 | 0.475 | 0.509 | 76461 | 88416
2025-06-19 | 170 | 1049 | PC | rm lf | 51.55 | 35.4 | 685.7 | 0.475 | 0.51 | 76689 | 88700
2025-06-20 | 171 | 1445 | PC | clr | 61.1 | 36.6 | 686.5 | 0.475 | 0.51 | 77010 | 89040
2025-06-21 | 172 | 1627 | PC | clr | 56.43 | 36.6 | 684.6 | 0.473 | 0.509 | 77053 | 89054
2025-06-22 | 173 | 1534 | PC | rm lf | 56.19 | 37.7 | 682.5 | 0.472 | 0.508 | 76682 | 88541
2025-06-23 | 174 | 1135 | PC | clr | 49.2 | 35.1 | 683.7 | 0.474 | 0.509 | 76551 | 88426
2025-06-24 | 175 | 905 | PC | rm lf | 48.62 | 35.3 | 683.1 | 0.474 | 0.509 | 76500 | 88315 | Plane flew a big low contrail over station 0905.
2025-06-25 | 176 | 1123 | PC | rm lf | 51.56 | 35.1 | 681.1 | 0.473 | 0.508 | 76506 | 88361 | Low winds equipment operations: 11:42-12:07Z- D6, 953, skidsteer, snowmo, welder (machines operating 11:42-11:49, welder missed radio call and operated 11:42-1207Z).
2025-06-26 | 177 | 940 | PC | rm lf | 52.99 | 35.5 | 676.5 | 0.471 | 0.506 | 76612 | 88431
2025-06-27 | 178 | 1422 | PC | clr | 30.81 | 35.6 | 676.5 | 0.471 | 0.506 | 76521 | 88284
2025-06-28 | 179 | 924 | PC | clr | 39.96 | 35.7 | 677 | 0.471 | 0.506 | 76536 | 88262 | 1005-1006Z: Planned power outage for mechanics training. No data impacted for instruments supported by UPS. 1010-1012Z: second planned power outage to switch back to main generator.
2025-06-29 | 180 | 1522 | PC | clr | 37.04 | 35.5 | 678.7 | 0.471 | 0.507 | 76833 | 88589
2025-06-30 | 181 | 1401 | PC | clr | 42.42 | 35.4 | 679.3 | 0.471 | 0.507 | 76541 | 88275 | 1518Z: Surface Ozone TSI filter changed one day early due to regular North/Low winds for 2 weeks.
2025-07-01 | 182 | 958 | PC | clr | 46.52 | 35.6 | 674.6 | 0.469 | 0.505 | 76500 | 88177 | Planned power outage due to generator issues last night, power swap from 1133-1135Z, again from 1224-1225Z. And again 1602-1604Z. Instruments supported by UPS.
2025-07-02 | 183 | 1006 | PC | rm lf | 38.37 | 35.4 | 673.9 | 0.47 | 0.504 | 76372 | 88016
2025-07-03 | 184 | 1128 | PC | clr | 43.09 | 35.6 | 673.4 | 0.469 | 0.504 | 76434 | 88015
2025-07-04 | 185 | 1056 | PC | clr | 41.53 | 35.5 | 674.1 | 0.469 | 0.504 | 76501 | 88110
2025-07-05 | 186 | 1631 | PC | clr | 31.74 | 35.5 | 675.7 | 0.47 | 0.505 | 76337 | 87920 | Northwinds operations: 1716-1735Z D6 Snow machine 953 953
2025-07-06 | 187 | 947 | PC | clr | 33 | 35.7 | 673.9 | 0.469 | 0.504 | 76449 | 88021
2025-07-07 | 188 | 1119 | PC | clr | 52.09 | 35.6 | 674.3 | 0.47 | 0.505 | 76350 | 87782 | 1602-1603Z: power swap back to main station generator. Instruments supported by UPS.
2025-07-08 | 189 | 1011 | PC | clr | 57.92 | 36.1 | 673.4 | 0.469 | 0.505 | 76343 | 87852
2025-07-09 | 190 | 1043 | PC | rm lf | 55.42 | 36 | 671 | 0.468 | 0.503 | 76547 | 88077 | LOW WINDS OPERATIONS 1015-1018Z: D6 HEAT CANNON, 5K
2025-07-10 | 191 | 1007 | PC | CLR | 48.14 | 35.7 | 669.4 | 0.467 | 0.502 | 76354 | 87802
2025-07-11 | 192 | 1101 | PC | rm lf | 36.49 | 35.3 | 677.8 | 0.471 | 0.507 | 76256 | 87709
2025-07-12 | 193 | 1028 | PC | rm lf | 43.76 | 35.7 | 680.3 | 0.472 | 0.508 | 76271 | 87698
2025-07-13 | 194 | 1229 | PC | rm ice | 35.01 | 36.5 | 684.5 | 0.473 | 0.509 | 76667 | 88128
2025-07-14 | 195 | 1039 | PC | clr | 47.95 | 35.5 | 689.4 | 0.476 | 0.511 | 76271 | 87616
2025-07-15 | 196 | 1031 | PC | rm ice | 39.79 | 36 | 692.3 | 0.478 | 0.512 | 76440 | 87783 | Northwind ops: x2 variances (welding, medical to heat isolation bay during day) from 0900-1400
2025-07-16 | 197 | 1047 | PC | rm lf | 47.09 | 35.4 | 689.4 | 0.477 | 0.511 | 76161 | 87472
2025-07-17 | 198 | 1148 | PC | rm lf | 51.98 | 35.2 | 687.9 | 0.476 | 0.511 | 76105 | 87422
2025-07-18 | 199 | 1009 | PC | clr | 46.96 | 35.5 | 687 | 0.475 | 0.511 | 76198 | 87472
2025-07-19 | 200 | 1019 | PC | rm ice | 49.2 | 35.7 | 686.9 | 0.475 | 0.51 | 76199 | 87477
2025-07-20 | 201 | 1139 | PC | clr | 50.71 | 35.7 | 682.3 | 0.472 | 0.508 | 76251 | 87516
2025-07-21 | 202 | 1034 | PC | clr | 47.44 | 35.5 | 681 | 0.473 | 0.508 | 76117 | 87284
2025-07-22 | 203 | 1044 | PC | clr | 38.8 | 35.6 | 687 | 0.475 | 0.51 | 76132 | 87328 | Ozone cal: Level 0: 0.5, level 1: 170.4, level 2: 42.5
2025-07-23 | 204 | 1017 | PC | clr | 46.18 | 35.7 | 686.3 | 0.475 | 0.51 | 76174 | 87327 | Low winds from 0943-1226Z, 1803-1908Z, North winds 1206-1217Z, 1851-1942Z, 2052-2150Z
2025-07-24 | 205 | 1025 | PC | clr | 49.01 | 35.6 | 680.7 | 0.472 | 0.508 | 76147 | 87248 | Low winds from 0823-1429Z, 1643-1643Z. North winds from 0051-0402Z, 0932-1230Z
2025-07-25 | 206 | 1013 | PC | clr | 49.38 | 35.5 | 678.6 | 0.471 | 0.507 | 76109 | 87193
2025-07-26 | 207 | 1202 | PC | clr | 44.95 | 35.6 | 679.4 | 0.471 | 0.507 | 76104 | 87180 | Low winds from 0223-0224Z. North winds from 2314-0054Z
2025-07-27 | 208 | 1857 | PC | clr | 46.36 | 36.7 | 682.3 | 0.472 | 0.508 | 76432 | 87480 | Low winds from 0520-0520Z, 0722-0741Z, 2147-2357Z. North winds from 0239-0520Z, 2202-2340Z
2025-07-28 | 209 | 1035 | PC | clr | 44.77 | 35.5 | 679.9 | 0.472 | 0.508 | 76013 | 87044
2025-07-29 | 210 | 1042 | PC | clr | 47.72 | 35.3 | 679.1 | 0.471 | 0.507 | 75988 | 86982 | Low winds 09:17Z to 09:56Z. Surface ozone cal 1131-1217Z: zero: 0.2, level 1: 168.8, level 2: 41.8
2025-07-30 | 211 | 1026 | PC | clr | 46.19 | 35.5 | 681 | 0.473 | 0.507 | 75985 | 86960 | Low winds 02:10Z to 02:11Z, north winds 09:19Z to 09:25Z 11:47Z to 13:28Z 19:54Z to 21:04Z
2025-07-31 | 212 | 1524 | PC | clr | 48.95 | 35.2 | 682.2 | 0.473 | 0.509 | 75892 | 86867
2025-08-01 | 213 | 1201 | PC | rm lf | 45.89 | 35.3 | 678.3 | 0.472 | 0.507 | 75855 | 86751
2025-08-02 | 214 | 1452 | PC | rm lf | 41.25 | 35.5 | 674.2 | 0.469 | 0.505 | 75921 | 86850
2025-08-03 | 215 | 1125 | PC | clr | 35.79 | 35.2 | 671.3 | 0.468 | 0.504 | 75862 | 86766
2025-08-04 | 216 | 1503 | PC | clr | 38.46 | 35.3 | 674.4 | 0.47 | 0.506 | 75878 | 86747
2025-08-05 | 217 | 958 | PC | rm lf | 40.35 | 35.9 | 675.4 | 0.469 | 0.505 | 75971 | 86812 | Surface ozone calibration: zero: 0.2, level 1: 171.1, level 2: 43.1
2025-08-06 | 218 | 1127 | PC | CLR | 41.04 | 35.6 | 675.7 | 0.469 | 0.506 | 75795 | 86621
2025-08-07 | 219 | 1058 | PC | clr | 37.14 | 35.5 | 676.2 | 0.47 | 0.506 | 75998 | 86810
2025-08-08 | 220 | 1006 | PC | RM LF | 49.6 | 36 | 677.1 | 0.471 | 0.506 | 75937 | 86717
2025-08-09 | 221 | 1226 | PC | CLR | 46.33 | 35.6 | 677.8 | 0.471 | 0.507 | 75912 | 86645
2025-08-10 | 222 | 1404 | PC | clr | 48.56 | 36.2 | 678.5 | 0.471 | 0.507 | 76053 | 86811
2025-08-11 | 223 | 1108 | PC | clr | 47.21 | 35.4 | 679.3 | 0.472 | 0.507 | 75765 | 86494
2025-08-12 | 224 | 1503 | PC | clr | 42.78 | 35.5 | 681.5 | 0.474 | 0.508 | 76013 | 86710 | Surface ozone calibration 1314Z: zero: 0.2, level 1: 169.4, level 2: 42.0
2025-08-13 | 225 | 1202 | PC | rm lf | 44.57 | 34.8 | 676.1 | 0.471 | 0.506 | 75379 | 85973
2025-08-14 | 226 | 1022 | PC | rm lf | 42.39 | 35.1 | 679 | 0.472 | 0.508 | 75630 | 86243
2025-08-15 | 227 | 909 | PC | clr | 39.85 | 35.6 | 676.6 | 0.471 | 0.507 | 75633 | 86263
2025-08-16 | 228 | 1132 | PC | clr | 42.03 | 35.1 | 675.7 | 0.47 | 0.506 | 75516 | 86097
2025-08-17 | 229 | 1419 | PC | clr | 35.47 | 34.7 | 686.1 | 0.475 | 0.511 | 75535 | 86089
2025-08-18 | 230 | 1007 | PC | clr | 43.87 | 35.1 | 961.5 | 0.478 | 0.513 | 75707 | 86247
2025-08-19 | 231 | 1126 | PC | rm lf | 39.37 | 35.7 | 697.7 | 0.481 | 0.515 | 75700 | 86214 | Surface ozone calibration 1157-1242Z: zero: -0.4, level 1: 169.4, level 2: 42.3
2025-08-20 | 232 | 1210 | PC | rm lf | 39.11 | 37.9 | 697.3 | 0.478 | 0.514 | 76219 | 86786
2025-08-21 | 233 | 1052 | PC | rm lf | 53.77 | 38.7 | 693 | 0.476 | 0.512 | 76399 | 86930
2025-08-22 | 234 | 1056 | PC | clr | 46.72 | 38.7 | 685.9 | 0.473 | 0.508 | 76178 | 86635
2025-08-23 | 235 | 1624 | PC | clr | 53.07 | 38 | 680.3 | 0.469 | 0.506 | 76266 | 86747
2025-08-24 | 236 | 1012 | PC | clr | 36.35 | 36.9 | 677.4 | 0.469 | 0.505 | 75944 | 86331
2025-08-25 | 237 | 1300 | PC | clr | 40.07 | 37.6 | 675.3 | 0.468 | 0.504 | 76116 | 86485 | Planned power swap, instruments supported by ups 1002-1004Z
2025-08-26 | 238 | 1028 | PC | rm lf | 46.05 | 36.2 | 679.4 | 0.47 | 0.506 | 75775 | 86123 | Planned power swap, instruments supported by ups 1101-1102Z. Surface ozone calibration 1125-1211Z. Zero: 0.2, level 1: 169.1, level 2: 42.2
2025-08-27 | 239 | 954 | PC | clr | 42.47 | 35.5 | 680.1 | 0.471 | 0.506 | 75503 | 85790
2025-08-28 | 240 | 952 | PC | CLR | 37.69 | 37.3 | 674.1 | 0.467 | 0.503 | 75951 | 86243
2025-08-29 | 241 | 1141 | PC | rm lf | 44.76 | 35.3 | 675.4 | 0.469 | 0.505 | 75371 | 85590
2025-08-30 | 242 | 1033 | PC | clr | 39.45 | 36.1 | 677.1 | 0.469 | 0.505 | 75615 | 85844
2025-08-31 | 243 | 1313 | PC | clr | 34.73 | 37.3 | 681.4 | 0.471 | 0.507 | 75981 | 86183
2025-09-01 | 244 | 1058 | PC | RM LF | 33.72 | 36.6 | 681.3 | 0.471 | 0.506 | 75768 | 85950
2025-09-02 | 245 | 1110 | PC | RM LF | 40.93 | 36 | 677.8 | 0.469 | 0.505 | 75509 | 85664 | Surface ozone calibration 1123-1210Z: zero: -0.7, level 1: 169.7, level 2: 42.6
2025-09-03 | 246 | 1023 | PC | CLR | 42.49 | 35.3 | 675.4 | 0.469 | 0.505 | 75321 | 85424
2025-09-04 | 247 | 1058 | PC | CLR | 45.41 | 35.8 | 679.1 | 0.47 | 0.506 | 75548 | 85671
2025-09-05 | 248 | 1234 | PC | CLR | 42.92 | 35.5 | 681.5 | 0.472 | 0.507 | 75410 | 85470
2025-09-06 | 249 | 1037 | PC | CLR | 44.84 | 35.6 | 682.2 | 0.472 | 0.508 | 75405 | 85467
2025-09-07 | 250 | 1445 | PC | RM LF | 46.79 | 36.1 | 680.6 | 0.471 | 0.507 | 75476 | 85491
2025-09-08 | 251 | 1208 | PC | CLR | 44.08 | 35.7 | 678.9 | 0.47 | 0.506 | 75354 | 85374
2025-09-09 | 252 | 1047 | PC | CLR | 41.03 | 35.5 | 671.7 | 0.467 | 0.503 | 75334 | 85282
2025-09-10 | 253 | 1454 | PC | CLR | 38.34 | 35.6 | 675.1 | 0.468 | 0.504 | 75280 | 85242
2025-09-11 | 254 | 1042 | PC | CLR | 37.18 | 35.5 | 680.5 | 0.471 | 0.506 | 75232 | 85151
2025-09-12 | 255 | 1128 | PC | rm lf | 36.36 | 35.6 | 683.3 | 0.473 | 0.508 | 75295 | 85212
2025-09-13 | 256 | 1104 | PC | CLR | 45.65 | 35.9 | 681.8 | 0.472 | 0.507 | 75331 | 85249
2025-09-14 | 257 | 1527 | PC | clr | 47 | 36.6 | 680.9 | 0.471 | 0.506 | 75466 | 85339
2025-09-15 | 258 | 1109 | PC | rm lf | 47.77 | 35.6 | 683 | 0.472 | 0.508 | 75195 | 85036
2025-09-16 | 259 | 1147 | PC | rm lf | 41.73 | 35.5 | 686.1 | 0.474 | 0.509 | 75121 | 84962
2025-09-17 | 260 | 1448 | PC | clr | 36.91 | 35.5 | 686.5 | 0.474 | 0.51 | 75214 | 85017
2025-09-18 | 261 | 1001 | PC | RM LF | 38.57 | 35.5 | 686.2 | 0.474 | 0.51 | 75228 | 85013
2025-09-19 | 262 | 1128 | PC | RM LF | 45.26 | 35.5 | 687.3 | 0.475 | 0.51 | 75230 | 84981
2025-09-20 | 263 | 1118 | PC | RM LF | 44.22 | 35.7 | 683.5 | 0.473 | 0.508 | 75129 | 84879
2025-09-21 | 264 | 1143 | PC | CLR | 42.77 | 36 | 679.9 | 0.471 | 0.507 | 75260 | 84975
2025-09-22 | 265 | 1230 | PC | rm lf | 42.46 | 36.8 | 678.5 | 0.469 | 0.505 | 75509 | 85268
2025-09-23 | 266 | 1127 | PC | RM LF | 40.2 | 37.6 | 678.7 | 0.469 | 0.505 | 75648 | 85372
2025-09-24 | 267 | 1138 | PC | RM LF | 39.4 | 36.4 | 677.9 | 0.469 | 0.506 | 75302 | 84998
2025-09-25 | 268 | 1120 | PC | RM LF | 39.21 | 36.1 | 670.2 | 0.466 | 0.502 | 75201 | 84841
2025-09-26 | 269 | 1134 | PC | RM LF | 39.22 | 36.3 | 666.5 | 0.464 | 0.5 | 75308 | 84818
2025-09-27 | 270 | 1212 | PC | RM LF | 46.93 | 36.1 | 668.3 | 0.465 | 0.501 | 75172 | 84786
2025-09-28 | 271 | 1451 | PC | RM LF | 41.48 | 36 | 669.1 | 0.466 | 0.502 | 75072 | 84675
2025-09-29 | 272 | 1247 | PC | RM LF | 37.81 | 36.1 | 673.3 | 0.468 | 0.503 | 75219 | 84781
2025-09-30 | 273 | 1417 | PC | CLR | 44.57 | 36.3 | 669.9 | 0.465 | 0.502 | 75305 | 84876
2025-10-01 | 274 | 1417 | PC | CLR | 38.56 | 35.9 | 667 | 0.464 | 0.501 | 75112 | 84657
2025-10-02 | 275 | 1200 | PC | RM LF | 37.02 | 36.2 | 664.3 | 0.463 | 0.499 | 75194 | 84709
2025-10-03 | 276 | 1119 | PC | RM LF | 39.2 | 36.2 | 662.9 | 0.463 | 0.498 | 75129 | 84604
2025-10-04 | 277 | 1149 | PC | RM LF | 38.85 | 36.6 | 663.9 | 0.463 | 0.499 | 75202 | 84685
2025-10-05 | 278 | 1619 | PC | RM LF | 35.01 | 36.6 | 664.7 | 0.463 | 0.5 | 75292 | 84779
2025-10-06 | 279 | 1222 | PC | RM LF | 34.61 | 36.1 | 664.2 | 0.463 | 0.5 | 75101 | 84557
2025-10-07 | 280 | 1143 | PC | RM LF | 37.29 | 36.1 | 663.1 | 0.463 | 0.499 | 75078 | 84488
2025-10-08 | 281 | 1155 | PC | - | 36.7 | 36.5 | 659.9 | 0.461 | 0.497 | 75209 | 84615 | CON-1 did not go to AWO
2025-10-09 | 282 | 1443 | PC | RM LF | 36.68 | 36 | 666.6 | 0.464 | 0.501 | 75046 | 84426
2025-10-10 | 283 | 1113 | PC | RM LF | 29.87 | 36.1 | 669.5 | 0.466 | 0.502 | 75109 | 84457
2025-10-11 | 284 | 1150 | PC | RM LF | 40.69 | 35.3 | 674.5 | 0.468 | 0.505 | 74776 | 84100
2025-10-12 | 285 | 1238 | GM, PC | CLR | 36.78 | 35 | 672.7 | 0.468 | 0.504 | 74742 | 84043
2025-10-13 | 286 | 1006 | GM, PC | RM LF | 34.27 | 35.2 | 674.3 | 0.468 | 0.504 | 74836 | 84132
2025-10-14 | 287 | 1021 | GM, PC | - | 31.22 | 35.6 | 670.2 | 0.466 | 0.502 | 74866 | 84141 | Weekly calibration: start:1130Z, end:1216Z, zero:-0.2, level 1:170.3, level 2:43.1, too windy to go on roof
2025-10-15 | 288 | 1630 | GM, PC | RM LF | 39.84 | 35.4 | 682.3 | 0.472 | 0.507 | 74800 | 84021
2025-10-16 | 289 | 1017 | GM | CLR | 46.24 | 35.2 | 686.2 | 0.474 | 0.509 | 74853 | 84091
2025-10-17 | 290 | 1109 | GM | CLR | 40.06 | 35.4 | 681.3 | 0.471 | 0.507 | 74851 | 84056
2025-10-18 | 291 | 1236 | GM | rm lf | 42.02 | 35.3 | 681.4 | 0.471 | 0.507 | 74559 | 83756
2025-10-19 | 292 | 1355 | GM | rm lf | 37 | 35.5 | 678.7 | 0.471 | 0.506 | 74739 | 83910
2025-10-20 | 293 | 1133 | GM | rm lf | 38.22 | 35.2 | 677.9 | 0.47 | 0.506 | 74538 | 83681
2025-10-21 | 294 | 1137 | GM | rm lf | 43.98 | 35.4 | 678.2 | 0.47 | 0.506 | 74781 | 83922 | Weekly calibration: start:1241Z, end:1326Z, Zero:0.2, Level 1:170.8, Level 2:42.7
2025-10-22 | 295 | 1652 | GM | clr | 40.76 | 35.5 | 675 | 0.469 | 0.505 | 74713 | 83837
2025-10-23 | 296 | 1248 | GM | rm lf | 33.01 | 35.6 | 676.9 | 0.47 | 0.505 | 74674 | 83784
2025-10-24 | 297 | 1103 | GM | rm lf | 40.39 | 35.3 | 678.9 | 0.47 | 0.506 | 74565 | 83643
2025-10-25 | 298 | 1101 | GM | clr | 45.34 | 35.8 | 670.5 | 0.466 | 0.503 | 74648 | 83736
2025-10-26 | 299 | 1328 | GM | rm lf | 31.77 | 35.8 | 671.1 | 0.467 | 0.503 | 74636 | 83699
2025-10-27 | 300 | 1548 | GM | rm lf | 42.9 | 35.8 | 671.3 | 0.467 | 0.503 | 74562 | 83606
2025-10-28 | 301 | 1543 | GM | rm lf | 35.47 | 36 | 672.5 | 0.467 | 0.503 | 74620 | 83642 | Weekly Calibration: Start: 1330Z, End: 1415Z, Zero: -0.0, Level 1: 171.8, Level 2: 43.5
2025-10-29 | 302 | 1346 | GM | rm lf | 39.62 | 35.9 | 671.4 | 0.466 | 0.503 | 74871 | 83889
2025-10-30 | 303 | 1453 | GM | rm lf | 34.54 | 35.8 | 674.1 | 0.468 | 0.503 | 74627 | 83584
2025-10-31 | 304 | 1647 | GM | rm lf | 37.55 | 35.9 | 673 | 0.467 | 0.503 | 74655 | 83639
2025-11-01 | 305 | 1347 | GM | rm hf | 39.01 | 36 | 670.9 | 0.467 | 0.502 | 74706 | 83648
2025-11-02 | 306 | 1441 | GM | clr | 39.8 | 35.9 | 664.7 | 0.464 | 0.5 | 74494 | 83382
2025-11-03 | 307 | 1943 | GM | - | 40.86 | 36.1 | 662.1 | 0.462 | 0.498 | 74665 | 83573 | No in person checks due to ICESat
2025-11-04 | 308 | 1252 | GM | rm lf | 40.51 | 36.1 | 664.5 | 0.463 | 0.499 | 74604 | 83482 | Weekly calibration: Start: 1326Z, End: 1412Z, Zero: -0.1, Level1: 172.6, Level2: 43.9
2025-11-05 | 309 | 1751 | GM | rm lf | 41.4 | 35.9 | 672.9 | 0.467 | 0.503 | 74508 | 83365
2025-11-06 | 310 | 1153 | HB | rm lf | 44.67 | 35.9 | 676.4 | 0.469 | 0.505 | 74523 | 83371
2025-11-07 | 311 | 1623 | GM | rm lf | 40.25 | 35.8 | 679.9 | 0.471 | 0.506 | 74407 | 83240
2025-11-08 | 312 | 1310 | GM | rm lf | 47.3 | 35.7 | 682.3 | 0.472 | 0.507 | 74437 | 83243
2025-11-09 | 313 | 1408 | HB | rm lf
2025-11-10 | 314 | 1252 | GM | rm lf | 45.11 | 35.9 | 681.8 | 0.472 | 0.507 | 74406 | 83139 | Weekly calibration: Start: 1419Z, End: , Zero: -0.6, Level 1: 171.6, Level2: 43.3
2025-11-11 | 315 | 1813 | GM | - | - | - | - | - | - | - | - | Did not go to AWO due to ICESat and could not login to FORGE acquisition webpage
2025-11-12 | 316 | 1714 | GM | rm lf | 43.6 | 35.7 | 681.5 | 0.471 | 0.507 | 74324 | 83033
2025-11-13 | 317 | 1310 | GM | clr | 49.16 | 35.9 | 683.1 | 0.472 | 0.508 | 74367 | 83067
2025-11-14 | 318 | 1400 | GM | lf | 46.45 | 35.7 | 680.6 | 0.471 | 0.507 | 74304 | 82977
2025-11-15 | 319 | 1254 | GM | rm lf | 44.39 | 35.7 | 685.3 | 0.473 | 0.508 | 74360 | 83034
2025-11-16 | 320 | 1341 | GM | rm lf | 42.71 | 35.8 | 683.4 | 0.472 | 0.508 | 74366 | 83045
2025-11-17 | 321 | 1706 | GM | rm lf | 49.6 | 35.2 | 688.1 | 0.475 | 0.51 | 74085 | 82698
2025-11-18 | 322 | 1434 | GM | rm lf | 46.64 | 35.6 | 687.3 | 0.475 | 0.51 | 74263 | 82845 | Weekly calibration: Start: 1308Z, End: 1353Z, Zero: -0.4, Level 1: 170.3, Level 2: 43.3
2025-11-19 | 323 | 1405 | GM | rm lf | 49.05 | 35.6 | 682.1 | 0.472 | 0.507 | 74271 | 82839
2025-11-20 | 324 | 1412 | GM | rm lf | 46.84 | 35.7 | 674.2 | 0.468 | 0.504 | 74253 | 82851
2025-11-21 | 325 |  | GM | rm lf | [7 empty] | Did not get these values
2025-11-22 | 326 | 1235 | GM | rm lf | 42.95 | 35.9 | 669.5 | 0.466 | 0.501 | 74208 | 82721
2025-11-23 | 327 | 1349 | GM | rm lf | 44.59 | 35.8 | 671.9 | 0.467 | 0.503 | 74165 | 82700
2025-11-24 | 328 | 1630 | GM | rm lf | 44.57 | 35.8 | 666.3 | 0.464 | 0.5 | 74116 | 82595
2025-11-25 | 329 | 1327 | GM | rm lf | 44.07 | 36.1 | 661.3 | 0.462 | 0.498 | 74188 | 82665 | Weekly calibration: Start: 1332Z, End: 1418Z, Zero: 0.0, Level 1: 171.4, Level 2: 44.0
2025-11-26 | 330 | 1748 | GM | rm lf | 39.88 | 35.7 | 661.4 | 0.462 | 0.498 | 74168 | 82610
2025-11-27 | 331 | [10 empty] | Did not go out to AWO and do not have remote access to FORGE system
2025-11-28 | 332 | 1329 | GM | rm lf | 40.61 | 35.7 | 661.3 | 0.462 | 0.498 | 74096 | 82506
2025-11-29 | 333 | 1306 | GM | rm lf | 39.63 | 35.6 | 663.4 | 0.463 | 0.499 | 74265 | 82703
2025-11-30 | 334 | 1310 | GM | rm lf | 42.05 | 35.6 | 661.4 | 0.462 | 0.498 | 74191 | 82625
2025-12-01 | 335 | 1700 | GM | rm lf | 42.27 | 35.9 | 664.5 | 0.463 | 0.5 | 74116 | 82523
2025-12-02 | 336 | 1829 | GM | - | 43 | 35.8 | 666.3 | 0.464 | 0.5 | 74048 | 82404 | Did not go out to AWO due to ICESat
2025-12-03 | 337 | 1739 | GM | rm lf | 45.64 | 35.2 | 670.5 | 0.467 | 0.502 | 73951 | 82274 | Weekly calibration: Start:1741Z, End: 1827Z, Zero: 0.1, Level1: 172.2, Level 2: 43.2
2025-12-04 | 338 | 1458 | GM | rm lf | 46.14 | 35.7 | 672.1 | 0.467 | 0.503 | 74161 | 82471
2025-12-05 | 339 | 1857 | GM | rm hf | 39.14 | 35.9 | 676.5 | 0.469 | 0.505 | 74070 | 82351
2025-12-06 | 340 | 1304 | GM | rm lf | 40.29 | 35.8 | 677.4 | 0.47 | 0.505 | 73988 | 82268
2025-12-07 | 341 | 1411 | GM | rm lf | 42.15 | 35.7 | 675.5 | 0.469 | 0.504 | 74005 | 82249
2025-12-08 | 342 | 1730 | GM | rm lf | 47.92 | 35.6 | 673.7 | 0.468 | 0.503 | 73931 | 82147
2025-12-09 | 343 | 1321 | HB | rm lf | 41.79 | 35.5 | 671 | 0.467 | 0.502 | 73985 | 82231 | Weekly calibration: Start:1444Z, End:1529Z, Zero:-0.0, Level 1:173.2, Level 2:43.5
2025-12-10 | 344 | 1757 | GM | rm lf | 43.01 | 35.8 | 666.6 | 0.464 | 0.5 | 73986 | 82206
2025-12-11 | 345 | 1345 | GM | rm lf | 43.73 | 35.9 | 668.5 | 0.465 | 0.502 | 73899 | 82091
2025-12-12 | 346 | 1601 | GM | - | 41.69 | 36.1 | 660.2 | 0.461 | 0.497 | 73989 | 82173
2025-12-13 | 347 | 1301 | GM | clr | 41.46 | 36 | 657.1 | 0.46 | 0.496 | 73923 | 82104
2025-12-14 | 348 | 1256 | HB | rm lf | 44.03 | 36.1 | 657 | 0.459 | 0.496 | 74048 | 82190
2025-12-15 | 349 | 1736 | GM | rm lf | 43.69 | 36.1 | 657.9 | 0.459 | 0.496 | 74143 | 82274
2025-12-16 | 350 | 1426 | GM | - | 38.85 | 36.4 | 655.8 | 0.459 | 0.495 | 74074 | 82166 | Weekly calibration: Start:1429Z, End:1514Z, Zero:0.4, Level1:172.0, Level2:44.5
2025-12-17 | 351 | 1628 | GM | rm lf | 39.47 | 36.5 | 659.4 | 0.461 | 0.496 | 73948 | 82017
2025-12-18 | 352 | 1324 | GM | rm lf | 34.75 | 36.2 | 663 | 0.462 | 0.498 | 74020 | 82075
2025-12-19 | 353 | 1318 | GM | rm lf | 39.24 | 36.2 | 661.9 | 0.461 | 0.498 | 74067 | 82156
2025-12-20 | 354 | 1241 | GM | rm lf | 41.47 | 36.3 | 666.2 | 0.463 | 0.5 | 73929 | 81956
2025-12-21 | 355 | 1224 | GM | - | 46.5 | 35.9 | 662.5 | 0.462 | 0.498 | 73793 | 81796 | Did not go to AWO due to Con 1
2025-12-22 | 356 | 1519 | GM | clr | 41.34 | 35.8 | 665.8 | 0.464 | 0.5 | 73800 | 81782
2025-12-23 | 357 | 1343 | GM | clr | 42.16 | 35.7 | 660.3 | 0.461 | 0.497 | 73690 | 81673 | Weekly calibration: Start:1420Z, End:1505Z, Zero:-0.3, Level 1:171.4, Level 2:43.7
2025-12-24 | 358 | 1246 | GM | rm lf | 42.33 | 35.8 | 659.5 | 0.461 | 0.497 | 73880 | 81870
2025-12-25 | 359 | 1814 | GM | - | 45.94 | 36 | 664.7 | 0.463 | 0.499 | 73782 | 81719 | Did not go out to AWO
2025-12-26 | 360 | 1717 | GM | rm lf | 37.69 | 35.3 | 675.8 | 0.469 | 0.504 | 73712 | 81633
2025-12-27 | 361 | 1245 | GM | - | 42.24 | 34.8 | 668.7 | 0.466 | 0.501 | 73427 | 81317 | Did not go out to AWO due to Con 1
2025-12-28 | 362 | 1306 | GM | rm lf | 39.93 | 35.4 | 681.4 | 0.472 | 0.507 | 73533 | 81428
2025-12-29 | 363 | 1454 | GM | rm lf | 45.75 | 35.4 | 682.6 | 0.472 | 0.508 | 73594 | 81486
2025-12-30 | 364 | 1315 | GM | clr | 40.64 | 35.7 | 677 | 0.47 | 0.505 | 73500 | 81320 | Weekly calibration: Start: 1357Z, End: 1442Z, Zero: 0.3, Level 1: 170.1, Level 2: 43.1
2025-12-31 | 365 | 1221 | GM | - | 41.72 | 35.2 | 684.5 | 0.473 | 0.508 | 73527 | 81361 | Did not go out to AWO
Note on F1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: G81 Thermo Box: Ozone (ppb)
Note on G1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Sample (C)
Note on H1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Sample (hPa) Reply: If pressure off from atmospheric pressure, inlet is plugged
Note on I1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Cell A (lpm)
Note on J1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Cell B (lpm)
Note on K1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Cell A (Hz)
Note on L1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Cell B (Hz)