## Logsheet
Year | Month | Day | DOY | DOY since 2012 | Science Tech Initials | H2 Gen. Act Pressure (PSI) | H2 Generator Water Level | H2 Gen. Flow (ml/min) | Methane Std 1 Pressure | Methane Std 2 Pressure |  | Inlet Line Flow Rate (L/min) | 4 Pumps On? |  | NMHC Std 1 Pressure | NMHC Std 2 Pressure | Trap Housing Purge Flow Rate (ml/min) | Picarro Purge Flow (ml/min) | Picarro dessicant | Picarro CC Standard (High) Pressure | Picarro CB Standard (Mid) Pressure | Picarro CA Standard (Low) Pressure | Graphs Look Good? | (Weekly) H2 Gen Desiccant Check | (Weekly) H2 Line Filter Desiccant Check | Daily Oxygen indicator (green) | H2 Cylinder pressure | Comments (Inlet Cleaning, Inlet Filter Changes, Water Refill, etc.
2012 | 4 | 26 | 117 | 117
2012 | 4 | 27 | 118 | 118 | [23 empty] | Refilled H2 gen to Max
2012 | 4 | 28 | 119 | 119 | JH | 60 |  | 54 | 2000 | 1925 |  | 2 |  |  | 1200 | 1575
2012 | 4 | 29 | 120 | 120 | AC | 60 |  | 47 | 2000 | 1925 |  | 2 |  |  | 1200 | 1575
2012 | 4 | 30 | 121 | 121 | AC | 60 |  | 51 | 2000 | 1950 |  | 2 |  |  | 1200 | 1600
2012 | 5 | 1 | 122 | 122 | AC | 60 |  | 50 | 2000 | 1900 |  | 2 |  |  | 1175 | 1550 | [11 empty] | Room temp now stable at 20C
2012 | 5 | 2 | 123 | 123 | EM | 60 |  | 50 | 2000 | 1925 |  | 2 |  |  | 1175 | 1575
2012 | 5 | 3 | 124 | 124 | AC | 60 |  | 52 | 2000 | 1900 |  | 2 |  |  | 1175 | 1575
2012 | 5 | 4 | 125 | 125 | AC | 60 |  | 50 | 2000 | 1900 |  | 2 |  |  | 1175 | 1575 | [11 empty] | Raised inlet on tower 1830-1935UTC
2012 | 5 | 5 | 126 | 126 | EM | 60 |  | 50 | 1990 | 1925 |  | 2 |  |  | 1175 | 1575
2012 | 5 | 6 | 127 | 127
2012 | 5 | 7 | 128 | 128 | AC | 60 |  | 52 | 2000 | 1925 |  | 2 |  |  | 1175 | 1575 | [11 empty] | Added ~1L water
2012 | 5 | 8 | 129 | 129 | AC | 60 |  | 53 | 2000 | 1925 |  | 2 |  |  | 1175 | 1575
2012 | 5 | 9 | 130 | 130 | AC | 60 |  | 50 | 2000 | 1925 |  | 2 |  |  | 1175 | 1575
2012 | 5 | 10 | 131 | 131 | AC | 60 |  | 51 | 2000 | 1925 |  | 2 |  |  | 1175 | 1575
2012 | 5 | 11 | 132 | 132 | AC | 60 |  | 51 | 2000 | 1925 |  | 2 |  |  | 1175 | 1575
2012 | 5 | 12 | 133 | 133 | AC/EM | 60 |  | 50 | 1990 | 1920 |  | 2 |  |  | 1175 | 1575
2012 | 5 | 13 | 134 | 134
2012 | 5 | 14 | 135 | 135 | EM/AC | 60 |  | 53 | 2000 | 1925 |  | 2 |  |  | 1175 | 1575 | [11 empty] | 1800utc going again after rebooting and Jacques restarting system remotely.
2012 | 5 | 15 | 136 | 136 | EM | 60 |  | 51 | 1990 | 1925 |  | 2 |  |  | 1175 | 1575
2012 | 5 | 16 | 137 | 137 | EM | 60 |  | 51 | 2000 | 1925 |  | 2 |  |  | 1175 | 1575
2012 | 5 | 17 | 138 | 138 | EM | 60 |  | 53 | 2000 | 1925 |  | 2 |  |  | 1175 | 1575
2012 | 5 | 18 | 139 | 139 | AC | 60 |  | 53 | 2000 | 1925 |  | 2 |  |  | 1175 | 1575
2012 | 5 | 19 | 140 | 140 | AC | 60 |  | 50 | 2000 | 1925 |  | 2 |  |  | 1175 | 1575
2012 | 5 | 20 | 141 | 141 | EM | 60 |  | 50 | 1975 | 1875 |  | 2 |  |  | 1150 | 1550
2012 | 5 | 21 | 142 | 142 | EM | 60 |  | 50 | 1975 | 1875 |  | 2 |  |  | 1150 | 1550
2012 | 5 | 22 | 143 | 143 | AC | 60 |  | 50 | 1975 | 1900 |  | 2 |  |  | 1175 | 1575
2012 | 5 | 23 | 144 | 144 | EM | 60 |  | 51 | 1975 | 1890 |  | 2 |  |  | 1160 | 1560
2012 | 5 | 24 | 145 | 145 | EM/AC | 60 |  | 50 | 1975 | 1900 |  | 2 |  |  | 1175 | 1575 | [11 empty] | 1838UTC: switched FID 1 gain to high filtered per DH request.
2012 | 5 | 25 | 146 | 146 | EM | 60 |  | 51 | 1975 | 1900 |  | 2 |  |  | 1160 | 1525
2012 | 5 | 26 | 147 | 147 | EM/AC | 60 |  | 49 | 2000 | 1925 |  | 2 |  |  | 1175 | 1575
2012 | 5 | 27 | 148 | 148 | EM | 60 |  | 53 | 2000 | 1925 |  | 2 |  |  | 1175 | 1575
2012 | 5 | 28 | 149 | 149 | EM | 60 |  | 52 | 1975 | 1925 |  | 2 |  |  | 1175 | 1575 | [11 empty] | 1306UTC: added 1L water
2012 | 5 | 29 | 150 | 150 | EM | 60 |  | 51 | 1975 | 1925 |  | 2 |  |  | 1175 | 1575
2012 | 5 | 30 | 151 | 151 | AC | 60 |  | 51 | 2000 | 1925 |  | 2 |  |  | 1175 | 1575 | [11 empty] | Station power outage 0315-0328utc
2012 | 5 | 31 | 152 | 152 | EM | 60 |  | 51 | 1975 | 1900 |  | 2 |  |  | 1175 | 1575 | [11 empty] | 1313UTC: switched FID 1 gain back to high (unfiltered) per DH request
2012 | 6 | 1 | 153 | 153 | AC | 60 |  | 53 | 1975 | 1900 |  | 2 |  |  | 1175 | 1575
2012 | 6 | 2 | 154 | 154 | EM | 60 |  | 51 | 1975 | 1900 |  | 2 |  |  | 1175 | 1575
2012 | 6 | 3 | 155 | 155 | EM | 60 |  | 52 | 1975 | 1900 |  | 2 |  |  | 1175 | 1575
2012 | 6 | 4 | 156 | 156 | EM | 60 |  | 52 | 1975 | 1900 |  | 2 |  |  | 1160 | 1560
2012 | 6 | 5 | 157 | 157 | AC | 60 |  | 52 | 1975 | 1900 |  | 2 |  |  | 1160 | 1550
2012 | 6 | 6 | 158 | 158 | EM | 60 |  | 51 | 1950 | 1875 |  | 2 |  |  | 1150 | 1550
2012 | 6 | 7 | 159 | 159 | EM | 60 |  | 53 | 1925 | 1875 |  | 2 |  |  | 1125 | 1525 | [11 empty] | Room temp cooler
2012 | 6 | 8 | 160 | 160 | AC | 60 |  | 49 | 1950 | 1875 |  | 2 |  |  | 1150 | 1550 | [11 empty] | *1902-1904UTC Power outage to switch to backup gen. On UPS for the two min.
2012 | 6 | 9 | 161 | 161 | EM | 60 |  | 53 | 1950 | 1875 |  | 2 |  |  | 1150 | 1550 | [11 empty] | *1526-1528UTC Power outage to switch back to normal power. On UPS for the two min. 1634-1636 power outage again. Working to restore stable power.
2012 | 6 | 10 | 162 | 162 | EM | 60 |  | 53 | 1950 | 1875 |  | 2 |  |  | 1150 | 1550
2012 | 6 | 11 | 163 | 163 | AC/EM | 60 |  | 52 | 1950 | 1875 |  | 2 |  |  | 1150 | 1550 | [11 empty] | Installed spare peltier cooler 1745
2012 | 6 | 12 | 164 | 164 | EM | 60 |  | 51 | 1950 | 1875 |  | 2 |  |  | 1150 | 1550
2012 | 6 | 13 | 165 | 165
2012 | 6 | 14 | 166 | 166
2012 | 6 | 15 | 167 | 167 | ac | 60 |  | 52 | 1950 | 1875 |  | 2 |  |  | 1150 | 1550
2012 | 6 | 16 | 168 | 168 | ac | 60 |  | 52 | 1975 | 1900 |  | 2 |  |  | 1175 | 1575
2012 | 6 | 17 | 169 | 169 | em | 60 |  | 53 | 1975 | 1900 |  | 2 |  |  | 1175 | 1575
2012 | 6 | 18 | 170 | 170 | ac | 60 |  | 52 | 1975 | 1900 |  | 2 |  |  | 1150 | 1550 | [11 empty] | 1225UTC Refilled H generator ~1L DI water
2012 | 6 | 19 | 171 | 171 | em | 60 |  | 51 | 1975 | 1900 |  | 2 |  |  | 1150 | 1560
2012 | 6 | 20 | 172 | 172 | ac | 60 |  | 51 | 1975 | 1900 |  | 2 |  |  | 1150 | 1550
2012 | 6 | 21 | 173 | 173 | em | 60 |  | 50 | 1975 | 1900 |  | 2 |  |  | 1150 | 1560
2012 | 6 | 22 | 174 | 174 | em | 60 |  | 50 | 1975 | 1900 |  | 2 |  |  | 1150 | 1560
2012 | 6 | 23 | 175 | 175 | AC | 60 |  | 50 | 1975 | 1875 |  | 2 |  |  | 1140 | 1550
2012 | 6 | 24 | 176 | 176 | ac | 60 |  | 53 | 1950 | 1875 |  | 2 |  |  | 1150 | 1550
2012 | 6 | 25 | 177 | 177 | em | 60 |  | 53 | 1950 | 1875 |  | 2 |  |  | 1150 | 1550
2012 | 6 | 26 | 178 | 178 | ac | 60 |  | 52 | 1950 | 1875 |  | 2 |  |  | 1150 | 1550
2012 | 6 | 27 | 179 | 179 | em | 60 |  | 49 | 1950 | 1875 |  | 2 |  |  | 1125 | 1540
2012 | 6 | 28 | 180 | 180 | em | 60 |  | 51 | 1960 | 1900 |  | 2 |  |  | 1140 | 1550
2012 | 6 | 29 | 181 | 181 | ac | 60 |  | 51 | 1975 | 1900 |  | 2 |  |  | 1150 | 1550
2012 | 6 | 30 | 182 | 182 | em | 60 |  | 53 | 2000 | 1950 |  | 2 |  |  | 1150 | 1575 | [11 empty] | Warmer in TAWO - Kramer instrument online
2012 | 7 | 1 | 183 | 183 | em | 60 |  | 51 | 2000 | 1950 |  | 2 |  |  | 1150 | 1575
2012 | 7 | 2 | 184 | 184 | ac | 60 |  | 50 | 2000 | 1925 |  | 2 |  |  | 1150 | 1550
2012 | 7 | 3 | 185 | 185 | em | 60 |  | 50 | 1975 | 1925 |  | 2 |  |  | 1125 | 1550
2012 | 7 | 4 | 186 | 186 | ac | 60 |  | 50 | 2000 | 1925 |  | 2 |  |  | 1125 | 1550
2012 | 7 | 5 | 187 | 187 | em | 60 |  | 50 | 2000 | 1925 |  | 2 |  |  | 1125 | 1560
2012 | 7 | 6 | 188 | 188 | ac | 60 |  | 52 | 2000 | 1950 |  | 2 |  |  | 1125 | 1550
2012 | 7 | 7 | 189 | 189 | ac | 60 |  | 54 | 2000 | 1950 |  | 2 |  |  | 1150 | 1600 | [11 empty] | BLD temps over 30C. Working on getting temps down and stable. Adc
2012 | 7 | 8 | 190 | 190 | em | 60 |  | 50 | 1960 | 1900 |  | 2 |  |  | 1110 | 1525
2012 | 7 | 9 | 191 | 191 | ac | 60 |  | 52 | 2000 | 1925 |  | 2 |  |  | 1125 | 1550
2012 | 7 | 10 | 192 | 192 | em | 60 |  | 52 | 2000 | 1950 |  | 2 |  |  | 1150 | 1550
2012 | 7 | 11 | 193 | 193 | em | 60 |  | 49 | 1975 | 1900 |  | 2 |  |  | 1125 | 1550
2012 | 7 | 12 | 194 | 194 | ac | 60 |  | 51 | 1950 | 1900 |  | 2 |  |  | 1100 | 1550
2012 | 7 | 13 | 195 | 195 | em | 60 |  | 50 | 2025 | 1975 |  | 2 |  |  | 1160 | 1600
2012 | 7 | 14 | 196 | 196 | em | 60 |  | 52 | 2000 | 1950 |  | 2 |  |  | 1140 | 1560 | [11 empty] | Added 1.3L DI water to H generator.
2012 | 7 | 15 | 197 | 197
2012 | 7 | 16 | 198 | 198 | ac | 60 |  | 50 | 2000 | 1950 |  | 2 |  |  | 1125 | 1550
2012 | 7 | 17 | 199 | 199 | em | 60 |  | 50 | 2000 | 1950 |  | 2 |  |  | 1125 | 1550
2012 | 7 | 18 | 200 | 200 | em | 60 |  | 51 | 1975 | 1900 |  | 2 |  |  | 1100 | 1525
2012 | 7 | 19 | 201 | 201 | em | 60 |  | 53 | 1975 | 1900 |  | 2 |  |  | 1100 | 1525
2012 | 7 | 20 | 202 | 202 | em | 60 |  | 53 | 1975 | 1925 |  | 2 |  |  | 1120 | 1540
2012 | 7 | 21 | 203 | 203 | em | 60 |  | 51 | 1975 | 1925 |  | 2 |  |  | 1125 | 1550
2012 | 7 | 22 | 204 | 204 | em | 60 |  | 52 | 2000 | 1950 |  | 2 |  |  | 1125 | 1550
2012 | 7 | 23 | 205 | 205 | em | 60 |  | 53 | 2000 | 1940 |  | 2 |  |  | 1125 | 1550
2012 | 7 | 24 | 206 | 206 | ac | 60 |  | 48 | 2000 | 1925 |  | 2 |  |  | 1125 | 1550
2012 | 7 | 25 | 207 | 207 | ac | 60 |  | 52 | 2000 | 1945 |  | 2 |  |  | 1125 | 1550
2012 | 7 | 26 | 208 | 208 | ac | 60 |  | 51 | 2000 | 1940 |  | 2 |  |  | 1125 | 1550
2012 | 7 | 27 | 209 | 209 | em | 60 |  | 52 | 1975 | 1900 |  | 2 |  |  | 1100 | 1525
2012 | 7 | 28 | 210 | 210 | ac | 60 |  | 53 | 1975 | 1900 |  | 2 |  |  | 1100 | 1550
2012 | 7 | 29 | 211 | 211 | ac | 60 |  | 52 | 2000 | 1950 |  | 2 |  |  | 1100 | 1550
2012 | 7 | 30 | 212 | 212 | ac | 60 |  | 51 | 1975 | 1925 |  | 2 |  |  | 1120 | 1540 | [11 empty] | new ventilation system installed, kinks being worked out
2012 | 7 | 31 | 213 | 213
2012 | 8 | 1 | 214 | 214
2012 | 8 | 2 | 215 | 215 | em | 60 |  | 53 | 1925 | 1860 |  | 2 |  |  | 1120 | 1540 | [11 empty] | Troubleshooting NMHC gc problems with Detlev
2012 | 8 | 3 | 216 | 216 | ac | 60 |  | 50 | 1925 | 1850 |  | 2 |  |  | 1100 | 1525 | [11 empty] | "
2012 | 8 | 4 | 217 | 217 | ac | 60 |  | 51 | 1960 | 1850 |  | 2 |  |  | 1100 | 1550
2012 | 8 | 5 | 218 | 218 | ac | 60 |  | 52 | 1925 | 1850 |  | 2 |  |  | 1050 | 1525
2012 | 8 | 6 | 219 | 219 | em | 60 |  | 52 | 1925 | 1860 |  | 2 |  |  | 1110 | 1525
2012 | 8 | 7 | 220 | 220 | em | 60 |  | 52 | 1900 | 1850 |  | 2 |  |  | 1100 | 1525
2012 | 8 | 8 | 221 | 221 | [23 empty] | Power outage 1252-1314UTC, then 1min @1530
2012 | 8 | 9 | 222 | 222 | ac | 60 |  | 52 | 1900 | 1825 |  | 2 |  |  | 1100 | 1525
2012 | 8 | 10 | 223 | 223 | ac | 60 |  | 52 | 1900 | 1800 |  | 2 |  |  | 1100 | 1525
2012 | 8 | 11 | 224 | 224 | ac | 60 |  | 51 | 1900 | 1750 |  | 0 |  |  | 1100 | 1500 | [11 empty] | Instrument is shut down.
2012 | 8 | 12 | 225 | 225
2012 | 8 | 13 | 226 | 226 | em | 60 |  | 52 | 1925 | 1750 |  | 2 |  |  | 1100 | 1540
2012 | 8 | 14 | 227 | 227 | em | 60 |  | 52 | 1900 | 1750 |  | 2 |  |  | 1100 | 1525
2012 | 8 | 15 | 228 | 228 | em | 60 |  | 52 | 1925 | 1740 |  | 2 |  |  | 1100 | 1530
2012 | 8 | 16 | 229 | 229
2012 | 8 | 17 | 230 | 230
2012 | 8 | 18 | 231 | 231 | jh | 60 |  | 50 | 1950 | 1775 |  | 2 | y |  | 1100 | 1525
2012 | 8 | 19 | 232 | 232 | jm | 60 |  | 50 | 1950 | 1750 |  | 2 | y |  | 1100 | 1500 | [11 empty] | Refilled hydrogen generator
2012 | 8 | 20 | 233 | 233 | lr | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1100 | 1550
2012 | 8 | 21 | 234 | 234 | jm | 60 |  | 51 | 1950 | 1800 |  | 2 | y |  | 1100 | 1550
2012 | 8 | 22 | 235 | 235 | jm | 60 |  | 50 | 1950 | 1750 |  | 2 | y |  | 1100 | 1550
2012 | 8 | 23 | 236 | 236 | jm | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1100 | 1550
2012 | 8 | 24 | 237 | 237 | jm | 60 |  | 52 | 1950 | 1800 |  | 2 | y |  | 1100 | 1550
2012 | 8 | 25 | 238 | 238 | jm | 60 |  | 51 | 1950 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 8 | 26 | 239 | 239 | lr | 60 |  | 51 | 1950 | 1775 |  | 2 | y |  | 1100 | 1525
2012 | 8 | 27 | 240 | 240 | jm | 60 |  | 51 | 1950 | 1800 |  | 2 | y |  | 1100 | 1550
2012 | 8 | 28 | 241 | 241 | jm | 60 |  | 53 | 1950 | 1800 |  | 2 | y |  | 1100 | 1550
2012 | 8 | 29 | 242 | 242 | jm | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1100 | 1550
2012 | 8 | 30 | 243 | 243 | jm | 60 |  | 52 | 1950 | 1800 |  | 2 | y |  | 1100 | 1550
2012 | 8 | 31 | 244 | 244 | jm | 60 |  | 52 | 1950 | 1800 |  | 2 | y |  | 1100 | 1550
2012 | 9 | 1 | 245 | 245 | jm | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1100 | 1550
2012 | 9 | 2 | 246 | 246 | jm | 60 |  | 50 | 1950 | 1800 |  | 2 | Y |  | 1100 | 1550
2012 | 9 | 3 | 247 | 247 | jm | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1100 | 1550
2012 | 9 | 4 | 248 | 248 | lr | 60 |  | 49 | 1950 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 9 | 5 | 249 | 249 | jm | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 9 | 6 | 250 | 250 | lr | 60 |  | 52 | 1950 | 1775 |  | 2 | y |  | 1100 | 1525
2012 | 9 | 7 | 251 | 251 | jm | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 9 | 8 | 252 | 252 | jm | 60 |  | 53 | 1950 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 9 | 9 | 253 | 253 | lr | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 9 | 10 | 254 | 254 | jm | 60 |  | 51 | 1950 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 9 | 11 | 255 | 255 | jm | 60 |  | 52 | 1950 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 9 | 12 | 256 | 256 | jm | 60 |  | 51 | 1950 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 9 | 13 | 257 | 257 | jm | 60 |  | 49 | 1950 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 9 | 14 | 258 | 258 | jm | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 9 | 15 | 259 | 259 | jm | 60 |  | 50 | 1975 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 9 | 16 | 260 | 260 | jm | 60 |  | 52 | 1975 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 9 | 17 | 261 | 261 | lr | 60 |  | 52 | 1975 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 9 | 18 | 262 | 262 | jm | 60 |  | 50 | 1975 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 9 | 19 | 263 | 263 | jm | 60 |  | 52 | 1975 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 9 | 20 | 264 | 264 | jm | 60 |  | 52 | 1975 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 9 | 21 | 265 | 265 | jm | 60 |  | 52 | 1975 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 9 | 22 | 266 | 266 | jm | 60 |  | 50 | 1975 | 1800 |  | 2 | y |  | 1100 | 1525 | [11 empty] | Topped off reservoir with water
2012 | 9 | 23 | 267 | 267 | lr | 60 |  | 52 | 1975 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 9 | 24 | 268 | 268 | jm | 60 |  | 51 | 1975 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 9 | 25 | 269 | 269 | jm | 60 |  | 51 | 1975 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 9 | 26 | 270 | 270 | jm | 60 |  | 52 | 1975 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 9 | 27 | 271 | 271 | lr | 60 |  | 52 | 1975 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 9 | 28 | 272 | 272 | jm | 60 |  | 50 | 1975 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 9 | 29 | 273 | 273 | lr | 60 |  | 51 | 1975 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 9 | 30 | 274 | 274 | jm | 60 |  | 49 | 1975 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 10 | 1 | 275 | 275 | jm | 60 |  | 50 | 1975 | 1800 |  | 2 | y |  | 1100 | 1525 | [11 empty] | Added DI water
2012 | 10 | 2 | 276 | 276 | lr | 60 |  | 50 | 1975 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 10 | 3 | 277 | 277 | jm | 60 |  | 50 | 1975 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 10 | 4 | 278 | 278 | lr | 60 |  | 51 | 1975 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 10 | 5 | 279 | 279 | jm | 60 |  | 51 | 1975 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 10 | 6 | 280 | 280 | jm | 60 |  | 51 | 1975 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 10 | 7 | 281 | 281 | lr | 60 |  | 52 | 1975 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 10 | 8 | 282 | 282 | jm | 60 |  | 50 | 1975 | 1800 |  | 2 | y |  | 1100 | 1525
2012 | 10 | 9 | 283 | 283 | jm | 60 |  | 52 | 1975 | 1775 |  | 2 | y |  | 1100 | 1525
2012 | 10 | 10 | 284 | 284
2012 | 10 | 11 | 285 | 285 | jm | 60 |  | 50 | 1975 | 1775 |  | 2 | y |  | 1100 | 1525
2012 | 10 | 12 | 286 | 286 | jm | 60 |  | 50 | 1975 | 1775 |  | 2 | y |  | 1100 | 1525
2012 | 10 | 13 | 287 | 287 | jm | 60 |  | 52 | 1975 | 1775 |  | 2 | y |  | 1075 | 1525
2012 | 10 | 14 | 288 | 288 | jm | 60 |  | 51 | 1975 | 1775 |  | 2 | y |  | 1075 | 1525
2012 | 10 | 15 | 289 | 289 | lr | 60 |  | 52 | 1975 | 1775 |  | 2 | y |  | 1075 | 1525
2012 | 10 | 16 | 290 | 290 | jm | 60 |  | 51 | 1975 | 1775 |  | 2 | y |  | 1075 | 1525
2012 | 10 | 17 | 291 | 291 | lr | 60 |  | 51 | 1975 | 1775 |  | 2 | y |  | 1075 | 1525
2012 | 10 | 18 | 292 | 292 | jm | 60 |  | 52 | 1975 | 1775 |  | 2 | y |  | 1075 | 1525
2012 | 10 | 19 | 293 | 293 | jm | 60 |  | 53 | 1975 | 1775 |  | 2 | y |  | 1075 | 1525
2012 | 10 | 20 | 294 | 294 | jm | 60 |  | 52 | 1975 | 1775 |  | 2 | y |  | 1075 | 1525
2012 | 10 | 21 | 295 | 295 | lr | 60 |  | 53 | 1950 | 1750 |  | 2 | y |  | 1075 | 1525
2012 | 10 | 22 | 296 | 296 | jm | 60 |  | 51 | 1950 | 1750 |  | 2 | y |  | 1075 | 1525 | [11 empty] | Added DI water
2012 | 10 | 23 | 297 | 297 | lr | 60 |  | 51 | 1950 | 1750 |  | 2 | y |  | 1075 | 1525
2012 | 10 | 24 | 298 | 298 | jm | 60 |  | 51 | 1950 | 1750 |  | 2 | y |  | 1075 | 1525
2012 | 10 | 25 | 299 | 299 | jm | 60 |  | 51 | 1950 | 1750 |  | 2 | y |  | 1075 | 1525
2012 | 10 | 26 | 300 | 300 | jm | 60 |  | 50 | 1950 | 1750 |  | 2 | y |  | 1075 | 1525
2012 | 10 | 27 | 301 | 301 | jm | 60 |  | 53 | 1950 | 1750 |  | 2 | y |  | 1050 | 1500
2012 | 10 | 28 | 302 | 302
2012 | 10 | 29 | 303 | 303 | jm | 60 |  | 51 | 1950 | 1750 |  | 2 | y |  | 1050 | 1500
2012 | 10 | 30 | 304 | 304 | jm | 60 |  | 51 | 1950 | 1750 |  | 2 | y |  | 1050 | 1500
2012 | 10 | 31 | 305 | 305 | jm | 60 |  | 50 | 1950 | 1750 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 1 | 306 | 306 | lr | 60 |  | 51 | 1950 | 1750 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 2 | 307 | 307 | jm | 60 |  | 48 | 1950 | 1750 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 3 | 308 | 308 | jm | 60 |  | 50 | 1950 | 1750 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 4 | 309 | 309 | lr | 60 |  | 51 | 1950 | 1750 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 5 | 310 | 310 | lr | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500 | [11 empty] | Added DI water
2012 | 11 | 6 | 311 | 311 | jm | 60 |  | 50 | 1950 | 1750 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 7 | 312 | 312 | bn | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 8 | 313 | 313 | ns | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 9 | 314 | 314 | ns | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 10 | 315 | 315 | ns | 60 |  | 51 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 11 | 316 | 316 | ns | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 12 | 317 | 317 | ns | 60 |  | 51 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 13 | 318 | 318 | bn | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 14 | 319 | 319 | ns | 60 |  | 51 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 15 | 320 | 320 | bn | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 16 | 321 | 321 | ns | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 17 | 322 | 322 | ns | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 18 | 323 | 323 | ns | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 19 | 324 | 324 | ns | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 20 | 325 | 325 | ns | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 21 | 326 | 326 | bn | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 22 | 327 | 327 | ns | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 23 | 328 | 328 | bn | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 24 | 329 | 329 | ns | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 25 | 330 | 330 | bn | 60 |  | 51 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500 | [11 empty] | Added DI water
2012 | 11 | 26 | 331 | 331 | ns | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 27 | 332 | 332 | ns | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 28 | 333 | 333 | ns | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 29 | 334 | 334 | bn | 60 |  | 51 | 1950 | 1775 |  | 2 | y |  | 1050 | 1500
2012 | 11 | 30 | 335 | 335 | ns | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 12 | 1 | 336 | 336 | bn | 60 |  | 51 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 12 | 2 | 337 | 337 | ns | 60 |  | 51 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 12 | 3 | 338 | 338 | bn | 60 |  | 51 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 12 | 4 | 339 | 339 | ns | 60 |  | 52 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 12 | 5 | 340 | 340 | bn | 60 |  | 51 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 12 | 6 | 341 | 341 | ns | 60 |  | 52 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 12 | 7 | 342 | 342 | bn | 60 |  | 51 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 12 | 8 | 343 | 343 | ns | 60 |  | 51 | 1950 | 1800 |  | 2 | y |  | 1050 | 1500
2012 | 12 | 9 | 344 | 344 | bn | 60 |  | 51 | 1950 | 1750 |  | 2 | y |  | 1000 | 1500
2012 | 12 | 10 | 345 | 345 | ns | 60 |  | 51 | 1950 | 1750 |  | 2 | y |  | 1000 | 1500
2012 | 12 | 11 | 346 | 346
2012 | 12 | 12 | 347 | 347 | ns | 60 |  | 51 | 1950 | 1750 |  | 2 | y |  | 1000 | 1500
2012 | 12 | 13 | 348 | 348 | bn | 60 |  | 51 | 1950 | 1750 |  | 2 | y |  | 1000 | 1500
2012 | 12 | 14 | 349 | 349 | ns | 60 |  | 51 | 1950 | 1750 |  | 2 | y |  | 1000 | 1500 | [11 empty] | Added DI water
2012 | 12 | 15 | 350 | 350 | bn | 60 |  | 52 | 1950 | 1750 |  | 2 | y |  | 1000 | 1500
2012 | 12 | 16 | 351 | 351 | ns | 60 |  | 50 | 1950 | 1750 |  | 2 | y |  | 1000 | 1500
2012 | 12 | 17 | 352 | 352 | bn | 60 |  | 50 | 1950 | 1750 |  | 2 | y |  | 1000 | 1500
2012 | 12 | 18 | 353 | 353 | ns | 60 |  | 51 | 1950 | 1800 |  | 2 | y |  | 1000 | 1500
2012 | 12 | 19 | 354 | 354 | bn | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1000 | 1500
2012 | 12 | 20 | 355 | 355 | ns | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1000 | 1500
2012 | 12 | 21 | 356 | 356 | bn | 60 |  | 51 | 1950 | 1800 |  | 2 | y |  | 1000 | 1500
2012 | 12 | 22 | 357 | 357 | ns | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1000 | 1500
2012 | 12 | 23 | 358 | 358 | bn | 60 |  | 51 | 1950 | 1800 |  | 2 | y |  | 1000 | 1450
2012 | 12 | 24 | 359 | 359 | ns | 60 |  | 52 | 1950 | 1800 |  | 2 | y |  | 1000 | 1500
2012 | 12 | 25 | 360 | 360
2012 | 12 | 26 | 361 | 361 | ns | 60 |  | 51 | 1950 | 1800 |  | 2 | y |  | 1000 | 1500
2012 | 12 | 27 | 362 | 362 | bn | 60 |  | 51 | 1950 | 1800 |  | 2 | y |  | 1000 | 1500
2012 | 12 | 28 | 363 | 363 | ns | 60 |  | 52 | 1950 | 1800 |  | 2 | y |  | 1000 | 1500
2012 | 12 | 29 | 364 | 364 | bn | 60 |  | 52 | 1950 | 1800 |  | 2 | y |  | 1000 | 1500
2012 | 12 | 30 | 365 | 365 | ns | 60 |  | 50 | 1950 | 1800 |  | 2 | y |  | 1000 | 1450
2012 | 12 | 31 | 366 | 366 | bn | 60 |  | 52 | 1950 | 1750 |  | 2 | y |  | 1000 | 1450
2013 | 1 | 1 | 1 | 367 | ns | 60 |  | 52 | 1950 | 1750 |  | 2 | y |  | 1000 | 1450
2013 | 1 | 2 | 2 | 368 | bn | 60 |  | 52 | 1950 | 1750 |  | 2 | y |  | 1000 | 1450 | [11 empty] | Added DI water
2013 | 1 | 3 | 3 | 369 | ns | 60 |  | 53 | 1950 | 1750 |  | 2 | y |  | 1000 | 1450
2013 | 1 | 4 | 4 | 370 | bn | 60 |  | 52 | 1950 | 1750 |  | 2 | y |  | 1000 | 1450
2013 | 1 | 5 | 5 | 371
2013 | 1 | 6 | 6 | 372 | bn | 60 |  | 52 | 1950 | 1750 |  | 2 | y |  | 1000 | 1450
2013 | 1 | 7 | 7 | 373 | ns | 60 |  | 53 | 1950 | 1750 |  | 2 | y |  | 1000 | 1450
2013 | 1 | 8 | 8 | 374 | bn | 60 |  | 52 | 1950 | 1750 |  | 2 | y |  | 1000 | 1450
2013 | 1 | 9 | 9 | 375 | ns | 60 |  | 50 | 1950 | 1750 |  | 2 | y |  | 1000 | 1450
2013 | 1 | 10 | 10 | 376 | bn | 60 |  | 52 | 1950 | 1750 |  | 2 | y |  | 1000 | 1450
2013 | 1 | 11 | 11 | 377 | ns | 60 |  | 51 | 1950 | 1750 |  | 2 | y |  | 1000 | 1450
2013 | 1 | 12 | 12 | 378 | bn | 60 |  | 51 | 1950 | 1750 |  | 2 | y |  | 1000 | 1450
2013 | 1 | 13 | 13 | 379 | ns | 60 |  | 51 | 1950 | 1750 |  | 2 | y |  | 1000 | 1450
2013 | 1 | 14 | 14 | 380 | bn | 60 |  | 51 | 1950 | 1750 |  | 2 | y |  | 1000 | 1450
2013 | 1 | 15 | 15 | 381 | bn | 60 |  | 52 | 1950 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 1 | 16 | 16 | 382 | ns | 60 |  | 51 | 1950 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 1 | 17 | 17 | 383 | bn | 60 |  | 52 | 1950 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 1 | 18 | 18 | 384
2013 | 1 | 19 | 19 | 385 | ns | 60 |  | 53 | 1950 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 1 | 20 | 20 | 386 | bn | 60 |  | 52 | 1950 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 1 | 21 | 21 | 387 | ns | 60 |  | 52 | 1950 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 1 | 22 | 22 | 388 | bn | 60 |  | 52 | 1950 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 1 | 23 | 23 | 389 | ns | 60 |  | 50 | 1950 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 1 | 24 | 24 | 390 | bn | 60 |  | 53 | 1950 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 1 | 25 | 25 | 391 | ns | 60 |  | 52 | 1950 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 1 | 26 | 26 | 392 | bn | 60 |  | 52 | 1950 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 1 | 27 | 27 | 393 | ns | 60 |  | 52 | 1950 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 1 | 28 | 28 | 394 | bn | 60 |  | 52 | 1900 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 1 | 29 | 29 | 395 | ns | 60 |  | 53 | 1900 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 1 | 30 | 30 | 396 | bn | 60 |  | 52 | 1900 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 1 | 31 | 31 | 397 | ns | 60 |  | 53 | 1900 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 2 | 1 | 32 | 398 | bn | 60 |  | 50 | 1900 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 2 | 2 | 33 | 399 | ns | 60 |  | 50 | 1900 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 2 | 3 | 34 | 400 | HM | 60 |  | 50 | 1900 | 1750 |  | 2 | y |  | 950 | 1450 | [11 empty] | Added DI water
2013 | 2 | 4 | 35 | 401 | HM | 60 |  | 52 | 1900 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 2 | 5 | 36 | 402 | WH | 60 |  | 50 | 1900 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 2 | 6 | 37 | 403 | HM | 60 |  | 51 | 1900 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 2 | 7 | 38 | 404 | HM | 60 |  | 52 | 1900 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 2 | 8 | 39 | 405 | HM | 60 |  | 51 | 1900 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 2 | 9 | 40 | 406 | WAH | 60 |  | 50 | 1900 | 1725 |  | 2 | y |  | 950 | 1450
2013 | 2 | 10 | 41 | 407 | HM | 60 |  | 52 | 1900 | 1725 |  | 2 | y |  | 950 | 1450
2013 | 2 | 11 | 42 | 408 | WAH | 60 |  | 50 | 1900 | 1725 |  | 2 | y |  | 950 | 1450
2013 | 2 | 12 | 43 | 409 | WAH | 60 |  | 50 | 1900 | 1725 |  | 2 | y |  | 950 | 1450
2013 | 2 | 13 | 44 | 410 | WAH | 60 |  | 49 | 1900 | 1725 |  | 2 | y |  | 950 | 1450
2013 | 2 | 14 | 45 | 411 | HM | 60 |  | 53 | 1900 | 1725 |  | 2 | y |  | 950 | 1450
2013 | 2 | 15 | 46 | 412 | HM | 60 |  | 53 | 1925 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 2 | 16 | 47 | 413 | HM | 60 |  | 53 | 1950 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 2 | 17 | 48 | 414 | HM | 60 |  | 52 | 1950 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 2 | 18 | 49 | 415 | WAH | 60 |  | 50 | 1920 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 2 | 19 | 50 | 416 | HM | 60 |  | 51 | 1925 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 2 | 20 | 51 | 417 | WAH | 60 |  | 51 | 1920 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 2 | 21 | 52 | 418 | HM | 60 |  | 50 | 1920 | 1750 |  | 2 | y |  | 950 | 1450
2013 | 2 | 22 | 53 | 419 | WAH | 60 |  | 50 | 1925 | 1740 |  | 2 | y |  | 950 | 1450
2013 | 2 | 23 | 54 | 420 | WAH | 60 |  | 52 | 1925 | 1740 |  | 2 | y |  | 940 | 1450
2013 | 2 | 24 | 55 | 421 | WAH | 60 |  | 51 | 1920 | 1740 |  | 2 | y |  | 925 | 1440
2013 | 2 | 25 | 56 | 422 | HM | 60 |  | 50 | 1920 | 1740 |  | 2 | y |  | 925 | 1450
2013 | 2 | 26 | 57 | 423 | WAH | 60 |  | 53 | 1920 | 1740 |  | 2 | y |  | 925 | 1450
2013 | 2 | 27 | 58 | 424 | HM | 60 |  | 51 | 1920 | 1740 |  | 2 | y |  | 925 | 1450
2013 | 2 | 28 | 59 | 425 | WAH | 60 |  | 53 | 1920 | 1740 |  | 2 | y |  | 920 | 1450
2013 | 3 | 1 | 60 | 426 | HM | 60 |  | 51 | 1920 | 1740 |  | 2 | y |  | 920 | 1450
2013 | 3 | 2 | 61 | 427 | HM | 60 |  | 52 | 1920 | 1740 |  | 2 | y |  | 920 | 1450
2013 | 3 | 3 | 62 | 428 | HM | 60 |  | 50 | 1920 | 1740 |  | 2 | y |  | 920 | 1440
2013 | 3 | 4 | 63 | 429 | WAH | 60 |  | 52 | 1940 | 1740 |  | 2 | y |  | 920 | 1440
2013 | 3 | 5 | 64 | 430 | HM | 60 |  | 51 | 1940 | 1740 |  | 2 | y |  | 920 | 1440
2013 | 3 | 6 | 65 | 431 | WAH | 60 |  | 51 | 1940 | 1740 |  | 2 | y |  | 920 | 1440 | [11 empty] | Water refill, H2 Gen had a pressure error around 1700UTC and has been in standby mode since
2013 | 3 | 7 | 66 | 432 | WAH | 0 |  | 0 | 1940 | 1740 |  | 2 | y |  | 920 | 1440 | [11 empty] | Restarted H2 gen @ 1330UTC. holding 60PSI. Flow = 105 and dropping slowly. @1810UTC FR = 85. @22:40 FR = 69
2013 | 3 | 8 | 67 | 433 | HM | 60 |  | 58 | 1920 | 1720 |  | 2 | y |  | 920 | 1440
2013 | 3 | 9 | 68 | 434 | HM | 60 |  | 50 | 1920 | 1720 |  | 2 | y |  | 910 | 1440
2013 | 3 | 10 | 69 | 435 | WAH | 60 |  | 50 | 1920 | 1720 |  | 2 | y |  | 910 | 1440
2013 | 3 | 11 | 70 | 436 | HM | 60 |  | 51 | 1920 | 1720 |  | 2 | y |  | 910 | 1440
2013 | 3 | 12 | 71 | 437 | HM | 60 |  | 53 | 1920 | 1720 |  | 2 | y |  | 910 | 1440
2013 | 3 | 13 | 72 | 438 | HM | 60 |  | 52 | 1920 | 1720 |  | 2 | y |  | 910 | 1440
2013 | 3 | 14 | 73 | 439 | WAH | 60 |  | 53 | 1920 | 1720 |  | 2 | y |  | 910 | 1440
2013 | 3 | 15 | 74 | 440 | HM | 60 |  | 54 | 1910 | 1720 |  | 2 | y |  | 910 | 1440
2013 | 3 | 16 | 75 | 441 | WAH | 60 |  | 53 | 1920 | 1730 |  | 2 | y |  | 910 | 1440
2013 | 3 | 17 | 76 | 442 | HM | 60 |  | 52 | 1920 | 1730 |  | 2 | y |  | 910 | 1440
2013 | 3 | 18 | 77 | 443 | HM | 60 |  | 52 | 1910 | 1730 |  | 2 | y |  | 910 | 1440
2013 | 3 | 19 | 78 | 444 | WAH | 60 |  | 52 | 1900 | 1710 |  | 2 | y |  | 900 | 1430
2013 | 3 | 20 | 79 | 445 | HM | 60 |  | 52 | 1900 | 1710 |  | 2 | y |  | 900 | 1430
2013 | 3 | 21 | 80 | 446 | HM | 60 |  | 52 | 1900 | 1710 |  | 2 | y |  | 900 | 1430
2013 | 3 | 22 | 81 | 447 | WAH | 60 |  | 53 | 1900 | 1710 |  | 2 | y |  | 900 | 1430
2013 | 3 | 23 | 82 | 448 | WAH | 60 |  | 51 | 1900 | 1710 |  | 2 | y |  | 900 | 1430
2013 | 3 | 24 | 83 | 449 | WAH | 60 |  | 51 | 1890 | 1700 |  | 2 | y |  | 900 | 1430
2013 | 3 | 25 | 84 | 450 | HM | 60 |  | 53 | 1890 | 1700 |  | 2 | y |  | 900 | 1430
2013 | 3 | 26 | 85 | 451 | HM | 60 |  | 50 | 1890 | 1700 |  | 2 | y |  | 900 | 1430
2013 | 3 | 27 | 86 | 452 | HM | 60 |  | 51 | 1890 | 1700 |  | 2 | y |  | 900 | 1430
2013 | 3 | 28 | 87 | 453 | HM | 60 |  | 52 | 1890 | 1700 |  | 2 | y |  | 900 | 1430
2013 | 3 | 29 | 88 | 454 | HM | 60 |  | 53 | 1890 | 1700 |  | 2 | y |  | 900 | 1430
2013 | 3 | 30 | 89 | 455 | HM | 60 |  | 52 | 1890 | 1700 |  | 2 | y |  | 900 | 1430 | [11 empty] | added DI water to hydrogen generator
2013 | 3 | 31 | 90 | 456 | HM | 60 |  | 51 | 1890 | 1700 |  | 2 | y |  | 900 | 1430
2013 | 4 | 1 | 91 | 457 | HM | 60 |  | 52 | 1900 | 1700 |  | 2 | y |  | 900 | 1430
2013 | 4 | 2 | 92 | 458 | WAH | 60 |  | 52 | 1900 | 1700 |  | 2 | y |  | 890 | 1440
2013 | 4 | 3 | 93 | 459 | WAH | 60 |  | 53 | 1900 | 1690 |  | 2 | y |  | 880 | 1430
2013 | 4 | 4 | 94 | 460 | WAH | 60 |  | 51 | 1900 | 1690 |  | 2 | y |  | 890 | 1440
2013 | 4 | 5 | 95 | 461 | HM | 60 |  | 53 | 1900 | 1690 |  | 2 | y |  | 890 | 1440
2013 | 4 | 6 | 96 | 462 | HM | 60 |  | 54 | 1900 | 1690 |  | 2 | y |  | 890 | 1440
2013 | 4 | 7 | 97 | 463 | WAH | 60 |  | 52 | 1890 | 1680 |  | 2 | y |  | 880 | 1430
2013 | 4 | 8 | 98 | 464 | HM | 60 |  | 53 | 1890 | 1680 |  | 2 | y |  | 880 | 1430
2013 | 4 | 9 | 99 | 465 | HM | 60 |  | 52 | 1890 | 1680 |  | 2 | y |  | 880 | 1430
2013 | 4 | 10 | 100 | 466 | HM | 60 |  | 52 | 1890 | 1680 |  | 2 | y |  | 880 | 1420
2013 | 4 | 11 | 101 | 467 | WAH | 60 |  | 51 | 1890 | 1680 |  | 2 | y |  | 880 | 1430
2013 | 4 | 12 | 102 | 468 | WAH | 60 |  | 51 | 1890 | 1680 |  | 2 | y |  | 880 | 1430
2013 | 4 | 13 | 103 | 469 | HM | 60 |  | 52 | 1890 | 1680 |  | 2 | y |  | 880 | 1430
2013 | 4 | 14 | 104 | 470 | HM | 60 |  | 52 | 1890 | 1680 |  | 2 | y |  | 880 | 1430
2013 | 4 | 15 | 105 | 471 | HM | 60 |  | 51 | 1890 | 1680 |  | 2 | y |  | 880 | 1430
2013 | 4 | 16 | 106 | 472 | WAH | 60 |  | 51 | 1890 | 1680 |  | 2 | y |  | 870 | 1430
2013 | 4 | 17 | 107 | 473 | WAH | 60 |  | 51 | 1890 | 1680 |  | 2 | y |  | 880 | 1430
2013 | 4 | 18 | 108 | 474 | HM | 60 |  | 50 | 1890 | 1680 |  | 2 | y |  | 880 | 1430
2013 | 4 | 19 | 109 | 475 | WAH | 60 |  | 51 | 1890 | 1680 |  | 2 | y |  | 870 | 1420
2013 | 4 | 20 | 110 | 476 | HM | 60 |  | 52 | 1890 | 1680 |  | 2 | y |  | 870 | 1420
2013 | 4 | 21 | 111 | 477 | WAH | 60 |  | 51 | 1890 | 1680 |  | 2 | y |  | 870 | 1420
2013 | 4 | 22 | 112 | 478 | SL | 60 |  | 50 | 1900 | 1700 |  | 2 | y |  | 860 | 1410
2013 | 4 | 23 | 113 | 479 | HM | 60 |  | 53 | 1890 | 1680 |  | 2 | y |  | 860 | 1410
2013 | 4 | 24 | 114 | 480 | WAH | 60 |  | 54 | 1890 | 1670 |  | 2 | y |  | 880 | 1430 | [11 empty] | added 1 quart DI water
2013 | 4 | 25 | 115 | 481 | DH | 60 |  | 50 | 1890 | 1670 |  | 2 | Y |  | 850 | 1410
2013 | 4 | 26 | 116 | 482 | DH | 60 |  | 52 | 1890 | 1680 |  | 2 | Y |  | 850 | 1410
2013 | 4 | 27 | 117 | 483 | HM | 60 |  | 52 | 1890 | 1680 |  | 2 | y |  | 1050 | 1440 | [11 empty] | switched NMHC Std 1 to NIST_2 std tank
2013 | 4 | 28 | 118 | 484 | HM | 60 |  | 51 | 1890 | 1680 |  | 2 | y |  | 1050 | 1440
2013 | 4 | 29 | 119 | 485 | SL | 60 |  | 50 | 1890 | 1680 |  | 2 | y |  | 1150 | 1430
2013 | 4 | 30 | 120 | 486 | SL | 60 |  | 54 | 1890 | 1680 |  | 2 | y |  | 1150 | 1430 | [11 empty] | Noticed that pressure was set to 58 - Reset to 60
2013 | 5 | 1 | 121 | 487 | SL | 60 |  | 50 | 1890 | 1680 |  | 2 | y |  | 1150 | 1430
2013 | 5 | 2 | 122 | 488 | SL | 60 |  | 54 | 1890 | 1680 |  | 2 | y |  | 1150 | 1430
2013 | 5 | 3 | 123 | 489 | WAH | 60 |  | 51 | 1880 | 1680 |  | 2 | y |  | 1140 | 1430
2013 | 5 | 4 | 124 | 490 | WAH | 60 |  | 52 | 1880 | 1670 |  | 2 | y |  | 1140 | 1420
2013 | 5 | 5 | 125 | 491 | SL | 60 |  | 50 | 1880 | 1670 |  | 2 | y |  | 1140 | 1410
2013 | 5 | 6 | 126 | 492 | WAH | 60 |  | 51 | 1880 | 1670 |  | 2 | y |  | 1140 | 1410
2013 | 5 | 7 | 127 | 493 | WAH | 60 |  | 50 | 1880 | 1680 |  | 2 | y |  | 1140 | 1420 | [11 empty] | Power outage from 0840-0905UTC.
2013 | 5 | 8 | 128 | 494 | SL | 60 |  | 50 | 1880 | 1680 |  | 2 | Y |  | 1140 | 1420 | [11 empty] | Power outage from 1040-1048UTC.
2013 | 5 | 9 | 129 | 495 | SL | 60 |  | 53 | 1880 | 1680 |  | 2 | Y |  | 1140 | 1420
2013 | 5 | 10 | 130 | 496 | SL | 60 |  | 51 | 1880 | 1680 |  | 2 | Y |  | 1140 | 1420
2013 | 5 | 11 | 131 | 497 | WAH | 60 |  | 51 | 1880 | 1680 |  | 2 | Y |  | 1140 | 1420
2013 | 5 | 12 | 132 | 498 | WAH | 60 |  | 51 | 1880 | 1690 |  | 2 | Y |  | 1140 | 1430
2013 | 5 | 13 | 133 | 499 | WAH | 60 |  | 51 | 1880 | 1690 |  | 2 | Y |  | 1140 | 1420
2013 | 5 | 14 | 134 | 500 | WAH | 60 |  | 51 | 1880 | 1690 |  | 2 | Y |  | 1140 | 1420
2013 | 5 | 15 | 135 | 501
2013 | 5 | 16 | 136 | 502 | WAH | 60 |  | 51 | 1880 | 1690 |  | 2 | Y |  | 1140 | 1420 | [11 empty] | Power outage from 1156-1159UTC and 2042-2048UTC
2013 | 5 | 17 | 137 | 503 | WAH | 60 |  | 51 | 1880 | 1690 |  | 2 | Y |  | 1140 | 1420
2013 | 5 | 18 | 138 | 504 | WAH | 60 |  | 51 | 1880 | 1690 |  | 2 | Y |  | 1140 | 1420
2013 | 5 | 19 | 139 | 505 | SL | 60 |  | 51 | 1880 | 1690 |  | 2 | Y |  | 1140 | 1420 | [11 empty] | Filled water in H2 Gen
2013 | 5 | 20 | 140 | 506 | WAH | 60 |  | 50 | 1880 | 1690 |  | 2 | Y |  | 1140 | 1420
2013 | 5 | 21 | 141 | 507 | WAH | 60 |  | 50 | 1880 | 1690 |  | 2 | Y |  | 1140 | 1410
2013 | 5 | 22 | 142 | 508 | WAH | 60 |  | 53 | 1880 | 1690 |  | 2 | Y |  | 1140 | 1410
2013 | 5 | 23 | 143 | 509 | SL | 60 |  | 51 | 1810 | 1640 |  | 2 | Y |  | 1140 | 1410
2013 | 5 | 24 | 144 | 510 | SL | 60 |  | 53 | 1870 | 1670 |  | 2 | Y |  | 1140 | 1410
2013 | 5 | 25 | 145 | 511 | SL | 60 |  | 50 | 1870 | 1670 |  | 2 | Y |  | 1140 | 1410
2013 | 5 | 26 | 146 | 512 | WAH | 60 |  | 51 | 1890 | 1710 |  | 2 | Y |  | 1150 | 1430
2013 | 5 | 27 | 147 | 513 | WAH | 60 |  | 51 | 1890 | 1710 |  | 2 | Y |  | 1160 | 1430 | [11 empty] | Filled water in H2 Gen
2013 | 5 | 28 | 148 | 514 | WAH | 60 |  | 50 | 1890 | 1710 |  | 2 | Y |  | 1160 | 1420
2013 | 5 | 29 | 149 | 515 | SL | 60 |  | 52 | 1890 | 1700 |  | 2 | Y |  | 1160 | 1420
2013 | 5 | 30 | 150 | 516 | WAH | 60 |  | 50 | 1890 | 1710 |  | 2 | Y |  | 1160 | 1440
2013 | 5 | 31 | 151 | 517 | SL | 60 |  | 51 | 1890 | 1700 |  | 2 | Y |  | 1160 | 1430
2013 | 6 | 1 | 152 | 518 | SL | 60 |  | 51 | 1890 | 1690 |  | 2 | Y |  | 1160 | 1430
2013 | 6 | 2 | 153 | 519 | DB | 60 |  | 50 | 1890 | 1690 |  | 2 | Y |  | 1160 | 1430
2013 | 6 | 3 | 154 | 520 | DB | 60 |  | 51 | 1890 | 1690 |  | 2 | Y |  | 1160 | 1430 | [11 empty] | Filled water in H2 Gen
2013 | 6 | 4 | 155 | 521 | ss | 60 |  | 50 | 1890 | 1690 |  | 2 | y |  | 1160 | 1430
2013 | 6 | 5 | 156 | 522 | DB | 60 |  | 51 | 1890 | 1690 |  | 2 | y |  | 1150 | 1430
2013 | 6 | 6 | 157 | 523 | SS | 60 |  | 50 | 1890 | 1690 |  | 2 | Y |  | 1160 | 1430
2013 | 6 | 7 | 158 | 524 | DB | 60 |  | 51 | 1890 | 1690 |  | 2 | Y |  | 1160 | 1430
2013 | 6 | 8 | 159 | 525 | ss | 60 |  | 51 | 1890 | 1690 |  | 2 | y |  | 1160 | 1430
2013 | 6 | 9 | 160 | 526 | DB | 60 |  | 51 | 1890 | 1690 |  | 2 | y |  | 1160 | 1430
2013 | 6 | 10 | 161 | 527 | ss | 60 |  | 50 | 1890 | 1680 |  | 2 | y |  | 1160 | 1430 | [11 empty] | Added water to H2 Gen
2013 | 6 | 11 | 162 | 528 | ss | 60 |  | 51 | 1890 | 1680 |  | 2 | y |  | 1160 | 1420
2013 | 6 | 12 | 163 | 529 | DB | 60 |  | 51 | 1890 | 1680 |  | 2 | Y |  | 1150 | 1420
2013 | 6 | 13 | 164 | 530 | DB | 60 |  | 51 | 1800 | 1680 |  | 2 | Y |  | 1150 | 1420
2013 | 6 | 14 | 165 | 531 | ss | 60 |  | 53 | 1890 | 1680 |  | 2 | y |  | 1160 | 1420
2013 | 6 | 15 | 166 | 532 | DB | 60 |  | 52 | 1890 | 1680 |  | 2 | Y |  | 1160 | 1420
2013 | 6 | 16 | 167 | 533 | ss | 60 |  | 53 | 1890 | 1680 |  | 2 | y |  | 1160 | 1420
2013 | 6 | 17 | 168 | 534 | DB | 60 |  | 50 | 1890 | 1680 |  | 2 | Y |  | 1160 | 1420
2013 | 6 | 18 | 169 | 535 | ss | 60 |  | 51 | 1890 | 1630 |  | 2 | y |  | 1160 | 1420
2013 | 6 | 19 | 170 | 536 | DB | 60 |  | 52 | 1890 | 1680 |  | 2 | y |  | 1160 | 1420
2013 | 6 | 20 | 171 | 537 | ss | 60 |  | 51 | 1890 | 1680 |  | 2 | y |  | 1160 | 1420
2013 | 6 | 21 | 172 | 538 | DB | 60 |  | 51 | 1800 | 1650 |  | 2 | y |  | 1160 | 1400
2013 | 6 | 22 | 173 | 539 | ss | 60 |  | 51 | 1890 | 1660 |  | 2 | y |  | 1130 | 1400
2013 | 6 | 23 | 174 | 540 | ss | 60 |  | 52 | 1890 | 1650 |  | 2 | y |  | 1130 | 1390
2013 | 6 | 24 | 175 | 541 | ss | 60 |  | 51 | 1890 | 1650 |  | 2 | y |  | 1130 | 1390
2013 | 6 | 25 | 176 | 542 | DB | 60 |  | 51 | 1890 | 1650 |  | 2 | Y |  | 1130 | 1390
2013 | 6 | 26 | 177 | 543 | ss | 60 |  | 51 | 1890 | 1650 |  | 2 | y |  | 1130 | 1390
2013 | 6 | 27 | 178 | 544 | DB | 60 |  | 52 | 1890 | 1650 |  | 2 | y |  | 1130 | 1390
2013 | 6 | 28 | 179 | 545 | ss | 60 |  | 53 | 1890 | 1650 |  | 2 | y |  | 1130 | 1400
2013 | 6 | 29 | 180 | 546 | DB | 60 |  | 52 | 1890 | 1650 |  | 2 | y |  | 1130 | 1390
2013 | 6 | 30 | 181 | 547 | ss | 60 |  | 51 | 1890 | 1650 |  | 2 | y |  | 1130 | 1390
2013 | 7 | 1 | 182 | 548 | DB | 60 |  | 51 | 1890 | 1650 |  | 2 | y |  | 1130 | 1390 | [11 empty] | added H2O
2013 | 7 | 2 | 183 | 549 | ss | 60 |  | 53 | 1890 | 1650 |  | 2 | y |  | 1130 | 1390
2013 | 7 | 3 | 184 | 550 | DB | 60 |  | 51 | 1890 | 1650 |  | 2 | y |  | 1110 | 1390
2013 | 7 | 4 | 185 | 551 | ss | 60 |  | 50 | 1890 | 1650 |  | 2 | y |  | 1150 | 1400
2013 | 7 | 5 | 186 | 552 | DB | 60 |  | 51 | 1890 | 1650 |  | 2 | y |  | 1150 | 1400
2013 | 7 | 6 | 187 | 553 | ss | 60 |  | 52 | 1890 | 1650 |  | 2 | y |  | 1150 | 1400
2013 | 7 | 7 | 188 | 554 | DB | 60 |  | 51 | 1890 | 1650 |  | 2 | y |  | 1150 | 1400
2013 | 7 | 8 | 189 | 555 | ss | 60 |  | 51 | 1890 | 1650 |  | 2 | y |  | 1150 | 1400
2013 | 7 | 9 | 190 | 556 | DB | 60 |  | 52 | 1890 | 1650 |  | 2 | y |  | 1150 | 1400
2013 | 7 | 10 | 191 | 557 | ss | 60 |  | 52 | 1850 | 1700 |  | 2 | y |  | 1150 | 1400 | [11 empty] | low H2O alarm went off, added H20
2013 | 7 | 11 | 192 | 558 | DB | 60 |  | 52 | 1850 | 1650 |  | 2 | y |  | 1150 | 1400
2013 | 7 | 12 | 193 | 559 | ss | 60 |  | 53 | 1890 | 1660 |  | 2 | y |  | 1150 | 1400
2013 | 7 | 13 | 194 | 560 | ss | 60 |  | 52 | 1890 | 1660 |  | 2 | y |  | 1150 | 1400
2013 | 7 | 14 | 195 | 561 | ss | 60 |  | 55 | 1890 | 1660 |  | 2 | y |  | 1150 | 1420
2013 | 7 | 15 | 196 | 562 | DB | 60 |  | 50 | 1890 | 1650 |  | 2 | y |  | 1150 | 1420 | [11 empty] | H2O level fine.
2013 | 7 | 16 | 197 | 563 | ss | 60 |  | 53 | 1890 | 1650 |  | 2 | y |  | 1150 | 1420
2013 | 7 | 17 | 198 | 564 | DB | 60 |  | 53 | 1890 | 1650 |  | 2 | y |  | 1150 | 1420
2013 | 7 | 18 | 199 | 565 | DB | 60 |  | 53 | 1890 | 1650 |  | 2 | y |  | 1150 | 1420
2013 | 7 | 19 | 200 | 566 | DB | 60 |  | 53 | 1860 | 1650 |  | 2 | y |  | 1130 | 1410
2013 | 7 | 20 | 201 | 567 | ss | 60 |  | 54 | 1860 | 1650 |  | 2 | y |  | 1130 | 1410
2013 | 7 | 21 | 202 | 568 | DB | 60 |  | 52 | 1860 | 1650 |  | 2 | y |  | 1130 | 1400
2013 | 7 | 22 | 203 | 569 | ss | 60 |  | 55 | 1860 | 1650 |  | 2 | y |  | 1130 | 1400 | [11 empty] | added 1L H2O
2013 | 7 | 23 | 204 | 570 | DB | 60 |  | 52 | 1860 | 1650 |  | 2 | y |  | 1130 | 1400
2013 | 7 | 24 | 205 | 571 | ss | 60 |  | 52 | 1860 | 1650 |  | 2 | y |  | 1130 | 1400
2013 | 7 | 25 | 206 | 572 | DB | 60 |  | 51 | 1860 | 1650 |  | 2 | y |  | 1130 | 1400
2013 | 7 | 26 | 207 | 573 | ss | 60 |  | 52 | 1860 | 1650 |  | 2 | y |  | 1130 | 1400
2013 | 7 | 27 | 208 | 574 | DB | 60 |  | 52 | 1860 | 1650 |  | 2 | y |  | 1130 | 1400
2013 | 7 | 28 | 209 | 575 | ss | 60 |  | 51 | 1860 | 1650 |  | 2 | y |  | 1130 | 1400
2013 | 7 | 29 | 210 | 576 | DB | 60 |  | 53 | 1890 | 1650 |  | 2 | y |  | 1130 | 1400
2013 | 7 | 30 | 211 | 577 | ss | 60 |  | 53 | 1890 | 1650 |  | 0 | n |  | 1130 | 1400 | [11 empty] | turned off for TAWO raise
2013 | 7 | 31 | 212 | 578 | DB | 60 |  | 53 | 1890 | 1650 |  | 0 | n |  | 1130 | 1400 | [11 empty] | turned off for TAWO raise
2013 | 8 | 1 | 213 | 579 | ss | 60 |  | 53 | 1890 | 1650 |  | 0 | n |  | 1130 | 1400 | [11 empty] | turned off for TAWO raise
2013 | 8 | 2 | 214 | 580 | DB | 60 |  | 52 | 1890 | 1650 |  | 0 | n |  | 1130 | 1400 | [11 empty] | turned off for TAWO raise
2013 | 8 | 3 | 215 | 581 | ss | 60 |  | 53 | 1890 | 1650 |  | 2 | y |  | 1130 | 1400
2013 | 8 | 4 | 216 | 582 | DB | 60 |  | 51 | 1900 | 1650 |  | 2 | y |  | 1130 | 1400
2013 | 8 | 5 | 217 | 583 | ss | 60 |  | 54 | 1890 | 1680 |  | 2 | y |  | 1130 | 1420 | [11 empty] | added 1L H20
2013 | 8 | 6 | 218 | 584 | DB | 60 |  | 52 | 1890 | 1680 |  | 2 | y |  | 1130 | 1420
2013 | 8 | 7 | 219 | 585 | ss | 60 |  | 54 | 1890 | 1680 |  | 2 | y |  | 1130 | 1420
2013 | 8 | 8 | 220 | 586 | DB | 60 |  | 52 | 1890 | 1680 |  | 2 | y |  | 1130 | 1410
2013 | 8 | 9 | 221 | 587 | ss | 60 |  | 53 | 1890 | 1680 |  | 2 | y |  | 1130 | 1410
2013 | 8 | 10 | 222 | 588 | DB | 60 |  | 52 | 1880 | 1680 |  | 2 | Y |  | 1130 | 1410
2013 | 8 | 11 | 223 | 589 | ss | 60 |  | 52 | 1880 | 1680 |  | 2 | Y |  | 1130 | 1410
2013 | 8 | 12 | 224 | 590 | DB | 60 |  | 52 | 1880 | 1680 |  | 2 | y |  | 1130 | 1410
2013 | 8 | 13 | 225 | 591 | ss | 60 |  | 53 | 1880 | 1680 |  | 2 | y |  | 1130 | 1410
2013 | 8 | 14 | 226 | 592 | DB | 60 |  | 52 | 1880 | 1680 |  | 2 | y |  | 1130 | 1410
2013 | 8 | 15 | 227 | 593 | SS | 60 |  | 53 | 1860 | 1650 |  | 2 | y |  | 1130 | 1410
2013 | 8 | 16 | 228 | 594 | DB | 60 |  | 52 | 1860 | 1650 |  | 2 | y |  | 1100 | 1410
2013 | 8 | 17 | 229 | 595 | ss | 60 |  | 52 | 1860 | 1650 |  | 2 | y |  | 1110 | 920 | [11 empty] | changed standard number 2 to BA std
2013 | 8 | 18 | 230 | 596 | DB | 60 |  | 53 | 1860 | 1650 |  | 2 | y |  | 1110 | 920
2013 | 8 | 19 | 231 | 597 | ss | 60 |  | 52 | 1860 | 1650 |  | 2 | y |  | 1110 | 920 | [11 empty] | added 1L H20
2013 | 8 | 20 | 232 | 598 | DB | 60 |  | 53 | 1860 | 1650 |  | 2 | y |  | 1110 | 910
2013 | 8 | 21 | 233 | 599 | DB | 60 |  | 55 | 1860 | 1650 |  | 2 | y |  | 1110 | 910
2013 | 8 | 22 | 234 | 600 | DB | 60 |  | 53 | 1860 | 1650 |  | 2 | y |  | 1110 | 910
2013 | 8 | 23 | 235 | 601 | DB | 60 |  | 55 | 1860 | 1650 |  | 2 | y |  | 1110 | 910
2013 | 8 | 24 | 236 | 602 | DB | 60 |  | 53 | 1860 | 1650 |  | 2 | y |  | 1110 | 910
2013 | 8 | 25 | 237 | 603 | DB | 60 |  | 53 | 1860 | 1650 |  | 2 | y |  | 1110 | 910
2013 | 8 | 26 | 238 | 604 | ss | 60 |  | 54 | 1860 | 1650 |  | 2 | y |  | 1100 | 900 | [11 empty] | water level is good
2013 | 8 | 27 | 239 | 605 | ss | 60 |  | 53 | 1860 | 1650 |  | 2 | y |  | 1100 | 900
2013 | 8 | 28 | 240 | 606 | ss | 60 |  | 54 | 1860 | 1650 |  | 2 | y |  | 1100 | 900
2013 | 8 | 29 | 241 | 607 | ss | 60 |  | 55 | 1860 | 1650 |  | 2 | y |  | 1100 | 900
2013 | 8 | 30 | 242 | 608 | DB | 60 |  | 54 | 1860 | 1650 |  | 2 | y |  | 1100 | 900
2013 | 8 | 31 | 243 | 609 | DB | 60 |  | 55 | 1860 | 1650 |  | 2 | y |  | 1100 | 900
2013 | 9 | 1 | 244 | 610 | ss | 60 |  | 55 | 1860 | 1650 |  | 2 | y |  | 1100 | 900
2013 | 9 | 2 | 245 | 611 | DB | 60 |  | 55 | 1860 | 1650 |  | 2 | y |  | 1100 | 900
2013 | 9 | 3 | 246 | 612 | DB | 60 |  | 53 | 1810 | 1600 |  | 2 | y |  | 1050 | 860
2013 | 9 | 4 | 247 | 613 | DB | 60 |  | 54 | 1850 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 9 | 5 | 248 | 614 | DB | 60 |  | 54 | 1850 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 9 | 6 | 249 | 615 | DB | 60 |  | 53 | 1850 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 9 | 7 | 250 | 616 | DB | 60 |  | 56 | 1850 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 9 | 8 | 251 | 617 | DB | 60 |  | 54 | 1850 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 9 | 9 | 252 | 618 | ss | 60 |  | 54 | 1850 | 1650 |  | 2 | y |  | 1050 | 900 | [11 empty] | added 1L of water
2013 | 9 | 10 | 253 | 619 | ss | 60 |  | 54 | 1850 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 9 | 11 | 254 | 620 | ss | 60 |  | 54 | 1850 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 9 | 12 | 255 | 621 | ss | 60 |  | 54 | 1850 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 9 | 13 | 256 | 622 | ss | 60 |  | 55 | 1850 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 9 | 14 | 257 | 623 | ss | 60 |  | 55 | 1850 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 9 | 15 | 258 | 624 | ss | 60 |  | 55 | 1840 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 9 | 16 | 259 | 625 | DB | 60 |  | 54 | 1820 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 9 | 17 | 260 | 626 | DB | 60 |  | 55 | 1820 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 9 | 18 | 261 | 627 | DB | 60 |  | 54 | 1820 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 9 | 19 | 262 | 628 | DB | 60 |  | 54 | 1820 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 9 | 20 | 263 | 629 | ss | 60 |  | 54 | 1840 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 9 | 21 | 264 | 630 | DB | 60 |  | 56 | 1820 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 9 | 22 | 265 | 631 | DB | 60 |  | 54 | 1820 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 9 | 23 | 266 | 632 | ss | 60 |  | 56 | 1840 | 1650 |  | 2 | y |  | 1050 | 900 | [11 empty] | added 1L H2O
2013 | 9 | 24 | 267 | 633 | ss | 60 |  | 55 | 1840 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 9 | 25 | 268 | 634 | ss | 60 |  | 55 | 1840 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 9 | 26 | 269 | 635 | DB | 60 |  | 54 | 1840 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 9 | 27 | 270 | 636 | DB | 60 |  | 54 | 1840 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 9 | 28 | 271 | 637 | DB | 60 |  | 53 | 1840 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 9 | 29 | 272 | 638 | DB | 60 |  | 54 | 1840 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 9 | 30 | 273 | 639 | DB | 60 |  | 55 | 1840 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 10 | 1 | 274 | 640 | DB | 60 |  | 54 | 1840 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 10 | 2 | 275 | 641 | DB | 60 |  | 53 | 1840 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 10 | 3 | 276 | 642 | ss | 60 |  | 54 | 1840 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 10 | 4 | 277 | 643 | DB | 60 |  | 54 | 1840 | 1650 |  | 2 | y |  | 1050 | 890
2013 | 10 | 5 | 278 | 644 | DB | 60 |  | 55 | 1840 | 1650 |  | 2 | y |  | 1050 | 890
2013 | 10 | 6 | 279 | 645 | DB | 60 |  | 55 | 1840 | 1650 |  | 2 | y |  | 1050 | 890
2013 | 10 | 7 | 280 | 646 | ss | 60 |  | 55 | 1840 | 1650 |  | 2 | y |  | 1080 | 880
2013 | 10 | 8 | 281 | 647 | ss | 60 |  | 55 | 1840 | 1650 |  | 2 | y |  | 1080 | 880
2013 | 10 | 9 | 282 | 648 | DB | 60 |  | 56 | 1840 | 1650 |  | 2 | y |  | 1080 | 880
2013 | 10 | 10 | 283 | 649 | ss | 60 |  | 55 | 1840 | 1650 |  | 2 | y |  | 1080 | 880
2013 | 10 | 11 | 284 | 650 | ss | 60 |  | 56 | 1840 | 1650 |  | 2 | y |  | 1080 | 880 | [11 empty] | added 1 L of H2O
2013 | 10 | 12 | 285 | 651 | ss | 60 |  | 55 | 1840 | 1650 |  | 2 | y |  | 1080 | 880
2013 | 10 | 13 | 286 | 652 | ss | 60 |  | 56 | 1840 | 1650 |  | 2 | y |  | 1080 | 880
2013 | 10 | 14 | 287 | 653 | ss | 60 |  | 55 | 1820 | 1650 |  | 2 | y |  | 1080 | 880
2013 | 10 | 15 | 288 | 654 | DB | 60 |  | 56 | 1820 | 1650 |  | 2 | y |  | 1080 | 880
2013 | 10 | 16 | 289 | 655 | DB | 60 |  | 55 | 1820 | 1650 |  | 2 | y |  | 1080 | 880
2013 | 10 | 17 | 290 | 656 | DB | 60 |  | 55 | 1820 | 1650 |  | 2 | y |  | 1080 | 880
2013 | 10 | 18 | 291 | 657 | DB | 60 |  | 54 | 1820 | 1650 |  | 2 | y |  | 1080 | 880
2013 | 10 | 19 | 292 | 658 | DB | 60 |  | 54 | 1820 | 1650 |  | 2 | y |  | 1080 | 880
2013 | 10 | 20 | 293 | 659 | DB | 60 |  | 54 | 1820 | 1650 |  | 2 | y |  | 1080 | 880
2013 | 10 | 21 | 294 | 660 | DB | 60 |  | 56 | 1820 | 1650 |  | 2 | y |  | 1080 | 880
2013 | 10 | 22 | 295 | 661 | ss | 60 |  | 56 | 1820 | 1650 |  | 2 | y |  | 1080 | 880
2013 | 10 | 23 | 296 | 662 | ss | 60 |  | 57 | 1820 | 1650 |  | 2 | Y |  | 1070 | 860
2013 | 10 | 24 | 297 | 663 | ss | 60 |  | 55 | 1820 | 1650 |  | 2 | y |  | 1070 | 860
2013 | 10 | 25 | 298 | 664 | DB | 60 |  | 56 | 1820 | 1650 |  | 2 | y |  | 1070 | 860
2013 | 10 | 26 | 299 | 665 | ss | 60 |  | 54 | 1820 | 1650 |  | 2 | y |  | 1070 | 860
2013 | 10 | 27 | 300 | 666 | bms | 60 |  | 54 | 1820 | 1650 |  | 2 | y |  | 1070 | 870 | [11 empty] | added 1/2 L
2013 | 10 | 28 | 301 | 667 | ss | 60 |  | 54 | 1820 | 1650 |  | 2 | y |  | 1070 | 870
2013 | 10 | 29 | 302 | 668 | bms | 60 |  | 53 | 1820 | 1650 |  | 2 | y |  | 1070 | 870
2013 | 10 | 30 | 303 | 669 | bms | 60 |  | 54 | 1820 | 1650 |  | 2 | y |  | 1070 | 870
2013 | 10 | 31 | 304 | 670 | bms | 60 |  | 55 | 1820 | 1600 |  | 2 | y |  | 1070 | 870 | [11 empty] | water ok
2013 | 11 | 1 | 305 | 671 | bms | 60 |  | 54 | 1820 | 1600 |  | 2 | y |  | 1070 | 870
2013 | 11 | 2 | 306 | 672 | jm | 60 |  | 56 | 1820 | 1600 |  | 2 | y |  | 1070 | 870 | [11 empty] | added 1/2 L
2013 | 11 | 3 | 307 | 673 | bms | 60 |  | 55 | 1810 | 1610 |  | 2 | y |  | 1070 | 860
2013 | 11 | 4 | 308 | 674 | bms | 60 |  | 55 | 1810 | 1610 |  | 2 | y |  | 1070 | 860
2013 | 11 | 5 | 309 | 675 | bms | 60 |  | 53 | 1810 | 1610 |  | 2 | y |  | 1070 | 860
2013 | 11 | 6 | 310 | 676 | jm | 60 |  | 53 | 1810 | 1610 |  | 2 | y |  | 1070 | 860
2013 | 11 | 7 | 311 | 677 | jm | 60 |  | 53 | 1810 | 1610 |  | 2 | y |  | 1070 | 860
2013 | 11 | 8 | 312 | 678 | jm | 60 |  | 53 | 1810 | 1610 |  | 2 | y |  | 1070 | 860 | [11 empty] | added 1/2 L
2013 | 11 | 9 | 313 | 679 | bms | 60 |  | 56 | 1810 | 1610 |  | 2 | y |  | 1070 | 860
2013 | 11 | 10 | 314 | 680 | jm | 60 |  | 53 | 1810 | 1610 |  | 2 | y |  | 1070 | 860
2013 | 11 | 11 | 315 | 681 | jm | 60 |  | 54 | 1810 | 1610 |  | 2 | y |  | 1070 | 860
2013 | 11 | 12 | 316 | 682 | bms | 60 |  | 55 | 1810 | 1610 |  | 2 | y |  | 1070 | 860
2013 | 11 | 13 | 317 | 683 | jm | 60 |  | 55 | 1810 | 1610 |  | 2 | y |  | 1070 | 860
2013 | 11 | 14 | 318 | 684 | jm | 60 |  | 53 | 1810 | 1610 |  | 2 | y |  | 1070 | 860 | [11 empty] | added 1/2 L
2013 | 11 | 15 | 319 | 685 | jm | 60 |  | 55 | 1810 | 1610 |  | 2 | y |  | 1070 | 860
2013 | 11 | 16 | 320 | 686 | jm | 60 |  | 54 | 1810 | 1610 |  | 2 | y |  | 1070 | 860
2013 | 11 | 17 | 321 | 687 | bms | 60 |  | 54 | 1810 | 1610 |  | 2 | y |  | 1070 | 860 | [11 empty] | water ok
2013 | 11 | 18 | 322 | 688 | jm | 60 |  | 55 | 1810 | 1610 |  | 2 | y |  | 1070 | 860
2013 | 11 | 19 | 323 | 689 | jm | 60 |  | 53 | 1810 | 1610 |  | 2 | y |  | 1070 | 900
2013 | 11 | 20 | 324 | 690 | jm | 60 |  | 54 | 1810 | 1610 |  | 2 | y |  | 1070 | 900
2013 | 11 | 21 | 325 | 691 | jm | 60 |  | 54 | 1810 | 1610 |  | 2 | y |  | 1070 | 900 | [11 empty] | Added 3/4 L water
2013 | 11 | 22 | 326 | 692 | jm | 60 |  | 55 | 1810 | 1610 |  | 2 | y |  | 1070 | 860
2013 | 11 | 23 | 327 | 693 | jm | 60 |  | 57 | 1810 | 1650 |  | 2 | y |  | 1050 | 860
2013 | 11 | 24 | 328 | 694 | jm | 60 |  | 56 | 1810 | 1650 |  | 2 | y |  | 1050 | 860
2013 | 11 | 25 | 329 | 695 | jm | 60 |  | 56 | 1810 | 1650 |  | 2 | y |  | 1050 | 900
2013 | 11 | 26 | 330 | 696 | jm | 60 |  | 55 | 1810 | 1650 |  | 2 | y |  | 1050 | 860
2013 | 11 | 27 | 331 | 697 | jm | 60 |  | 55 | 1810 | 1650 |  | 2 | y |  | 1050 | 860 | [11 empty] | Added 1/2 L H2O
2013 | 11 | 28 | 332 | 698 | [23 empty] | Did not visit TAWO due to high winds
2013 | 11 | 29 | 333 | 699 | jm | 60 |  | 55 | 1810 | 1650 |  | 2 | y |  | 1050 | 860
2013 | 11 | 30 | 334 | 700 | jm | 60 |  | 56 | 1810 | 1630 |  | 2 | y |  | 1050 | 860
2013 | 12 | 1 | 335 | 701 | bms | 60 |  | 56 | 1810 | 1630 |  | 2 | y |  | 1050 | 850
2013 | 12 | 2 | 336 | 702 | jm | 60 |  | 55 | 1810 | 1630 |  | 2 | y |  | 1050 | 850
2013 | 12 | 3 | 337 | 703 | jm | 60 |  | 59 | 1725 | 1550 |  | 2 | y |  | 975 | 840
2013 | 12 | 4 | 338 | 704 | jm | 60 |  | 55 | 1725 | 1550 |  | 2 | y |  | 975 | 840
2013 | 12 | 5 | 339 | 705 | bms | 60 |  | 56 | 1725 | 1550 |  | 2 | y |  | 1000 | 825 | [11 empty] | Jennie is working on FID flame problem with Jacques.
2013 | 12 | 6 | 340 | 706 | jm | 60 |  | 56 | 1800 | 1600 |  | 2 | y |  | 1025 | 860
2013 | 12 | 7 | 341 | 707 | jm | 60 |  | 54 | 1800 | 1600 |  | 2 | y |  | 1025 | 860
2013 | 12 | 8 | 342 | 708 | jm | 60 |  | 53 | 1800 | 1640 |  | 2 | y |  | 1050 | 860
2013 | 12 | 9 | 343 | 709 | jm | 60 |  | 54 | 1810 | 1640 |  | 2 | y |  | 1050 | 860
2013 | 12 | 10 | 344 | 710 | jm | 60 |  | 54 | 1810 | 1650 |  | 2 | y |  | 1050 | 860
2013 | 12 | 11 | 345 | 711 | jm | 60 |  | 53 | 1810 | 1650 |  | 2 | y |  | 1050 | 860
2013 | 12 | 12 | 346 | 712 | bms | 60 |  | 55 | 1810 | 1625 |  | 2 | y |  | 1050 | 860
2013 | 12 | 13 | 347 | 713 | [23 empty] | Did not visit
2013 | 12 | 14 | 348 | 714 | jm | 60 |  | 53 | 1810 | 1625 |  | 2 | y |  | 1050 | 860 | [11 empty] | Added 1/2 L H2O
2013 | 12 | 15 | 349 | 715 | bms | 60 |  | 53 | 1810 | 1625 |  | 2 | y |  | 1025 | 860
2013 | 12 | 16 | 350 | 716 | jm | 60 |  | 53 | 1800 | 1610 |  | 2 | y |  | 1025 | 860
2013 | 12 | 17 | 351 | 717 | jm | 60 |  | 54 | 1800 | 1625 |  | 2 | y |  | 1025 | 860
2013 | 12 | 18 | 352 | 718 | bms | 60 |  | 60 | 1810 | 1620 |  | 2 | y |  | 1025 | 850
2013 | 12 | 19 | 353 | 719 | bms | 60 |  | 59 | 1810 | 1620 |  | 2 | y |  | 1050 | 850
2013 | 12 | 20 | 354 | 720 | [23 empty] | no visit, high winds
2013 | 12 | 21 | 355 | 721 | jm | 60 |  | 59 | 1810 | 1650 |  | 2 | y |  | 1050 | 860 | [11 empty] | H2O fine
2013 | 12 | 22 | 356 | 722 | jm | 60 |  | 59 | 1810 | 1650 |  | 2 | y |  | 1050 | 860
2013 | 12 | 23 | 357 | 723 | jm | 60 |  | 59 | 1810 | 1650 |  | 2 | y |  | 1050 | 860
2013 | 12 | 24 | 358 | 724 | jm | 60 |  | 58 | 1810 | 1650 |  | 2 | y |  | 1050 | 860
2013 | 12 | 25 | 359 | 725 | bms | 60 |  | 60 | 1800 | 1610 |  | 2 | y |  | 1020 | 850
2013 | 12 | 26 | 360 | 726 | jm | 60 |  | 60 | 1800 | 1640 |  | 2 | y |  | 1000 | 850
2013 | 12 | 27 | 361 | 727 | bms | 60 |  | 62 | 1810 | 1620 |  | 2 | y |  | 1050 | 850
2013 | 12 | 28 | 362 | 728 | jm | 60 |  | 61 | 1810 | 1625 |  | 2 | y |  | 1050 | 850 | [11 empty] | Added 1/2L H2O
2013 | 12 | 29 | 363 | 729 | bms | 60 |  | 60 | 1810 | 1625 |  | 2 | y |  | 1050 | 850
2013 | 12 | 30 | 364 | 730 | jm | 60 |  | 61 | 1820 | 1625 |  | 2 | y |  | 1050 | 860
2013 | 12 | 31 | 365 | 731 | jm | 60 |  | 62 | 1800 | 1610 |  | 2 | y |  | 1000 | 850
2014 | 1 | 1 | 1 | 732 | bms | 60 |  | 60 | 1800 | 1610 |  | 2 | y |  | 1020 | 850
2014 | 1 | 2 | 2 | 733 | bms | 60 |  | 61 | 1800 | 1610 |  | 2 | y |  | 1020 | 850
2014 | 1 | 3 | 3 | 734 | jm | 60 |  | 61 | 1800 | 1610 |  | 2 | y |  | 1000 | 850
2014 | 1 | 4 | 4 | 735 | jm | 60 |  | 61 | 1800 | 1610 |  | 2 | y |  | 1000 | 850
2014 | 1 | 5 | 5 | 736 | jm | 60 |  | 60 | 1800 | 1610 |  | 2 | y |  | 1000 | 850
2014 | 1 | 6 | 6 | 737 | jm | 60 |  | 60 | 1800 | 1625 |  | 2 | y |  | 1000 | 860
2014 | 1 | 7 | 7 | 738 | jm | 60 |  | 60 | 1800 | 1625 |  | 2 | y |  | 1000 | 860
2014 | 1 | 8 | 8 | 739 | jm | 60 |  | 62 | 1800 | 1625 |  | 2 | y |  | 1000 | 860
2014 | 1 | 9 | 9 | 740 | jm | 60 |  | 62 | 1800 | 1625 |  | 2 | y |  | 1000 | 850
2014 | 1 | 10 | 10 | 741 | jm | 60 |  | 61 | 1790 | 1600 |  | 2 | y |  | 1000 | 850
2014 | 1 | 11 | 11 | 742 | jm | 60 |  | 61 | 1800 | 1600 |  | 2 | y |  | 1000 | 860
2014 | 1 | 12 | 12 | 743 | bms | 60 |  | 63 | 1800 | 1600 |  | 2 | y |  | 1050 | 850
2014 | 1 | 13 | 13 | 744 | jm | 60 |  | 61 | 1800 | 1625 |  | 2 | y |  | 1000 | 850
2014 | 1 | 14 | 14 | 745 | bms | 60 |  | 60 | 1800 | 1620 |  | 2 | y |  | 1050 | 840
2014 | 1 | 15 | 15 | 746 | bms | 60 |  | 61 | 1800 | 1620 |  | 2 | y |  | 1050 | 840
2014 | 1 | 16 | 16 | 747 | bms | 60 |  | 62 | 1800 | 1620 |  | 2 | y |  | 1050 | 840
2014 | 1 | 17 | 17 | 748 | bms | 60 |  | 61 | 1800 | 1620 |  | 2 | y |  | 1050 | 840
2014 | 1 | 18 | 18 | 749 | jm | 60 |  | 63 | 1800 | 1620 |  | 2 | y |  | 1050 | 840 | [11 empty] | Added 1/2L H2O
2014 | 1 | 19 | 19 | 750 | jm | 60 |  | 62 | 1800 | 1620 |  | 2 | y |  | 1050 | 840
2014 | 1 | 20 | 20 | 751 | jm | 60 |  | 61 | 1800 | 1620 |  | 2 | y |  | 1050 | 850
2014 | 1 | 21 | 21 | 752 | bms | 60 |  | 61 | 1790 | 1610 |  | 2 | y |  | 1025 | 830
2014 | 1 | 22 | 22 | 753 | jm | 60 |  | 62 | 1800 | 1610 |  | 2 | y |  | 1025 | 840 | [11 empty] | Added 1/2L H2O
2014 | 1 | 23 | 23 | 754 | jm | 60 |  | 62 | 1800 | 1610 |  | 2 | y |  | 1025 | 850
2014 | 1 | 24 | 24 | 755 | jm | 60 |  | 62 | 1800 | 1610 |  | 2 | y |  | 1000 | 850
2014 | 1 | 25 | 25 | 756 | bms | 60 |  | 61 | 1790 | 1610 |  | 2 | y |  | 1020 | 850
2014 | 1 | 26 | 26 | 757 | bms | 60 |  | 63 | 1790 | 1620 |  | 2 | y |  | 1020 | 850
2014 | 1 | 27 | 27 | 758 | jm | 60 |  | 62 | 1790 | 1625 |  | 2 | y |  | 1000 | 840
2014 | 1 | 28 | 28 | 759 | jm | 60 |  | 61 | 1800 | 1625 |  | 2 | y |  | 1000 | 850
2014 | 1 | 29 | 29 | 760 | bms | 60 |  | 63 | 1790 | 1620 |  | 2 | y |  | 1000 | 850
2014 | 1 | 30 | 30 | 761 | bms | 60 |  | 64 | 1790 | 1620 |  | 2 | y |  | 1020 | 840
2014 | 1 | 31 | 31 | 762 | [23 empty] | did not visit due to high winds
2014 | 2 | 1 | 32 | 763 | jm | 60 |  | 62 | 1800 | 1625 |  | 2 | y |  | 1000 | 850 | [11 empty] | added 1/2L H2O
2014 | 2 | 2 | 33 | 764 | jm | 60 |  | 61 | 1800 | 1625 |  | 2 | y |  | 1000 | 850
2014 | 2 | 3 | 34 | 765 | jm | 60 |  | 62 | 1800 | 1625 |  | 2 | y |  | 1000 | 850
2014 | 2 | 4 | 35 | 766 | jm | 60 |  | 61 | 1800 | 1625 |  | 2 | y |  | 1000 | 850
2014 | 2 | 5 | 36 | 767 | jm | 60 |  | 63 | 1800 | 1625 |  | 2 | y |  | 1000 | 850
2014 | 2 | 6 | 37 | 768 | bms | 60 |  | 62 | 1790 | 1610 |  | 2 | y |  | 1025 | 840
2014 | 2 | 7 | 38 | 769 | bms | 60 |  | 65 | 1790 | 1610 |  | 2 | y |  | 1025 | 840
2014 | 2 | 8 | 39 | 770 | jm | 60 |  | 62 | 1790 | 1610 |  | 2 | y |  | 1025 | 840 | [11 empty] | Added 1/2L H2O.
2014 | 2 | 9 | 40 | 771 | bms | 60 |  | 62 | 1780 | 1610 |  | 2 | y |  | 1020 | 830
2014 | 2 | 10 | 41 | 772 | jm | 60 |  | 63 | 1800 | 1625 |  | 2 | y |  | 1025 | 840
2014 | 2 | 11 | 42 | 773 | jm | 60 |  | 63 | 1790 | 1625 |  | 2 | y |  | 1000 | 840
2014 | 2 | 12 | 43 | 774 | jm | 60 |  | 62 | 1790 | 1625 |  | 2 | y |  | 1000 | 840
2014 | 2 | 13 | 44 | 775 | bms | 60 |  | 63 | 1775 | 1590 |  | 2 | y |  | 1000 | 820
2014 | 2 | 14 | 45 | 776 | bms | 60 |  | 65 | 1780 | 1600 |  | 2 | y |  | 1000 | 830
2014 | 2 | 15 | 46 | 777 | bms | 60 |  | 62 | 1780 | 1610 |  | 2 | y |  | 1000 | 850
2014 | 2 | 16 | 47 | 778 | jm | 60 |  | 63 | 1780 | 1610 |  | 2 | y |  | 1000 | 850
2014 | 2 | 17 | 48 | 779 | jm | 60 |  | 63 | 1790 | 1610 |  | 2 | y |  | 1000 | 850
2014 | 2 | 18 | 49 | 780 | [23 empty] | no rounds, bad weather
2014 | 2 | 19 | 50 | 781 | bms | 60 |  | 63 | 1780 | 1610 |  | 2 | y |  | 1000 | 820
2014 | 2 | 20 | 51 | 782 | bms | 60 |  | 62 | 1780 | 1610 |  | 2 | y |  | 1000 | 830 | [11 empty] | added 1/3L H2O; beeping alarm and message "Service Reqd" prior to adding H2O
2014 | 2 | 21 | 52 | 783 | jm | 60 |  | 62 | 1790 | 1610 |  | 2 | y |  | 1000 | 840
2014 | 2 | 22 | 53 | 784 | aw | 60 |  | 59 | 1800 | 1600 |  | 2 | y |  | 1000 | 850
2014 | 2 | 23 | 54 | 785 | aw | 60 |  | 61 | 1800 | 1625 |  | 2 | y |  | 1000 | 800
2014 | 2 | 24 | 55 | 786 | aw | 60 |  | 61 | 1800 | 1600 |  | 2 | y |  | 1000 | 800
2014 | 2 | 25 | 56 | 787 | jm | 60 |  | 61 | 1790 | 1600 |  | 2 | y |  | 1000 | 840
2014 | 2 | 26 | 57 | 788 | lr | 60 | x | 62 | 1775 | 1600 |  | 2 | y |  | 1000 | 825
2014 | 2 | 27 | 58 | 789 | lr | 60 | x | 62 | 1775 | 1600 |  | 2 | y |  | 1000 | 825
2014 | 2 | 28 | 59 | 790 | lr | 60 | x | 61 | 1775 | 1600 |  | 2 | y |  | 1000 | 825
2014 | 3 | 1 | 60 | 791 | aw | 60 | x | 62 | 1775 | 1600 |  | 2 | y |  | 1000 | 800
2014 | 3 | 2 | 61 | 792 | lr | 60 | x | 62 | 1775 | 1600 |  | 2 | y |  | 1000 | 850
2014 | 3 | 3 | 62 | 793 | lr | 60 | x | 62 | 1775 | 1600 |  | 2 | y |  | 1000 | 850
2014 | 3 | 4 | 63 | 794 | aw | 60 | x | 61 | 1800 | 1600 |  | 2 | y |  | 1000 | 800
2014 | 3 | 5 | 64 | 795 | aw | 60 | x | 61 | 1775 | 1600 |  | 2 | y |  | 1000 | 800
2014 | 3 | 6 | 65 | 796 | aw | 60 | x | 62 | 1775 | 1600 |  | 2 | y |  | 1000 | 775
2014 | 3 | 7 | 66 | 797 | aw | 60 | x | 62 | 1775 | 1600 |  | 2 | y |  | 1000 | 850
2014 | 3 | 8 | 67 | 798 | aw | 60 | x | 63 | 1775 | 1625 |  | 2 | y |  | 1000 | 825 | [11 empty] | refilled H2 Generator with 3/4L DI water.
2014 | 3 | 9 | 68 | 799 | aw | 60 | x | 61 | 1775 | 1610 |  | 2 | y |  | 1000 | 775
2014 | 3 | 10 | 69 | 800 | lr | 60 | x | 61 | 1775 | 1610 |  | 2 | y |  | 1000 | 825
2014 | 3 | 11 | 70 | 801 | aw | 60 | x | 61 | 1775 | 1625 |  | 2 | y |  | 1000 | 825
2014 | 3 | 12 | 71 | 802 | lr | 60 | x | 62 | 1775 | 1625 |  | 2 | y |  | 1000 | 825
2014 | 3 | 13 | 72 | 803 | aw | 60 | x | 63 | 1800 | 1625 |  | 2 | y |  | 975 | 800
2014 | 3 | 14 | 73 | 804 | aw | 60 | x | 62 | 1775 | 1600 |  | 2 | y |  | 975 | 825
2014 | 3 | 15 | 74 | 805 | aw | 60 | x | 62 | 1800 | 1650 |  | 2 | y |  | 975 | 825
2014 | 3 | 16 | 75 | 806 | aw | 60 | x | 63 | 1750 | 1600 |  | 2 | y |  | 975 | 825
2014 | 3 | 17 | 76 | 807 | aw | 60 | x | 82 | 1700 | 1550 |  | 2 | y |  | 900 | 800
2014 | 3 | 18 | 77 | 808 | aw | [22 empty] | unable to make trip to tawo 2* weather conditions
2014 | 3 | 19 | 78 | 809 | aw | 60 | x | 62 | 1775 | 1600 |  | 2 | y |  | 975 | 850
2014 | 3 | 20 | 79 | 810 | aw | 60 | x | 63 | 1775 | 1600 |  | 2 | y |  | 975 | 825 | [11 empty] | refilled water in H2 gen
2014 | 3 | 21 | 80 | 811 | aw | 60 | x | 62 | 1775 | 1600 |  | 2 | y |  | 975 | 825 | [11 empty] | brief power outage this morning at 1224 UTC
2014 | 3 | 22 | 81 | 812 | lr | 60 | x | 62 | 1775 | 1600 |  | 2 | y |  | 975 | 825
2014 | 3 | 23 | 82 | 813 | lr | 60 | x | 62 | 1775 | 1600 |  | 2 | y |  | 975 | 825
2014 | 3 | 24 | 83 | 814 | lr | 60 | RED | 62 | 1775 | 1600 |  | 2 | y |  | 975 | 825 | [11 empty] | H2 generator error, filled with water
2014 | 3 | 25 | 84 | 815 | lr | 60 | x | 62 | 1775 | 1600 |  | 2 | y |  | 975 | 825
2014 | 3 | 26 | 85 | 816 | lr | 60 | x | 62 | 1775 | 1600 |  | 2 | y |  | 975 | 825
2014 | 3 | 27 | 86 | 817 | aw | 60 | x | 62 | 1775 | 1600 |  | 2 | y |  | 975 | 825
2014 | 3 | 28 | 87 | 818 | lr | 60 | x | 62 | 1775 | 1600 |  | 2 | y |  | 975 | 825
2014 | 3 | 29 | 88 | 819 | aw | 60 | x | 63 | 1775 | 1600 |  | 2 | y |  | 975 | 825
2014 | 3 | 30 | 89 | 820 | aw | 60 | x | 63 | 1775 | 1600 |  | 2 | y |  | 975 | 825
2014 | 3 | 31 | 90 | 821 | aw | 60 | x | 62 | 1775 | 1600 |  | 2 | y |  | 975 | 825
2014 | 4 | 1 | 91 | 822 | aw | 60 | x | 62 | 1775 | 1600 |  | 2 | y |  | 975 | 825
2014 | 4 | 2 | 92 | 823 | aw | 60 | x | 63 | 1775 | 1600 |  | 2 | y |  | 975 | 825
2014 | 4 | 3 | 93 | 824 | aw | 60 | x | 64 | 1775 | 1600 |  | 2 | y |  | 975 | 825
2014 | 4 | 4 | 94 | 825 | aw | 60 | x | 64 | 1775 | 1600 |  | 2 | y |  | 975 | 825 | [11 empty] | added water to H2Gen.
2014 | 4 | 5 | 95 | 826 | lr | 60 | x | 66 | 1775 | 1600 |  | 2 | y |  | 975 | 825
2014 | 4 | 6 | 96 | 827 | lr | 60 | x | 64 | 1775 | 1600 |  | 2 | y |  | 975 | 825
2014 | 4 | 7 | 97 | 828 | lr | 60 | x | 62 | 1775 | 1600 |  | 2 | y |  | 975 | 825
2014 | 4 | 8 | 98 | 829 | aw | 60 | x | 63 | 1775 | 1600 |  | 2 | y |  | 975 | 825
2014 | 4 | 9 | 99 | 830 | aw | 60 | x | 63 | 1775 | 1600 |  | 2 | y |  | 975 | 800
2014 | 4 | 10 | 100 | 831 | lr | 60 | x | 63 | 1775 | 1600 |  | 2 | y |  | 975 | 800
2014 | 4 | 11 | 101 | 832 | aw | 60 | x | 64 | 1775 | 1600 |  | 2 | y |  | 975 | 800
2014 | 4 | 12 | 102 | 833 | lr | 60 | x | 64 | 1775 | 1600 |  | 2 | y |  | 975 | 800
2014 | 4 | 13 | 103 | 834 | aw | 60 | x | 64 | 1775 | 1600 |  | 2 | y |  | 975 | 800
2014 | 4 | 14 | 104 | 835 | aw | 60 | x | 64 | 1750 | 1600 |  | 2 | y |  | 975 | 825
2014 | 4 | 15 | 105 | 836 | aw | 60 | x | 64 | 1775 | 1600 |  | 2 | y |  | 975 | 800 | [11 empty] | added hater to H2Gen, lifted red lid to photograph FIDs per PI request
2014 | 4 | 16 | 106 | 837 | aw | 60 | x | 64 | 1775 | 1600 |  | 2 | y |  | 950 | 800 | [11 empty] | per PI request, switched BNC connectors, photographed inside of oven
2014 | 4 | 17 | 107 | 838 | aw | 60 | x | 64 | 1750 | 1600 |  | 2 | y |  | 950 | 825
2014 | 4 | 18 | 108 | 839 | aw | 60 | x | 64 | 1775 | 1600 |  | 2 | y |  | 950 | 825
2014 | 4 | 19 | 109 | 840 | lr | 60 | x | 60 | 1775 | 1600 |  | 2 | y |  | 950 | 825 | [11 empty] | power outage from 1036-1050 UTC.
2014 | 4 | 20 | 110 | 841 | aw | 60 | x | 63 | 1775 | 1600 |  | 2 | y |  | 975 | 825 | [11 empty] | Per PI instruction: BNC connectors returned to usual configuration, H2 generator and scrubber desiccants checked and photographed, H2 tank turned on.
2014 | 4 | 21 | 111 | 842 | aw | 60 | x | 60 | 1775 | 1600 |  | 2 | y |  | 975 | 825
2014 | 4 | 22 | 112 | 843 | aw | error | x | error | 1775 | 1600 |  | 2 | y |  | 975 | 825 | [11 empty] | H2 generator beeping and showing error, PI notified. Instrument running on H2 cylinder, PSI 2150, delivery pressure 40psi
2014 | 4 | 23 | 113 | 844 | lr | error | x | error | 1775 | 1600 |  | 2 | y |  | 975 | 825 | [11 empty] | H2 generator beeping and showing error, PI notified. Instrument running on H2 cylinder, PSI 2150, delivery pressure 40psi
2014 | 4 | 24 | 114 | 845 | aw | standby | x | standby | 1700 | 1600 |  | 2 | y |  | 975 | 825 | [11 empty] | H2 generator placed in standby mode, H2 cylinder 2100psi, delivery pressure 40psi
2014 | 4 | 25 | 115 | 846 | lr | standby | x | standby | 1700 | 1600 |  | 2 | y |  | 975 | 825 | [11 empty] | H2 cylinder 2100psi, delivery pressure 40psi
2014 | 4 | 26 | 116 | 847
2014 | 4 | 27 | 117 | 848 | aw | standby | x | standby | 1700 | 1600 |  | 2 | y |  | 975 | 825 | [11 empty] | H2 cylinder 2150psi, delivery pressure 40psi
2014 | 4 | 28 | 118 | 849 | aw | standby | x | standby | 1750 | 1600 |  | 2 | y |  | 975 | 825 | [11 empty] | H2 cylinder 2025psi, delivery pressure 40psi
2014 | 4 | 29 | 119 | 850 | aw | standby | x | standby | 1725 | 1550 |  | 2 | y |  | 975 | 825 | [11 empty] | H2 cylinder 2050psi, delivery pressure 40psi
2014 | 4 | 30 | 120 | 851 | lr | standby | x | standby | 1725 | 1550 |  | 2 | y |  | 975 | 825 | [11 empty] | H2 cylinder 2050psi, delivery pressure 40psi
2014 | 5 | 1 | 121 | 852 | aw | standby | x | standby | 1725 | 1550 |  | 2 | y |  | 975 | 825 | [11 empty] | H2 cylinder 1950psi, delivery pressure 40psi
2014 | 5 | 2 | 122 | 853 | aw | standby | x | standby | 1750 | 1550 |  | 2 | y |  | 975 | 825 | [11 empty] | H2 cylinder 1850psi, delivery pressure 40psi
2014 | 5 | 3 | 123 | 854 | aw | standby | x | standby | 1750 | 1550 |  | 2 | y |  | 975 | 825 | [11 empty] | H2 cylinder 1850psi, delivery pressure 40psi
2014 | 5 | 4 | 124 | 855 | lr | standby | x | standby | 1750 | 1550 |  | 2 | y |  | 975 | 825 | [11 empty] | H2 cylinder 1850psi, delivery pressure 40psi
2014 | 5 | 5 | 125 | 856 | lr | standby | x | standby | 1750 | 1550 |  | 2 | y |  | 975 | 825 | [11 empty] | H2 cylinder 1850psi, delivery pressure 40psi
2014 | 5 | 6 | 126 | 857 | lr | standby | x | standby | 1750 | 1550 |  | 2 | y |  | 975 | 825 | [11 empty] | H2 cylinder 1800psi, delivery pressure 40psi
2014 | 5 | 7 | 127 | 858 | lr | standby | x | standby | 1750 | 1550 |  | 2 | y |  | 975 | 825 | [11 empty] | H2 cylinder 1800psi, delivery pressure 40psi
2014 | 5 | 8 | 128 | 859 | aw | standby | x | standby | 1750 | 1550 |  | 2 | y |  | 975 | 825 | [11 empty] | H2 cylinder 1750psi, delivery pressure 40psi
2014 | 5 | 9 | 129 | 860 | lr | standby | x | standby | 1750 | 1550 |  | 2 | y |  | 975 | 825 | [11 empty] | H2 cylinder 1700psi, delivery pressure 40psi
2014 | 5 | 10 | 130 | 861 | aw | standby | x | standby | 1750 | 1550 |  | 2 | y |  | 975 | 825 | [11 empty] | H2 cylinder 1650 psi, delivery pressure 40psi
2014 | 5 | 11 | 131 | 862 | lr | standby | x | standby | 1750 | 1550 |  | 2 | y |  | 975 | 825 | [11 empty] | H2 cylinder 1650 psi, delivery pressure 40psi
2014 | 5 | 12 | 132 | 863 | aw | standby | x | standby | 1750 | 1550 |  | 2 | y |  | 975 | 825 | [11 empty] | H2 cylinder 1550 psi, delivery pressure 40 psi
2014 | 5 | 13 | 133 | 864 | lr | standby | x | standby | 1750 | 1550 |  | 2 | y |  | 975 | 825 | [11 empty] | H2 cylinder 1550 psi, delivery pressure 40 psi
2014 | 5 | 14 | 134 | 865 | aw | standby | x | standby | 1750 | 1550 |  | 2 | y |  | 975 | 825 | [11 empty] | H2 cylinder 1525 psi, delivery pressure 42 psi
2014 | 5 | 15 | 135 | 866 | aw | standby | x | standby | 1700 | 1550 |  | 2 | y |  | 975 | 800 | [11 empty] | H2 cylinder 1500psi, delivery pressure 42psi
2014 | 5 | 16 | 136 | 867 | lr | 60 | x | 103 | 1700 | 1550 |  | 2 | y |  | 975 | 800
2014 | 5 | 17 | 137 | 868 | lr | 60 | x | 47 | 1700 | 1500 |  | 2 | y |  | 975 | 800
2014 | 5 | 18 | 138 | 869 | lr | standby | x | standby | 1700 | 1500 |  | 2 | y |  | 975 | 800
2014 | 5 | 19 | 139 | 870 | lr | standby | x | standby | 1700 | 1500 |  | 2 | y |  | 975 | 800 | [11 empty] | H2 cylinder 1425psi, delivery pressure 42psi
2014 | 5 | 20 | 140 | 871 | lr | standby | x | standby | 1700 | 1500 |  | 2 | y |  | 975 | 800 | [11 empty] | H2 cylinder 1425psi, delivery pressure 42psi
2014 | 5 | 21 | 141 | 872 | lr | standby | x | standby | 1700 | 1500 |  | 2 | y |  | 975 | 800 | [11 empty] | H2 cylinder 1450psi, delivery pressure 42psi
2014 | 5 | 22 | 142 | 873 | lr | standby | x | standby | 1700 | 1500 |  | 2 | y |  | 975 | 800 | [11 empty] | H2 cylinder 1450psi, delivery pressure 42psi
2014 | 5 | 23 | 143 | 874 | aw | standby | x | standby | 1700 | 1550 |  | 2 | y |  | 950 | 800 | [11 empty] | H2 cylinder pressure 1350psi, delivery pressure 40psi
2014 | 5 | 24 | 144 | 875 | lr | standby | x | standby | 1700 | 1550 |  | 2 | y |  | 950 | 800 | [11 empty] | H2 cylinder pressure 1350psi, delivery pressure 40psi
2014 | 5 | 25 | 145 | 876 | lr | standby | x | standby | 1700 | 1550 |  | 2 | y |  | 950 | 800 | [11 empty] | H2 cylinder pressure 1350psi, delivery pressure 40psi
2014 | 5 | 26 | 146 | 877 | lr | standby | x | standby | 1700 | 1550 |  | 2 | y |  | 950 | 800 | [11 empty] | H2 cylinder pressure 1300psi, delivery pressure 40psi
2014 | 5 | 27 | 147 | 878 | aw | standby | x | standby | 1700 | 1550 |  | 2 | y |  | 950 | 800 | [11 empty] | H2 cylinder pressure1275psi, delivery pressure 41psi
2014 | 5 | 28 | 148 | 879 | lr | standby | x | standby | 1700 | 1550 |  | 2 | y |  | 950 | 800 | [11 empty] | H2 cylinder pressure1275psi, delivery pressure 41psi
2014 | 5 | 29 | 149 | 880 | aw | standby | x | standby | 1700 | 1550 |  | 2 | y |  | 950 | 800 | [11 empty] | H2 cylinder pressure 1225psi, delivery pressure 41psi @ 1530 swapped power supplies and turned H2 generator to stand by. LR
2014 | 5 | 30 | 150 | 881 | lr | standby | x | standby | 1700 | 1550 |  | 2 | y |  | 950 | 800 | [11 empty] | H2 cylinder pressure 1225psi, delivery pressure 41psi
2014 | 5 | 31 | 151 | 882 | aw | standby | x | standby | 1700 | 1550 |  | 2 | y |  | 925 | 800 | [11 empty] | H2 cylinder pressure 1175psi, delivery pressure 42psi
2014 | 6 | 1 | 152 | 883 | aw | standby | x | standby | 1700 | 1550 |  | 2 | y |  | 950 | 800 | [11 empty] | H2 cylinder pressure 1150psi, delivery pressure 41psi
2014 | 6 | 2 | 153 | 884 | lr | standby | x | standby | 1700 | 1550 |  | 2 | y |  | 950 | 800 | [11 empty] | H2 cylinder pressure 1200psi, delivery pressure 41psi
2014 | 6 | 3 | 154 | 885 | lr | 60 | x | ? | 1700 | 1550 |  | 2 | y |  | 950 | 800 | [11 empty] | Switch H2 Generator on @1130UTC. LR
2014 | 6 | 4 | 155 | 886 | lr | 60 | x | 55 | 1700 | 1550 |  | 2 | y |  | 950 | 800
2014 | 6 | 5 | 156 | 887 | lr | 60 | x | 55 | 1700 | 1550 |  | 2 | y |  | 950 | 800
2014 | 6 | 6 | 157 | 888 | aw | 60 | x | 56 | 1700 | 1550 |  | 2 | y |  | 900 | 800
2014 | 6 | 7 | 158 | 889 | smc | 60 | x | 55 | 1700 | 1550 |  | 2 | y |  | 900 | 800 | [11 empty] | Turnover training, first time doing rounds
2014 | 6 | 8 | 159 | 890 | dp | 60 | x | 54 | 1700 | 1550 |  | 2 | y |  | 900 | 800 | [11 empty] | Domi first rounds
2014 | 6 | 9 | 160 | 891 | smc | 60 | x | 55 | 1700 | 1550 |  | 2 | y |  | 900 | 800
2014 | 6 | 10 | 161 | 892 | dp | 60 | x | 55 | 1700 | 1550 |  | 2 | y |  | 900 | 800 | [11 empty] | Added some H20
2014 | 6 | 11 | 162 | 893 | SMC | 60 | x | 54 | 1700 | 1550 |  | 2 | y |  | 900 | 790
2014 | 6 | 12 | 163 | 894 | dp | 60 | x | 55 | 1700 | 1550 |  | 2 | y |  | 900 | 800 | [11 empty] | Added H20
2014 | 6 | 13 | 164 | 895 | DP | 60 | X | 55 | 1700 | 1550 |  | 2 | Y |  | 900 | 800
2014 | 6 | 14 | 165 | 896 | dp | 60 | x | 55 | 1700 | 1550 |  | 2 | y |  | 900 | 800
2014 | 6 | 15 | 166 | 897 | dp | 60 | x | 54 | 1700 | 1550 |  | 2 | y |  | 900 | 800
2014 | 6 | 16 | 167 | 898 | smc | 60 | x | 57 | 1700 | 1550 | ` | 2 | y |  | 900 | 800
2014 | 6 | 17 | 168 | 899 | dp | 60 | x | 55 | 1700 | 1550 |  | 2 | y |  | 900 | 800 | [11 empty] | Hydrogen generator was off due to low water, added water, won't let it get that low again.
2014 | 6 | 18 | 169 | 900 | SMC | 60 | x | 55 | 1700 | 1550 |  | 2 | y |  | 900 | 800
2014 | 6 | 19 | 170 | 901 | dp | 60 | x | 55 | 1700 | 1550 |  | 2 | y |  | 900 | 800
2014 | 6 | 20 | 171 | 902 | dp | 60 | x | 55 | 1700 | 1550 |  | 2 | y |  | 900 | 800
2014 | 6 | 21 | 172 | 903 | SMC | 60 | x | 50 | 1700 | 1550 |  | 2 | Y |  | 900 | 800
2014 | 6 | 22 | 173 | 904 | dp | 60 | x | 53 | 1700 | 1550 |  | 2 | Y |  | 900 | 800
2014 | 6 | 23 | 174 | 905 | dp | 60 | x | 53 | 1700 | 1550 |  | 2 | Y |  | 900 | 800
2014 | 6 | 24 | 175 | 906 | SMC | 60 | x | 53 | 1700 | 1550 |  | 2 | y |  | 900 | 800
2014 | 6 | 25 | 176 | 907 | smc | 60 | X | 53 | 1700 | 1550 |  | 2 | Y |  | 900 | 800
2014 | 6 | 26 | 177 | 908 | SMC | 60 | x | 53 | 1700` | 1550 |  | 2 | y |  | 900 | 800
2014 | 6 | 27 | 178 | 909 | SMC | 60 | X | 52 | 1700 | 1550 |  | 2 | Y |  | 900 | 800
2014 | 6 | 28 | 179 | 910 | DP | 60 | X | 53 | 1700 | 1550 |  | 2 | Y |  | 900 | 800
2014 | 6 | 29 | 180 | 911 | SMC | 60 | X | 53 | 1700 | 1550 |  | 2 | Y |  | 900 | 800
2014 | 6 | 30 | 181 | 912 | DP | 60 | X | 52 | 1700 | 1550 |  | 2 | Y |  | 900 | 800
2014 | 7 | 1 | 182 | 913 | SMC | 60 | x | 52 | 1700 | 1550 |  | 2 | y |  | 900 | 800
2014 | 7 | 2 | 183 | 914 | DP | 60 | x | 52 | 1700 | 1550 |  | 2 | y |  | 900 | 800
2014 | 7 | 3 | 184 | 915 | SMC | 60 | x | 53 | 1700 | 1550 |  | 2 | y |  | 900 | 800
2014 | 7 | 4 | 185 | 916 | SMC | 60 | x | 53 | 1700 | 1550 |  | 2 | y |  | 900 | 800
2014 | 7 | 5 | 186 | 917 | DP | 60 | x | 52 | 1700 | 1550 |  | 2 | Y |  | 900 | 800
2014 | 7 | 6 | 187 | 918 | DP | 60 | x | 53 | 1690 | 1525 |  | 2 | y |  | 900 | 800
2014 | 7 | 7 | 188 | 919 | DP | 60 | x | 53 | 1690 | 1525 |  | 2 | y |  | 900 | 800
2014 | 7 | 8 | 189 | 920 | SMC | 60 | x | 53 | 1690 | 1525 |  | 2 | y |  | 900 | 800
2014 | 7 | 9 | 190 | 921 | SMC | 60 | x | 54 | 1690 | 1525 |  | 2 | y |  | 900 | 790
2014 | 7 | 10 | 191 | 922 | DP | 60 | x | 54 | 1690 | 1525 |  | 2 | y |  | 900 | 790
2014 | 7 | 11 | 192 | 923 | DP | 60 | x | 54 | 1690 | 1525 |  | 2 | y |  | 900 | 790
2014 | 7 | 12 | 193 | 924 | SMC | 60 | X | 54 | 1690 | 1525 |  | 2 | Y |  | 900 | 790
2014 | 7 | 13 | 194 | 925 | SMC | 60 | x | 52 | 1690 | 1525 |  | 2 | y |  | 900 | 790
2014 | 7 | 14 | 195 | 926 | DP | 60 | x | 52 | 1690 | 1525 |  | 2 | y |  | 900 | 790
2014 | 7 | 15 | 196 | 927 | DP | 60 | x | 52 | 1690 | 1525 |  | 2 | y |  | 900 | 790
2014 | 7 | 16 | 197 | 928 | DP/SMC | 60 | x | 52 | 1690 | 1525 |  | 2 | y |  | 900 | 790
2014 | 7 | 17 | 198 | 929 | SMC | 60 | x | 53 | 1690 | 1525 |  | 2 | y |  | 900 | 790
2014 | 7 | 18 | 199 | 930 | DP | 60 | x | 52 | 1690 | 1525 |  | 2 | y |  | 900 | 790
2014 | 7 | 19 | 200 | 931 | SMC | 60 | x` | 52 | 1690 | 1525 |  | 2 | y |  | 900 | 790
2014 | 7 | 20 | 201 | 932 | DP | 60 | x` | 53 | 1690 | 1525 |  | 2 | y |  | 900 | 790
2014 | 7 | 21 | 202 | 933 | SMC | 60 | x | 53 | 1690 | 1525 |  | 2 | Y |  | 900 | 790
2014 | 7 | 22 | 203 | 934 | DP | 60 | x | 52 | 1690 | 1515 |  | 2 | y |  | 900 | 790
2014 | 7 | 23 | 204 | 935 | SMC | 60 | x | 53 | 1690 | 1515 |  | 2 | y |  | 900 | 790
2014 | 7 | 24 | 205 | 936 | DP | 60 | x | 52 | 1690 | 1515 |  | 2 | y |  | 900 | 790
2014 | 7 | 25 | 206 | 937 | SMC | 60 | x | 53 | 1690 | 1515 |  | 2 | y |  | 900 | 790
2014 | 7 | 26 | 207 | 938 | DP | 60 | x | 53 | 1690 | 1515 |  | 2 | y |  | 900 | 790 | [11 empty] | Water refill
2014 | 7 | 27 | 208 | 939 | DP | 60 | x | 53 | 1690 | 1515 |  | 2 | y |  | 900 | 790
2014 | 7 | 28 | 209 | 940 | SMC | 60 | x | 52 | 1690 | 1515 |  | 2 | y |  | 900 | 790
2014 | 7 | 29 | 210 | 941 | SMC | 60 | x | 52 | 1690 | 1515 |  | 2 | y |  | 900 | 790
2014 | 7 | 30 | 211 | 942 | SMC | 60 | x | 52 | 1690 | 1515 |  | 2 | y |  | 900 | 790
2014 | 7 | 31 | 212 | 943 | DP | 60 | x | 52 | 1690 | 1515 |  | 2 | y |  | 900 | 790
2014 | 8 | 1 | 213 | 944 | SMC | 60 | x | 53 | 1690 | 1515 |  | 2 | y |  | 900 | 790
2014 | 8 | 2 | 214 | 945 | SMC | 60 | x | 53 | 1690 | 1515 |  | 2 | y |  | 900 | 790
2014 | 8 | 3 | 215 | 946 | DP | 60 | x | 53 | 1690 | 1515 |  | 2 | y |  | 900 | 790
2014 | 8 | 4 | 216 | 947 | SMC | 60 | x | 52 | 1690 | 1515 |  | 2 | y |  | 900 | 790 | [11 empty] | water refill
2014 | 8 | 5 | 217 | 948 | SMC | 60 | x | 53 | 1690 | 1515 |  | 2 | y |  | 900 | 790
2014 | 8 | 6 | 218 | 949 | DP | 60 | x | 52 | 1690 | 1515 |  | 2 | y |  | 900 | 790
2014 | 8 | 7 | 219 | 950 | SMC | 60 | x | 52 | 1650 | 1515 |  | 2 | y |  | 900 | 790
2014 | 8 | 8 | 220 | 951 | SMC | 60 | x | 52 | 1650 | 1515 |  | 2 | y |  | 900 | 790
2014 | 8 | 9 | 221 | 952 | SMC | 60 | x | 52 | 1650 | 1515 |  | 2 | y |  | 900 | 790 | [11 empty] | water refill
2014 | 8 | 10 | 222 | 953 | SMC | 60 | x | 53 | 1650 | 1515 |  | 2 | y |  | 900 | 790
2014 | 8 | 11 | 223 | 954 | DP | 60 | x | 53 | 1650 | 1515 |  | 2 | y |  | 900 | 790
2014 | 8 | 12 | 224 | 955 | SMC | 60 | x | 53 | 1650 | 1500 |  | 2 | y |  | 900 | 750
2014 | 8 | 13 | 225 | 956 | SMC | 60 | x | 54 | 1650 | 1500 |  | 2 | y |  | 900 | 750
2014 | 8 | 14 | 226 | 957 | SMC | 60 | x | 53 | 1650 | 1500 |  | 2 | y |  | 900 | 750
2014 | 8 | 15 | 227 | 958 | DP | 60 | x | 53 | 1650 | 1500 |  | 2 | y |  | 900 | 750
2014 | 8 | 16 | 228 | 959 | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance
2014 | 8 | 17 | 229 | 960 | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance
2014 | 8 | 18 | 230 | 961 | DP | 60 | x | 53 | 1650 | 1500 | Maintenance | 2 | y | Maintenance | 900 | 750
2014 | 8 | 19 | 231 | 962 | SMC | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | [8 empty] | Maintenance
2014 | 8 | 20 | 232 | 963 | SMC | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance | Maintenance
2014 | 8 | 21 | 233 | 964 | DP | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 8 | 22 | 234 | 965 | SMC | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 8 | 23 | 235 | 966 | DP | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 8 | 24 | 236 | 967 | SMC | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 8 | 25 | 237 | 968 | SMC | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750 | [11 empty] | *added water
2014 | 8 | 26 | 238 | 969 | DP | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 8 | 27 | 239 | 970 | SMC | 60 | x | 47 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 8 | 28 | 240 | 971 | DP | 60 | x | 47 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 8 | 29 | 241 | 972 | DP | 60 | x | 47 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 8 | 30 | 242 | 973 | DP | 60 | x | 47 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 8 | 31 | 243 | 974 | DP | 60 | x | 47 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 1 | 244 | 975 | DP | 60 | x | 47 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 2 | 245 | 976 | SMC | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 3 | 246 | 977 | SMC | 60 | x | 47 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 4 | 247 | 978 | SMC | 60 | x | 47 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 5 | 248 | 979 | SMC | 60 | x | 46 | 1700 | 1550 |  | 2 | Y |  | 850 | 750
2014 | 9 | 6 | 249 | 980 | SMC | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 7 | 250 | 981 | SMC | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 8 | 251 | 982 | SMC | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 9 | 252 | 983 | DP | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750 | [11 empty] | *add water to hydrogen generator
2014 | 9 | 10 | 253 | 984 | SMC | 60 | x | 47 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 11 | 254 | 985 | SMC | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 12 | 255 | 986 | DP | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 13 | 256 | 987 | SMC | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 14 | 257 | 988 | DP | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 15 | 258 | 989 | DP | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 16 | 259 | 990 | DP | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 17 | 260 | 991 | DP | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 18 | 261 | 992 | DP | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750 | [11 empty] | Added water to hydrogen generator.
2014 | 9 | 19 | 262 | 993 | DP | 60 | x | 45 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 20 | 263 | 994 | SMC | 60 | x | 45 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 21 | 264 | 995 | SMC | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 22 | 265 | 996 | SMC | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 23 | 266 | 997 | SMC | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 24 | 267 | 998 | SMC | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 25 | 268 | 999 | DP | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 26 | 269 | 1000 | DP | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750 | [11 empty] | Added water to hydrogen generator.
2014 | 9 | 27 | 270 | 1001 | SMC | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 28 | 271 | 1002 | DP | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 29 | 272 | 1003 | SMC | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 9 | 30 | 273 | 1004 | SMC | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 850 | 750
2014 | 10 | 1 | 274 | 1005 | DP | 60 | x | 45 | 1700 | 1550 |  | 2 | y |  | 825 | 750
2014 | 10 | 2 | 275 | 1006 | SMC | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 825 | 750
2014 | 10 | 3 | 276 | 1007 | DP | 60 | x | 45 | 1700 | 1550 |  | 2 | y |  | 825 | 750
2014 | 10 | 4 | 277 | 1008 | SMC | 60 | x | 45 | 1700 | 1550 |  | 2 | y |  | 825 | 750
2014 | 10 | 5 | 278 | 1009 | SMC | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 825 | 750
2014 | 10 | 6 | 279 | 1010 | DP | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 825 | 750
2014 | 10 | 7 | 280 | 1011 | DP | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 825 | 750
2014 | 10 | 8 | 281 | 1012 | SMC | 60 | x | 47 | 1700 | 1550 |  | 2 | y |  | 825 | 750
2014 | 10 | 9 | 282 | 1013 | SMC | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 825 | 750
2014 | 10 | 10 | 283 | 1014 | DP | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 825 | 750
2014 | 10 | 11 | 284 | 1015 | HJ | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 825 | 725 | [11 empty] | added water at daily rounds in a.m. alarm went off around 1400, more water was added.Error "H2 Outlet ACT 1.0 PSI *LOCK*" continued at 2030, LR performed a hard reset. System appeards normal -HJ & LR
2014 | 10 | 12 | 285 | 1016 | HJ | 60 | x | 47 | 1700 | 1550 |  | 2 | y |  | 825 | 725
2014 | 10 | 13 | 286 | 1017 | HJ | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 825 | 725
2014 | 10 | 14 | 287 | 1018 | DP | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 825 | 725
2014 | 10 | 15 | 288 | 1019 | HJ | 60 | x | 46 | 1700 | 1550 |  | 2 | y |  | 800 | 725 | [11 empty] | Planned power outage from 12:10-12:30UTC for turnover training
2014 | 10 | 16 | 289 | 1020 | LR | 60 | x | 45 | 1700 | 1550 |  | 2 | y |  | 800 | 725
2014 | 10 | 17 | 290 | 1021 | HJ | 60 | X | 45 | 1690 | 1550 |  | 2 | Y |  | 800 | 725
2014 | 10 | 18 | 291 | 1022 | LR | 60 | x | 45 | 1690 | 1550 |  | 2 | y |  | 800 | 725
2014 | 10 | 19 | 292 | 1023 | HJ | 60 | X | 46 | 1690 | 1550 |  | 2 | Y |  | 800 | 725
2014 | 10 | 20 | 293 | 1024 | HJ | 60 | X | 46 | 1690 | 1550 |  | 2 | Y |  | 800 | 725
2014 | 10 | 21 | 294 | 1025 | HJ | 60 | X | 45 | 1690 | 1550 |  | 2 | Y |  | 800 | 725
2014 | 10 | 22 | 295 | 1026 | HJ | 60 | x | 45 | 1690 | 1550 |  | 2 | y |  | 800 | 725
2014 | 10 | 23 | 296 | 1027 | HJ | 60 | X | locked | 1690 | 1550 |  | 2 | y |  | 800 | 725 | [11 empty] | H2 geneator reads "float chamber act1.4psi *lock* set 60psi flow". LR reset at 1140UTC, seems to be running fine since. At 1900UTC, switched to H2 cylinder (1000psi). Descant in scrubber appeared clear suggesting that it is no longer working. LR
2014 | 10 | 24 | 297 | 1028 | HJ |  |  |  | 1690 | 1550 |  | 2 | y |  | 800 | 725 | [11 empty] | H2 cyliner at 1000psi/45psi
2014 | 10 | 25 | 298 | 1029 | LR |  |  |  | 1690 | 1550 |  | 2 | y |  | 800 | 725 | [11 empty] | H2 cyliner at 1000psi/45psi
2014 | 10 | 26 | 299 | 1030 | LR |  |  |  | 1690 | 1550 |  | 2 | y |  | 800 | 725 | [11 empty] | H2 cyliner at 1000psi/45psi
2014 | 10 | 27 | 300 | 1031 | HJ |  |  |  | 1690 | 1550 |  | 2 | y |  | 800 | 725 | [11 empty] | H2 cyliner at 1000psi/45psi
2014 | 10 | 28 | 301 | 1032 | LR |  |  |  | 1690 | 1550 |  | 2 | y |  | 800 | 725 | [11 empty] | H2 cyliner at 1000psi/45psi
2014 | 10 | 29 | 302 | 1033 | LR |  |  |  | 1690 | 1550 |  | 2 | y |  | 800 | 725 | [11 empty] | H2 cyliner at 950psi/45psi
2014 | 10 | 30 | 303 | 1034 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 725 | [11 empty] | H2 cyliner at 900psi/45psi
2014 | 10 | 31 | 304 | 1035 | LR |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 725 | [11 empty] | H2 cyliner at 900psi/45psi
2014 | 11 | 1 | 305 | 1036 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 725 | [11 empty] | H2 cyliner at 890psi/45psi
2014 | 11 | 2 | 306 | 1037 | LR |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 725 | [11 empty] | H2 cyliner at 850psi/45psi
2014 | 11 | 3 | 307 | 1038 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 725 | [11 empty] | H2 cyliner at 800psi/45psi
2014 | 11 | 4 | 308 | 1039 | LR |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 725 | [11 empty] | H2 cyliner at 800psi/45psi
2014 | 11 | 5 | 309 | 1040 | HJ |  |  |  | 1690 | 1550 |  | 2 | y |  | 800 | 725 | [11 empty] | H2 cylinder at 800/45psi
2014 | 11 | 6 | 310 | 1041 | HJ |  |  |  | 1690 | 1550 |  | 2 | y |  | 800 | 725 | [11 empty] | H2 cylinder at 775/45psi
2014 | 11 | 7 | 311 | 1042 | LR |  |  |  | 1690 | 1550 |  | 2 | y |  | 800 | 725 | [11 empty] | H2 cylinder at 775/45psi
2014 | 11 | 8 | 312 | 1043 | HJ |  |  |  | 1690 | 1550 |  | 2 | y |  | 800 | 725 | [11 empty] | H2 cylinder at 700/45psi
2014 | 11 | 9 | 313 | 1044 | HJ |  |  |  | 1690 | 1550 |  | 2 | y |  | 800 | 725 | [11 empty] | H2 cylinder at 650/45psi
2014 | 11 | 10 | 314 | 1045 | LR |  |  |  | 1690 | 1550 |  | 2 | y |  | 800 | 725 | [11 empty] | H2 cylinder at 650/45psi. Brief power outage at 0700UTC
2014 | 11 | 11 | 315 | 1046 | LR |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 725 | [11 empty] | H2 cylinder at 650/45psi. Planned power outages from 1259-1301UTC & 1303-1305UTC
2014 | 11 | 12 | 316 | 1047 | LR |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 725 | [11 empty] | H2 Cylinder at 650/45 psi
2014 | 11 | 13 | 317 | 1048 | HJ |  |  |  | 1690 | 1550 |  | 2 | y |  | 800 | 725 | [11 empty] | H2 cylinder at 600/45psi
2014 | 11 | 14 | 318 | 1049 | LR |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 725 | [11 empty] | Planned power outage from 1257-1301UTC. H2 cylinder at 600/45psi
2014 | 11 | 15 | 319 | 1050 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 725 | [11 empty] | H2 cylinder at 600/45psi
2014 | 11 | 16 | 320 | 1051 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 725 | [11 empty] | H2 cylinder at 575/45psi
2014 | 11 | 17 | 321 | 1052 | LR |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 725 | [11 empty] | H2 cylinder at 575/45psi
2014 | 11 | 18 | 322 | 1053 | LR |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 725 | [11 empty] | H2 cylinder at 550/45psi
2014 | 11 | 19 | 323 | 1054 | HJ | [22 empty] | 30+ knot winds, did not travel to TAWO
2014 | 11 | 20 | 324 | 1055 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 725 | [11 empty] | H2 cylinder at 500/45psi
2014 | 11 | 21 | 325 | 1056 | [23 empty] | 40+ knot winds, did not travel to TAWO
2014 | 11 | 22 | 326 | 1057 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 450/45psi
2014 | 11 | 23 | 327 | 1058 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 725 | [11 empty] | H2 cylinder at 425/45psi
2014 | 11 | 24 | 328 | 1059 | LR |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 725 | [11 empty] | H2 cylinder at 425/45psi
2014 | 11 | 25 | 329 | 1060 | LR |  |  |  | 1690 | 1550 |  | 2 | N |  | 800 | 725 | [11 empty] | H2 cylinder at 425/45psi. One pump not running
2014 | 11 | 26 | 330 | 1061 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 725 | [11 empty] | H2 cylinder at 350/45psi.
2014 | 11 | 27 | 331 | 1062 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 725 | [11 empty] | H2 cylinder at 300/45psi.
2014 | 11 | 28 | 332 | 1063 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 725 | [11 empty] | H2 cylinder at 300/45psi.
2014 | 11 | 29 | 333 | 1064 | HJ | [22 empty] | High winds, did not travel to TAWO
2014 | 11 | 30 | 334 | 1065 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 725 | [11 empty] | H2 cylinder at 300/45psi.
2014 | 12 | 1 | 335 | 1066 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 725 | [11 empty] | H2 cylinder at 250/45psi.
2014 | 12 | 2 | 336 | 1067 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 725 | [11 empty] | H2 cylinder at 200/45psi.
2014 | 12 | 3 | 337 | 1068 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 725 | [11 empty] | H2 cylinder at 200/45psi.
2014 | 12 | 4 | 338 | 1069 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 200/45psi.
2014 | 12 | 5 | 339 | 1070 | LR |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 200/45psi. Changed tank at 1500utc, new pressure 2150psi.
2014 | 12 | 6 | 340 | 1071 | LR |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 2050psi.
2014 | 12 | 7 | 341 | 1072 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 2000/50psi.
2014 | 12 | 8 | 342 | 1073 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 2000/47psi.
2014 | 12 | 9 | 343 | 1074 | LR |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 2000/47psi.
2014 | 12 | 10 | 344 | 1075 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 1950/50 psi.
2014 | 12 | 11 | 345 | 1076 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 1900/50 psi.
2014 | 12 | 12 | 346 | 1077 | LR |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 1900/50 psi. Adjusted delivery pressure to 45psi. Vestibule cold due to Westerly wind.
2014 | 12 | 13 | 347 | 1078 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 1850/40 psi.
2014 | 12 | 14 | 348 | 1079
2014 | 12 | 15 | 349 | 1080 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 1950/40 psi.
2014 | 12 | 16 | 350 | 1081 | LR |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 1850/45 psi.
2014 | 12 | 17 | 351 | 1082 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 1825/45 psi.
2014 | 12 | 18 | 352 | 1083
2014 | 12 | 19 | 353 | 1084 | LR |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 1875/40 psi.
2014 | 12 | 20 | 354 | 1085 | LR |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 1850/40 psi.
2014 | 12 | 21 | 355 | 1086
2014 | 12 | 22 | 356 | 1087 | LR |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 1825/42 psi.
2014 | 12 | 23 | 357 | 1088 | [23 empty] | Did not travel to TAWO
2014 | 12 | 24 | 358 | 1089 | HJ |  |  |  | 1650 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 1650/50psi
2014 | 12 | 25 | 359 | 1090 | LR |  |  |  | 1650 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 1750/45psi
2014 | 12 | 26 | 360 | 1091 | HJ |  |  |  | 1650 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 1700/45psi
2014 | 12 | 27 | 361 | 1092 | [23 empty] | High winds, did not travel to TAWO
2014 | 12 | 28 | 362 | 1093 | HJ |  |  |  | 1650 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 1600/45psi
2014 | 12 | 29 | 363 | 1094 | LR |  |  |  | 1650 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 1600/45psi
2014 | 12 | 30 | 364 | 1095 | HJ |  |  |  | 1650 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 1590/48psi
2014 | 12 | 31 | 365 | 1096 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 1600/43psi
2015 | 1 | 1 | 1 | 1097 | HJ |  |  |  | 1690 | 1550 |  | 2 | y |  | 800 | 700 | [11 empty] | H2 cylinder at 1600/45psi
2015 | 1 | 2 | 2 | 1098 | LR |  |  |  | 1690 | 1550 |  | 2 | y |  | 800 | 700 | [11 empty] | H2 cylinder at 1550/45psi
2015 | 1 | 3 | 3 | 1099 |  |  |  |  | 1690 | 1550 |  | 2 | y |  | 800 | 700 | [11 empty] | H2 cylinder at 1525/45psi
2015 | 1 | 4 | 4 | 1100 | LR |  |  |  | 1690 | 1550 |  | 2 | y |  | 800 | 690
2015 | 1 | 5 | 5 | 1101 | LR |  |  |  | 1690 | 1550 |  | 2 | y |  | 800 | 690
2015 | 1 | 6 | 6 | 1102 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 1500/45psi
2015 | 1 | 7 | 7 | 1103
2015 | 1 | 8 | 8 | 1104 | LR |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 1450/45psi
2015 | 1 | 9 | 9 | 1105 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 1450/45psi
2015 | 1 | 10 | 10 | 1106
2015 | 1 | 11 | 11 | 1107
2015 | 1 | 12 | 12 | 1108 | LR |  |  |  | 1690 | 1550 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 1400/45psi
2015 | 1 | 13 | 13 | 1109 | HJ |  |  |  | 1690 | 1550 |  | 2 | Y |  | 750 | 675 | [11 empty] | H2 cylinder at 1350/48psi
2015 | 1 | 14 | 14 | 1110 | LR |  |  |  | 1690 | 1550 |  | 2 | Y |  | 750 | 675 | [11 empty] | H2 cylinder at 1300/48psi
2015 | 1 | 15 | 15 | 1111 | LR |  |  |  | 1690 | 1550 |  | 2 | Y |  | 750 | 675 | [11 empty] | H2 cylinder at 1250/48psi
2015 | 1 | 16 | 16 | 1112 | HJ |  |  |  | 1620 | 1500 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 1300/48psi
2015 | 1 | 17 | 17 | 1113 |  |  |  |  | 1625 | 1500 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 1250/50psi
2015 | 1 | 18 | 18 | 1114 | [23 empty] | Did not travel to TAWO
2015 | 1 | 19 | 19 | 1115 | [23 empty] | High winds, did not travel to TAWO
2015 | 1 | 20 | 20 | 1116 |  |  |  |  | 1610 | 1500 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 1200/48
2015 | 1 | 21 | 21 | 1117 | [23 empty] | High winds, did not travel to TAWO
2015 | 1 | 22 | 22 | 1118 |  |  |  |  | 1610 | 1500 |  | 2 | y |  | 800 | 700 | [11 empty] | H2 cylinder at 1150/48 psi
2015 | 1 | 23 | 23 | 1119 | [23 empty] | High winds, did not travel to TAWO
2015 | 1 | 24 | 24 | 1120 |  |  |  |  | 1600 | 1500 |  | 2 | y |  | 800 | 700 | [11 empty] | H2 cylinder at 1100/50 psi
2015 | 1 | 25 | 25 | 1121
2015 | 1 | 26 | 26 | 1122 |  |  |  |  | 1600 | 1500 |  | 2 | y |  | 800 | 650 | [11 empty] | H2 cylinder at 1050/52psi
2015 | 1 | 27 | 27 | 1123 |  |  |  |  | 1450 | 1350 |  | 2 | Y |  | 700 | 600 | [11 empty] | TAWO temps freezing upon arrival. H2 cylinder at 1000/55 psi
2015 | 1 | 28 | 28 | 1124 |  |  |  |  | 1625 | 1500 |  | 2 | Y |  | 750 | 650 | [11 empty] | H2 cylinder at 1050/50psi
2015 | 1 | 29 | 29 | 1125 | HJ |  |  |  | 1625 | 1500 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 1000/55psi. Checked H2 generator, still holds pressure after TAWO freeze.
2015 | 1 | 30 | 30 | 1126
2015 | 1 | 31 | 31 | 1127 | LR |  |  |  | 1625 | 1500 |  | 2 | Y |  | 800 | 700 | [11 empty] | H2 cylinder at 900/55psi.
2015 | 2 | 1 | 32 | 1128
2015 | 2 | 2 | 33 | 1129 | HJ |  |  |  | 1625 | 1500 |  | 2 | Y |  | 750 | 700 | [11 empty] | H2 cylinder at 900/55psi.
2015 | 2 | 3 | 34 | 1130 | [23 empty] | Winds at 30+ knots, did not travel to TAWO
2015 | 2 | 4 | 35 | 1131 | HJ |  |  |  | 1625 | 1500 |  | 2 | Y |  | 750 | 700 | [11 empty] | H2 cylinder at 850/55psi.
2015 | 2 | 5 | 36 | 1132
2015 | 2 | 6 | 37 | 1133 | HJ |  |  |  | 1610 | 1500 |  | 2 | Y |  | 750 | 650 | [11 empty] | H2 cylinder at 800/57psi
2015 | 2 | 7 | 38 | 1134 | LR |  |  |  | 1610 | 1500 |  | 2 | Y |  | 750 | 650 | [11 empty] | H2 cylinder at 800/57psi
2015 | 2 | 8 | 39 | 1135
2015 | 2 | 9 | 40 | 1136
2015 | 2 | 10 | 41 | 1137 | HJ |  |  |  | 1610 | 1500 |  | 2 | y |  | 750 | 700 | [11 empty] | H2 cylinder at 725/58psi
2015 | 2 | 11 | 42 | 1138
2015 | 2 | 12 | 43 | 1139 | [23 empty] | Planned power outage 1305-1307, 1328-1331 UTC
2015 | 2 | 13 | 44 | 1140 | LR |  |  |  | 1610 | 1500 |  | 2 | y |  | 750 | 700 | [11 empty] | H2 cylinder at 725/58psi
2015 | 2 | 14 | 45 | 1141 | JJ |  |  |  | 1600 | 1500 |  | 2 | y |  | 750 | 660 | [11 empty] | 660/55 h2
2015 | 2 | 15 | 46 | 1142 | JJ/YT |  |  |  | 1600 | 1500 |  | 2 | y |  | 750 | 660 | [11 empty] | 650/55
2015 | 2 | 16 | 47 | 1143 | JJ/YT |  |  |  | 1600 | 1500 |  | 2 | Y |  | 750 | 650 | [11 empty] | 600/57
2015 | 2 | 17 | 48 | 1144 | JJ/YT |  |  |  | 1600 | 1490 |  | 2 | y |  | 750 | 650 | [11 empty] | 570/57
2015 | 2 | 18 | 49 | 1145
2015 | 2 | 19 | 50 | 1146 | YT |  |  |  | 1640 | 1510 |  | 2 | y |  | 750 | 650 | [11 empty] | 550/55
2015 | 2 | 20 | 51 | 1147 | YT/JJ |  |  |  | 1650 | 1550 |  | 2 | y |  | 750 | 660 | [11 empty] | 550/55
2015 | 2 | 21 | 52 | 1148 | YT/JJ |  |  |  | 1600 | 1500 |  | 2 | y |  | 750 | 650 | [11 empty] | 540/56
2015 | 2 | 22 | 53 | 1149
2015 | 2 | 23 | 54 | 1150 | JJ/YT |  |  |  | 1600 | 1480 |  | 2 | y |  | 750 | 650 | [11 empty] | 500/60
2015 | 2 | 24 | 55 | 1151 | YT/JJ |  |  |  | 1600 | 1490 |  | 2 | y |  | 750 | 650 | [11 empty] | 450/57
2015 | 2 | 25 | 56 | 1152 | YT/JJ |  |  |  | 1600 | 1480 |  | 2 | y |  | 730 | 650 | [11 empty] | 450/57
2015 | 2 | 26 | 57 | 1153 | YT/JJ |  |  |  | 1600 | 1490 |  | 2 | y |  | 750 | 650 | [11 empty] | 470/57
2015 | 2 | 27 | 58 | 1154 | YT/JJ |  |  |  | 1560 | 1430 |  | 2 | y |  | 740 | 630 | [11 empty] | 420/58
2015 | 2 | 28 | 59 | 1155
2015 | 3 | 1 | 60 | 1156 | YT/JJ |  |  |  | 1600 | 1500 |  | 2 | y |  | 750 | 650 | [11 empty] | 390/55
2015 | 3 | 2 | 61 | 1157 | YT/JJ |  |  |  | 1610 | 1480 |  | 2 | y |  | 740 | 650 | [11 empty] | 360/59
2015 | 3 | 3 | 62 | 1158 | YT/JJ |  |  |  | 1600 | 1500 |  | 2 | y |  | 740 | 650 | [11 empty] | 350/57. Replaced float & dessicant on H2 gen. Switched to H2 gen (60psi) at 17:30. H2 gen flow 45ml/min. Pressure gauge on the rack 35psi.
2015 | 3 | 4 | 63 | 1159 | YT/JJ | 60 | x | 48 | 1600 | 1480 |  | 2 | y |  | 740 | 640
2015 | 3 | 5 | 64 | 1160 | YT/JJ | 60 | x | 48 | 1600 | 1490 |  | 2 | y |  | 740 | 650
2015 | 3 | 6 | 65 | 1161 | YT/JJ | 60 | * | 49 | 1630 | 1500 |  | 2 | y |  | 730 | 650
2015 | 3 | 7 | 66 | 1162 | JJ | 60 | * | 47 | 1640 | 1520 |  | 2 | y |  | 740 | 650
2015 | 3 | 8 | 67 | 1163 | JJ/YT | 60 | * | 47 | 1620 | 1500 |  | 2 | y |  | 730 | 650
2015 | 3 | 9 | 68 | 1164 | YT/JJ | 60 | * | 48 | 1630 | 1510 |  | 2 | y |  | 750 | 650
2015 | 3 | 10 | 69 | 1165 | YT/JJ | 60 | * | 48 | 1620 | 1500 |  | 2 | y |  | 730 | 640
2015 | 3 | 11 | 70 | 1166
2015 | 3 | 12 | 71 | 1167 | YT/JJ | 60 | * | 54 | 1630 | 1500 |  | 2 | y |  | 740 | 650
2015 | 3 | 13 | 72 | 1168
2015 | 3 | 14 | 73 | 1169 | YT/JJ | 60 | * | 50 | 1600 | 1480 |  | 2 | y |  | 750 | 640
2015 | 3 | 15 | 74 | 1170
2015 | 3 | 16 | 75 | 1171 | YT/JJ | 60 | * | 49 | 1600 | 1490 |  | 2 | y |  | 720 | 650
2015 | 3 | 17 | 76 | 1172
2015 | 3 | 18 | 77 | 1173 | YT/JJ | 60 | * | 48 | 1600 | 1490 |  | 2 | y |  | 720 | 645
2015 | 3 | 19 | 78 | 1174 | YT/JJ | 60 | * | 49 | 1600 | 1470 |  | 2 | y |  | 750 | 630 | [11 empty] | 13:30 Refilled 1L water.
2015 | 3 | 20 | 79 | 1175 | YT/JJ | 60 | * | 47 | 1600 | 1450 |  | 2 | y |  | 710 | 640
2015 | 3 | 21 | 80 | 1176 | YT/JJ | 60 | * | 47 | 1600 | 1470 |  | 2 | y |  | 740 | 640
2015 | 3 | 22 | 81 | 1177
2015 | 3 | 23 | 82 | 1178 | YT/JJ | 60 | * | 46 | 1600 | 1470 |  | 2 | y |  | 710 | 640
2015 | 3 | 24 | 83 | 1179 | YT/JJ | 60 | * | 47 | 1600 | 1450 |  | 2 | y |  | 710 | 640
2015 | 3 | 25 | 84 | 1180 | YT/JJ | 60 | * | 46 | 1600 | 1450 |  | 2 | y |  | 710 | 640
2015 | 3 | 26 | 85 | 1181 | YT | 60 | * | 47 | 1600 | 1470 |  | 2 | y |  | 740 | 630
2015 | 3 | 27 | 86 | 1182 | YT/JJ | 60 | * | 47 | 1600 | 1480 |  | 2 | y |  | 740 | 630
2015 | 3 | 28 | 87 | 1183 | YT/JJ | 60 | * | 46 | 1600 | 1440 |  | 2 | y |  | 710 | 630
2015 | 3 | 29 | 88 | 1184
2015 | 3 | 30 | 89 | 1185 | YT/JJ | 60 | * | 47 | 1600 | 1480 |  | 2 | y |  | 730 | 630
2015 | 3 | 31 | 90 | 1186 | YT/JJ | 60 | * | 45 | 1610 | 1480 |  | 2 | y |  | 710 | 640
2015 | 4 | 1 | 91 | 1187 | YT/JJ | 60 | * | 46 | 1600 | 1480 |  | 2 | y |  | 730 | 620
2015 | 4 | 2 | 92 | 1188 | YT/JJ | 60 | * | 47 | 1600 | 1480 |  | 2 | y |  | 710 | 630
2015 | 4 | 3 | 93 | 1189 | YT/JJ | 60 | * | 47 | 1590 | 1470 |  | 2.1 | y |  | 730 | 620
2015 | 4 | 4 | 94 | 1190 | YT/JJ | 60 | * | 46 | 1590 | 1430 |  | 2.1 | y |  | 720 | 630
2015 | 4 | 5 | 95 | 1191
2015 | 4 | 6 | 96 | 1192 | YT/JJ | 60 | * | 46 | 1590 | 1420 |  | 2.1 | y |  | 710 | 630 | [11 empty] | 17:30 Refilled 1L water.
2015 | 4 | 7 | 97 | 1193 | YT | 60 | * | 46 | 1580 | 1460 |  | 2.1 | y |  | 720 | 610
2015 | 4 | 8 | 98 | 1194 | YT/JJ | 60 | * | 47 | 1580 | 1460 |  | 2.1 | y |  | 730 | 610
2015 | 4 | 9 | 99 | 1195 | JJ/YT | 60 | * | 46 | 1590 | 1430 |  | 2.2 | y |  | 720 | 620
2015 | 4 | 10 | 100 | 1196 | JJ | 60 | * | 47 | 1590 | 1440 |  | 2.1 | y |  | 720 | 620
2015 | 4 | 11 | 101 | 1197 | JJ/YT | 60 | * | 47 | 1590 | 1460 |  | 2.1 | y |  | 720 | 630
2015 | 4 | 12 | 102 | 1198
2015 | 4 | 13 | 103 | 1199 | JJ/YT | 60 | * | 46 | 1580 | 1470 |  | 2.1 | y |  | 720 | 610
2015 | 4 | 14 | 104 | 1200 | JJ/YT | 60 | * | 47 | 1580 | 1420 |  | 1.9 | y |  | 700 | 620
2015 | 4 | 15 | 105 | 1201
2015 | 4 | 16 | 106 | 1202 | YT/JJ | 60 | * | 47 | 1570 | 1460 |  | 2 | y |  | 710 | 620
2015 | 4 | 17 | 107 | 1203 | JJ/YT | 60 | * | 47 | 1580 | 1420 |  | 2 | y |  | 710 | 620
2015 | 4 | 18 | 108 | 1204 | JJ/YT | 60 | * | 48 | 1580 | 1470 |  | 2 | y |  | 710 | 620
2015 | 4 | 19 | 109 | 1205
2015 | 4 | 20 | 110 | 1206 | JJ/YT | 60 | * | 55 | 1580 | 1410 |  | 2 | Y |  | 700 | 620
2015 | 4 | 21 | 111 | 1207 | JJ/YT | 60 | * | 46 | 1580 | 1460 |  | 2 | y |  | 710 | 610
2015 | 4 | 22 | 112 | 1208 | JJ/YT | 60 | * | 47 | 1580 | 1410 |  | 2 | y |  | 690 | 610
2015 | 4 | 23 | 113 | 1209 | JJ/YT | 60 | * | 47 | 1580 | 1470 |  | 2 | y |  | 710 | 620
2015 | 4 | 24 | 114 | 1210 | JJ/YT | 60 | * | 49 | 1590 | 1420 |  | 2 | y |  | 700 | 620
2015 | 4 | 25 | 115 | 1211 | YT/JJ | 60 | refill | 47 | 1580 | 1460 |  | 2 | y |  | 710 | 620 | [11 empty] | 12:00 Refilled 1L water.
2015 | 4 | 26 | 116 | 1212 | YT/JJ | 60 | * | 47 | 1600 | 1480 |  | 2 | y |  | 730 | 610
2015 | 4 | 27 | 117 | 1213 | YT/JJ | 60 | * | 47 | 1600 | 1480 |  | 2 | y |  | 730 | 610
2015 | 4 | 28 | 118 | 1214 | JJ/YT | 60 | * | 47 | 1600 | 1500 |  | 2 | y |  | 700 | 620 | [11 empty] | Jacques arrived :)
2015 | 4 | 29 | 119 | 1215 | YT/JJ | 60 | * | 46 | 1600 | 1480 |  | 2 | y |  | 710 | 610
2015 | 4 | 30 | 120 | 1216 | JJ/YT | 60 | * | 47 | 1600 | 1500 |  | 2 | y |  | 700 | 1200 | [11 empty] | NMHC Std 2 replaced by Jacques.
2015 | 5 | 1 | 121 | 1217 | JJ/YT | 60 | * | 48 | 1590 | 1450 |  | 2 | y |  | 700 | 1230 | [11 empty] | H2 generator moved to under the desk and back on UPS.
2015 | 5 | 2 | 122 | 1218
2015 | 5 | 3 | 123 | 1219
2015 | 5 | 4 | 124 | 1220
2015 | 5 | 5 | 125 | 1221 | JJ/YT | 60 | * | 48 | 1570 | 1400 |  | 2 | y |  | 700 | 1260
2015 | 5 | 6 | 126 | 1222 | JJ/YT | 60 | * | 48 | 1560 | 1440 |  | 2.1 | y |  | 700 | 1250
2015 | 5 | 7 | 127 | 1223 | JJ/YT | 61 | * | 48 | 1560 | 1390 |  | 2.2 | y |  | 680 | 1260 | [11 empty] | Set pressure was accidentally 61, but reset it to 60.
2015 | 5 | 8 | 128 | 1224 | JJ/YT | 60 | * | 49 | 1560 | 1440 |  | 2.1 | y |  | 700 | 1250
2015 | 5 | 9 | 129 | 1225 | JJ/YT | 60 | * | 48 | 1570 | 1440 |  | 2.2 | y |  | 690 | 1250 | [11 empty] | 11:20 Refill IL deionized water.
2015 | 5 | 10 | 130 | 1226
2015 | 5 | 11 | 131 | 1227 | JJ/YT | 60 | * | 48 | 1560 | 1430 |  | 2 | y |  | 700 | 1250
2015 | 5 | 12 | 132 | 1228 | JJ/YT | 60 | * | 49 | 1540 | 1430 |  | 2.1 | y |  | 680 | 1240
2015 | 5 | 13 | 133 | 1229 | JJ/YT | 60 | * | 47 | 1560 | 1440 |  | 2.1 | y |  | 700 | 1250
2015 | 5 | 14 | 134 | 1230 | JJ/YT | 60 | * | 47 | 1560 | 1440 |  | 2.1 | y |  | 680 | 1240
2015 | 5 | 15 | 135 | 1231 | JJ | 60 | * | 48 | 1560 | 1430 |  | 2.1 | y |  | 680 | 1240
2015 | 5 | 16 | 136 | 1232
2015 | 5 | 17 | 137 | 1233
2015 | 5 | 18 | 138 | 1234 | YT | 60 | * | 47 | 1550 | 1430 |  | 2.1 | y |  | 690 | 1240
2015 | 5 | 19 | 139 | 1235 | YT | 60 | * | 49 | 1560 | 1430 |  | 2.1 | y |  | 690 | 1240
2015 | 5 | 20 | 140 | 1236 | JJ/YT | 60 | * | 48 | 1570 | 1430 |  | 2.1 | y |  | 670 | 1250
2015 | 5 | 21 | 141 | 1237 | JJ/YT | 60 | * | 48 | 1560 | 1430 |  | 2.2 | y |  | 670 | 1250
2015 | 5 | 22 | 142 | 1238 | JJ/YT | 60 | * | 48 | 1560 | 1430 |  | 2.2 | y |  | 670 | 1250
2015 | 5 | 23 | 143 | 1239 | YT/JJ | 60 | refilled | 49 | 1550 | 1430 |  | 2.1 | y |  | 690 | 1240 | [11 empty] | 12:54 Refilled 1L deionized H2O in H2 Generator
2015 | 5 | 24 | 144 | 1240
2015 | 5 | 25 | 145 | 1241 | JJ/YT | 60 | * | 48 | 1560 | 1430 |  | 2.2 | y |  | 680 | 1250
2015 | 5 | 26 | 146 | 1242
2015 | 5 | 27 | 147 | 1243 | JJ/YT | 60 | * | 49 | 1560 | 1430 |  | 2.1 | y |  | 670 | 1240
2015 | 5 | 28 | 148 | 1244 | JJ/YT | 60 | * | 49 | 1550 | 1430 |  | 2.1 | y |  | 690 | 1240
2015 | 5 | 29 | 149 | 1245 | JJ/YT | 60 | * | 48 | 1560 | 1430 |  | 2.1 | y |  | 670 | 1240
2015 | 5 | 30 | 150 | 1246
2015 | 5 | 31 | 151 | 1247
2015 | 6 | 1 | 152 | 1248 | JJ/YT | 60 | * | 48 | 1550 | 1430 |  | 2.1 | y |  | 690 | 1240
2015 | 6 | 2 | 153 | 1249 | JJ/YT | 60 | * | 48 | 1550 | 1430 |  | 2.1 | y |  | 660 | 1240 | [11 empty] | identified a burned yellow wire insulation inside electronics rack by the connector for H2O B.
2015 | 6 | 3 | 154 | 1250 | JJ/YT | 60 | * | 50 | 1550 | 1430 |  | 2.2 | Y |  | 660 | 1250 | [11 empty] | 16:30 Jacques shut down the system.17 UTC Replaced the crimp female socket for H2O B.
2015 | 6 | 4 | 155 | 1251 | CV/NS | 60 | * | 50 | 1550 | 1430 |  | 2.2 | y |  | 690 | 1250
2015 | 6 | 5 | 156 | 1252 | CV/NS | 60 | * | 50 | 1550 | 1440 |  | 2.2 | y |  | 650 | 1240
2015 | 6 | 6 | 157 | 1253 | CV/NS | 60 | * | 49 | 1550 | 1430 |  | 2.2 | y |  | 675 | 1250 | [11 empty] | planned generator test/power outage at 1659, bck on 1701, off at 1707, and bck on at 1708.
2015 | 6 | 7 | 158 | 1254 | CV/NS | 60 | * | 49 | 1540 | 1430 |  | 2.1 | y |  | 650 | 1240
2015 | 6 | 8 | 159 | 1255 | CV/NS | 60 | * | 48 | 1550 | 1425 |  | 2.3 | y |  | 650 | 1250
2015 | 6 | 9 | 160 | 1256 | CV/NS | 60 | * | 48 | 1520 | 1420 |  | 2.2 | y |  | 650 | 1240
2015 | 6 | 10 | 161 | 1257 | CV/NS | 60 | * | 49 | 1530 | 1430 |  | 2.2 | y |  | 650 | 1240
2015 | 6 | 11 | 162 | 1258 | CV/NS | 60 | * | 48 | 1530 | 1430 |  | 2.2 | y |  | 650 | 1240
2015 | 6 | 12 | 163 | 1259 | CV/NS | 60 | * | 47 | 1550 | 1430 |  | 2.3 | y |  | 675 | 1225 | [11 empty] | 11:24 Refilled 1/2 L deionized H2O in H2 Generator
2015 | 6 | 13 | 164 | 1260 | CV/NS | 60 | * | 50 | 1520 | 1420 |  | 2.2 | Y |  | 650 | 1230
2015 | 6 | 14 | 165 | 1261 | CV/NS | 60 | * | 49 | 1550 | 1420 |  | 2.2 | y |  | 675 | 1250
2015 | 6 | 15 | 166 | 1262 | CV/NS | 60 | * | 49 | 1520 | 1420 |  | 2.2 | y |  | 650 | 1240
2015 | 6 | 16 | 167 | 1263 | CV | 60 | * | 51 | 1520 | 1420 |  | 2.2 | y |  | 650 | 1240
2015 | 6 | 17 | 168 | 1264 | CV/NS | 60 | * | 52 | 1550 | 1425 |  | 2.2 | y |  | 650 | 1220
2015 | 6 | 18 | 169 | 1265 | CV | 60 | * | 52 | 1550 | 1425 |  | 2.2 | y |  | 675 | 1220
2015 | 6 | 19 | 170 | 1266 | CV/NS | 60 | * | 53 | 1520 | 1430 |  | 2.2 | y |  | 650 | 1220
2015 | 6 | 20 | 171 | 1267 | CV/NS | 60 | * | 52 | 1540 | 1420 |  | 2.2 | y |  | 650 | 1220
2015 | 6 | 21 | 172 | 1268 | NS | 60 | * | 52 | 1520 | 1420 |  | 2.2 | y |  | 650 | 1220
2015 | 6 | 22 | 173 | 1269 | CV/NS | 60 | * | 51 | 1550 | 1420 |  | 2.2 | y |  | 650 | 1220 | [11 empty] | 11:37 Refilled 1/2 L deionized H2O in H2 Generator
2015 | 6 | 23 | 174 | 1270 | NS | 60 | * | 51 | 1520 | 1420 |  | 2.2 | y |  | 650 | 1210
2015 | 6 | 24 | 175 | 1271 | NS | 60 | * | 57 | 1520 | 1420 |  | 2.2 | y |  | 650 | 1220
2015 | 6 | 25 | 176 | 1272 | NS | 60 | * | 57 | 1520 | 1420 |  | 2.2 | y |  | 650 | 1220
2015 | 6 | 26 | 177 | 1273 | NS | 60 | * | 50 | 1520 | 1420 |  | 2.2 | y |  | 650 | 1220
2015 | 6 | 27 | 178 | 1274 | NS | 60 | * | 53 | 1520 | 1420 |  | 2.2 | y |  | 650 | 1220
2015 | 6 | 28 | 179 | 1275 | NS/HJ | 60 | * | 53 | 1550 | 1400 |  | 2 | y |  | 675 | 1250
2015 | 6 | 29 | 180 | 1276 | NS/HJ | 60 | * | 51 | 1550 | 1425 |  | 2 | y |  | 650 | 1250 | [11 empty] | 13:10 Refilled 1 L deionized H2O in H2 Generator
2015 | 6 | 30 | 181 | 1277 | HJ | 60 | * | 52 | 1520 | 1420 |  | 2 | y |  | 650 | 1250 | [11 empty] | NS climbned tower 1108-1114 UTC. Reset zeno clock
2015 | 7 | 1 | 182 | 1278 | NS | 60 | * | 52 | 1510 | 1410 |  | 2.2 | y |  | 650 | 1220
2015 | 7 | 2 | 183 | 1279 | NS/HJ | 60 | * | 52 | 1500 | 1400 |  | 2 | y |  | 650 | 1200
2015 | 7 | 3 | 184 | 1280 | NS/HJ | 60 | * | 52 | 1500 | 1400 |  | 2 | y |  | 650 | 1200
2015 | 7 | 4 | 185 | 1281 | HJ | 60 | * | 53 | 1510 | 1440 |  | 2 | Y |  | 650 | 1200
2015 | 7 | 5 | 186 | 1282 | NS | 60 | * | 53 | 1510 | 1410 |  | 2.2 | 7 |  | 650 | 1220
2015 | 7 | 6 | 187 | 1283 | HJ/NS | 60 | * | 52 | 1510 | 1400 |  | 2.2 | y |  | 650 | 1200
2015 | 7 | 7 | 188 | 1284 | NS | 60 | * | 52 | 1510 | 1400 |  | 2.2 | y |  | 650 | 1200
2015 | 7 | 8 | 189 | 1285 | HJ | 60 | * | 53 | 1500 | 1400 |  | 2.2 | y |  | 650 | 1200
2015 | 7 | 9 | 190 | 1286 | NS | 60 | * | 53 | 1510 | 1400 |  | 2.2 | y |  | 650 | 1200
2015 | 7 | 10 | 191 | 1287 | NS | 60 | * | 53 | 1500 | 1400 |  | 2 | y |  | 650 | 1200
2015 | 7 | 11 | 192 | 1288 | NS | 60 | * | 53 | 1510 | 1400 |  | 2.2 | y |  | 650 | 1200 | [11 empty] | 11:05 Refilled 1 L deionized H2O in H2 Generator
2015 | 7 | 12 | 193 | 1289 | HJ | 60 | * | 53 | 1500 | 1400 |  | 2.2 | y |  | 650 | 1200 | [11 empty] | Planned power outage 1600-1602, 1613-1614
2015 | 7 | 13 | 194 | 1290 | NS/HJ | 60 | * | 52 | 1500 | 1400 |  | 2 | y |  | 650 | 1200
2015 | 7 | 14 | 195 | 1291 | NS | 60 | * | 52 | 1500 | 1400 |  | 2.2 | y |  | 650 | 1200
2015 | 7 | 15 | 196 | 1292 | NS | 60 | * | 52 | 1510 | 1400 |  | 2.2 | y |  | 650 | 1200
2015 | 7 | 16 | 197 | 1293 | HJ | 60 | * | 53 | 1500 | 1400 |  | 2 | Y |  | 650 | 1200 | [11 empty] | HJ climbed tower 1140-1150utc
2015 | 7 | 17 | 198 | 1294 | NS | 60 | * | 52 | 1500 | 1400 |  | 2 | y |  | 650 | 1200 | [11 empty] | HJ & NS climbed tower 1133-1315UT
2015 | 7 | 18 | 199 | 1295 | HJ/NS/AC | 60 | * | 52 | 1500 | 1400 |  | 2 | y |  | 650 | 1200 | [11 empty] | NS climbed tower 1055-1105 utc
2015 | 7 | 19 | 200 | 1296 | HJ/NS | OFFLINE | [21 empty] | Jacques powered down at 2036utc
2015 | 7 | 20 | 201 | 1297 |  | OFFLINE | [21 empty] | Turned of cugc computer 1117, switched off H2 generator 1156
2015 | 7 | 21 | 202 | 1298 |  | OFFLINE | [21 empty] | powered off Computer 2 1235 utc
2015 | 7 | 22 | 203 | 1299 |  | OFFLINE
2015 | 7 | 23 | 204 | 1300 |  | OFFLINE
2015 | 7 | 24 | 205 | 1301 |  | OFFLINE
2015 | 7 | 25 | 206 | 1302 | hj/ns | OFFLINE | [21 empty] | instrument offline for TAWO work
2015 | 7 | 26 | 207 | 1303 | hj/ns | OFFLINE
2015 | 7 | 27 | 208 | 1304 | hj/ns | OFFLINE
2015 | 7 | 28 | 209 | 1305 | hj/ns | OFFLINE
2015 | 7 | 29 | 210 | 1306 | hj/ns | OFFLINE
2015 | 7 | 30 | 211 | 1307 | HJ | OFFLINE | [21 empty] | 1L of DI water to H2 generator.
2015 | 7 | 31 | 212 | 1308 | hj/ns | OFFLINE
2015 | 8 | 1 | 213 | 1309 | hj/ns | OFFLINE
2015 | 8 | 2 | 214 | 1310 | NS | 60 | * | 55 | 1500 | 1400 |  |  | 2 |  | 650 | 1200 | [11 empty] | Still offline, just checking tank pressures
2015 | 8 | 3 | 215 | 1311 | HJ | 60 | * | 53 | 1500 | 1400 |  | OFFLINE | 2 |  | 650 | 1200 | [11 empty] | Still offline, just checking tank pressures. NS climbed tower 1110-1120utc.Heavy equpiment at TAWO 1150-1915utc
2015 | 8 | 4 | 216 | 1312
2015 | 8 | 5 | 217 | 1313 | NS | 60 | * | 51 | 1500 | 1400 |  |  | 2 |  | 650 | 1200 | [11 empty] | Offline, but expecting to restart later today
2015 | 8 | 6 | 218 | 1314 | HJ | 60 | * | 51 | 1500 | 1400 |  | 2 | y |  | 650 | 1200 | [11 empty] | restarted yesterday remotely by Domi
2015 | 8 | 7 | 219 | 1315 | NS | 60 | * | 49 | 1500 | 1400 |  | 2 | y |  | 650 | 1200
2015 | 8 | 8 | 220 | 1316 | NS/HJ | 60 | * | 49 | 1500 | 1400 |  | 2 | y |  | 650 | 1200 | [11 empty] | NS & HJ on tower 1215-1445utc for inlet raising.
2015 | 8 | 9 | 221 | 1317 | HJ | 60 | * | 46 | 1500 | 1400 |  | 2 | Y |  | 650 | 1200
2015 | 8 | 10 | 222 | 1318 | NS/HJ | 60 | * | 46 | 1500 | 1400 |  | 2 | Y |  | 650 | 1200
2015 | 8 | 11 | 223 | 1319 | NS/HJ | 60 | * | 46 | 1500 | 1400 |  | 2 | y |  | 650 | 1200
2015 | 8 | 12 | 224 | 1320 | NS/HJ | 60 | * | 46 | 1500 | 1400 |  | 2 | Y |  | 650 | 1200 | [11 empty] | Closed valve on Methane Std 1, 22 psi left in regulator.
2015 | 8 | 13 | 225 | 1321 | HJ | 60 | * | 46 | 1510 | 1410 |  | 2 | y |  | 650 | 1200
2015 | 8 | 14 | 226 | 1322 | NS | 60 | * | 46 | 1510 | 1400 |  | 2 | Y |  | 650 | 1200 | [11 empty] | NS & HJ climbed otwer 1145-1345utc, raised inlet.
2015 | 8 | 15 | 227 | 1323 | HJ | 60 | Added1L | 46 | 1500 | 1400 |  | 2 | y |  | 650 | 1200 | [11 empty] | Added 1L of DI water, DH arrived
2015 | 8 | 16 | 228 | 1324 | NS | 60 | * | 54 | 1500 | 1380 |  | 2 | y |  | 650 | 1200 | [11 empty] | Methane Std 1 @ 21 psi in regulator.
2015 | 8 | 17 | 229 | 1325
2015 | 8 | 18 | 230 | 1326
2015 | 8 | 19 | 231 | 1327
2015 | 8 | 20 | 232 | 1328 | HJ | 60 | * | 53 | 1690 | 1410 |  | 2 | y |  | 650 | 1200 | 80 | [5 empty] | y |  |  |  |  | NS climbed tower 1047-1057utc.
2015 | 8 | 21 | 233 | 1329 | HJ/NS | 60 | * | 52 | 1680 | 1400 |  | 2 | y |  | 650 | 1200 | 80 | [5 empty] | y
2015 | 8 | 22 | 234 | 1330 | NS | 60 | * | 52 | 1660 | 1390 |  | 2 | y |  | 650 | 1200 | 85 | [5 empty] | y
2015 | 8 | 23 | 235 | 1331 | HJ | 60 | * | 53 | 1660 | 1400 |  | 2 | Y |  | 650 | 1200 | 85 | [5 empty] | Y
2015 | 8 | 24 | 236 | 1332 | NS | 60 | * | 52 | 1660 | 1400 |  | 2 | Y |  | 650 | 1200 | 85 | [5 empty] | Y |  |  |  |  | HJ climbed tower 11:30-11:42utc. Removed snow from inlet.
2015 | 8 | 25 | 237 | 1333 | NS | 60 | * | 53 | 1660 | 1400 |  | 2 | Y |  | 650 | 1200 | 85 | [5 empty] | Y
2015 | 8 | 26 | 238 | 1334 | HJ | 60 | Added1L | 53 | 1660 | 1400 |  | 2 | Y |  | 640 | 1200 | 85 | [5 empty] | Y
2015 | 8 | 27 | 239 | 1335 | NS | 60 | * | 53 | 1660 | 1390 |  | 2.2 | Y |  | 650 | 1190 | 85 | [5 empty] | Y
2015 | 8 | 28 | 240 | 1336 | NS | 60 | * | 53 | 1660 | 1400 |  | 2.2 | Y |  | 640 | 1200 | 85 | [5 empty] | Y
2015 | 8 | 29 | 241 | 1337 | NS | 60 | * | 53 | 1660 | 1400 |  | 2.2 | Y |  | 640 | 1200 | 85 | [5 empty] | Y
2015 | 8 | 30 | 242 | 1338 | NS | 60 | * | 53 | 1660 | 1400 |  | 2.2 | Y |  | 640 | 1200 | 85 | [5 empty] | Y
2015 | 8 | 31 | 243 | 1339 | NS | 60 | * | 54 | 1660 | 1400 |  | 2.2 | Y |  | 640 | 1200 | 85 | [5 empty] | Y
2015 | 9 | 1 | 244 | 1340 | HJ | 60 | * | 52 | 1690 | 1400 |  | 2 | y |  | 640 | 1200 | 85 | [5 empty] | Y
2015 | 9 | 2 | 245 | 1341 | NS | 60 | * | 53 | 1660 | 1400 |  | 2.2 | Y |  | 640 | 1200 | 85 | [5 empty] | Y |  |  |  |  | Snow removed from inlet at 11:20UT. NS climbed tower 1115-1135utc
2015 | 9 | 3 | 246 | 1342 | NS | 60 | * | 53 | 1660 | 1400 |  | 2.2 | Y |  | 640 | 1200 | 85 | [5 empty] | Y
2015 | 9 | 4 | 247 | 1343 | HJ | 60 | * | 53 | 1650 | 1400 |  | 2 | Y |  | 640 | 1190 | 85 | [5 empty] | Y
2015 | 9 | 5 | 248 | 1344 | NS | 60 | * | 52 | 1650 | 1390 |  | 2.2 | Y |  | 640 | 1190 | 85 | [5 empty] | Y
2015 | 9 | 6 | 249 | 1345 | HJ | 60 | * | 53 | 1650 | 1390 |  | 2 | y |  | 640 | 1190 | 85 | [5 empty] | y
2015 | 9 | 7 | 250 | 1346 | HJ | 60 | * | 52 | 1650 | 1390 |  | 2 | y |  | 640 | 1200 | 85 | [5 empty] | y
2015 | 9 | 8 | 251 | 1347 | NS | 60 | * | 52 | 1650 | 1360 |  | 2.2 | Y |  | 640 | 1190 | 80 | [5 empty] | Y
2015 | 9 | 9 | 252 | 1348 | HJ | 60 | * | 52 | 1650 | 1375 |  | 2 | Y |  | 620 | 1190 | 85 | [5 empty] | Y
2015 | 9 | 10 | 253 | 1349 | NS/HJ | 60 | * | 52 | 1650 | 1375 |  | 2 | Y |  | 625 | 1190 | 85 | [5 empty] | Y |  |  |  |  | Added DI water to H2 Gen.
2015 | 9 | 11 | 254 | 1350 | NS | 60 | * | 54 | 1650 | 1360 |  | 2.2 | Y |  | 600 | 1190 | 85 | [5 empty] | Y
2015 | 9 | 12 | 255 | 1351 | HJ | 60 | * | 52 | 1650 | 1360 |  | 2 | Y |  | 610 | 1175 | 85 | [5 empty] | Y
2015 | 9 | 13 | 256 | 1352 | NS | 60 | * | 53 | 1650 | 1360 |  | 2.2 | Y |  | 600 | 1170 | 78 | [5 empty] | Y
2015 | 9 | 14 | 257 | 1353 | NS | 60 | * | 52 | 1650 | 1360 |  | 2 | Y |  | 600 | 1190 | 90 | [5 empty] | Y |  |  |  |  | HJ climbed tower from 1117-1131utc. Removed snow from inlet.
2015 | 9 | 15 | 258 | 1354 | NS | 60 | * | 52 | 1660 | 1360 |  | 2.2 | Y |  | 600 | 1190 | 90 | [5 empty] | Y
2015 | 9 | 16 | 259 | 1355 | NS | 60 | * | 53 | 1660 | 1360 |  | 2.2 | Y |  | 600 | 1190 | 90 | [5 empty] | Y
2015 | 9 | 17 | 260 | 1356 | HJ | 60 | * | 52 | 1650 | 1350 |  | 2 | y |  | 620 | 1190 | 85 | [5 empty] | y |  |  |  |  | NS climbed tower 1251-1300 utc, removed snow from inlet.
2015 | 9 | 18 | 261 | 1357 | HJ | 60 | * | 52 | 1650 | 1350 |  | 2 | y |  | 620 | 1190 | 85 | [5 empty] | y
2015 | 9 | 19 | 262 | 1358 | HJ | 60 | * | 51 | 1650 | 1340 |  | 2 | Y |  | 620 | 1190 | 85 | [5 empty] | y
2015 | 9 | 20 | 263 | 1359 | HJ | 60 | * | 53 | 1650 | 1340 |  | 2 | Y |  | 610 | 1190 | 85 | [5 empty] | Y
2015 | 9 | 21 | 264 | 1360 | NS | 60 | * | 52 | 1650 | 1340 |  | 2.2 | Y |  | 610 | 1190 | 85 | [5 empty] | Y
2015 | 9 | 22 | 265 | 1361 | NS | 60 | * | 52 | 1650 | 1340 |  | 2.2 | Y |  | 610 | 1190 | 90 | [5 empty] | Y
2015 | 9 | 23 | 266 | 1362 | HJ | 60 | Added1L | 52 | 1650 | 1330 |  | 2 | y |  | 610 | 1190 | 85 | [5 empty] | Y
2015 | 9 | 24 | 267 | 1363 | HJ | 60 | * | 52 | 1650 | 1325 |  | 2 | Y |  | 610 | 1190 | 85 | [5 empty] | Y
2015 | 9 | 25 | 268 | 1364 | NS | 60 | * | 53 | 1650 | 1320 |  | 2 | Y |  | 610 | 1190 | 85 | [5 empty] | Y
2015 | 9 | 26 | 269 | 1365 | NS | 60 | * | 52 | 1650 | 1320 |  | 2 | Y |  | 610 | 1190 | 85 | [5 empty] | Y
2015 | 9 | 27 | 270 | 1366 | NS | 60 | * | 52 | 1650 | 1320 |  | 2 | Y |  | 610 | 1190 | 90 | [5 empty] | Y
2015 | 9 | 28 | 271 | 1367 | NS/HJ | 60 | * | 52 | 1660 | 1320 |  | 2 | Y |  | 610 | 1190 | 85 | [5 empty] | Y
2015 | 9 | 29 | 272 | 1368 | NS | 60 | * | 52 | 1660 | 1310 |  | 2 | Y |  | 610 | 1170 | 85 | [5 empty] | Y
2015 | 9 | 30 | 273 | 1369 | NS | 60 | * | 51 | 1660 | 1310 |  | 2 | Y |  | 610 | 1170 | 85 | [5 empty] | Y
2015 | 10 | 1 | 274 | 1370 | HJ | 60 | * | 52 | 1650 | 1300 |  | 2 | Y |  | 610 | 1170 | 85 | [5 empty] | Y
2015 | 10 | 2 | 275 | 1371 | HJ | 60 | * | 52 | 1650 | 1300 |  | 2 | Y |  | 600 | 1150 | 85 | [5 empty] | Y |  |  |  |  | HJ climber tower 1151-1208 UTC
2015 | 10 | 3 | 276 | 1372 | NS | 60 | * | 52 | 1640 | 1280 |  | 2 | Y |  | 600 | 1160 | 80 | [5 empty] | Y
2015 | 10 | 4 | 277 | 1373 | HJ | 60 | * | 53 | 1640 | 1280 |  | 2 | Y |  | 600 | 1160 | 85 | [5 empty] | Y
2015 | 10 | 5 | 278 | 1374 | NS | 60 | * | 53 | 1610 | 1250 |  | 2 | Y |  | 600 | 1140 | 85 | [5 empty] | Y
2015 | 10 | 6 | 279 | 1375 | NS/HJ | 60 | * | 51 | 1650 | 1280 |  | 2 | Y |  | 610 | 1160 | 85 | [5 empty] | Y
2015 | 10 | 7 | 280 | 1376 | NS/HJ | 60 | * | 54 | 1625 | 1260 |  | 2 | Y |  | 600 | 1150 | 85 | [5 empty] | Y
2015 | 10 | 8 | 281 | 1377 | HJ | 60 | * | 53 | 1650 | 1260 |  | 2 | Y |  | 600 | 1160 | 85 | [5 empty] | Y |  |  |  |  | NS climbed tower 1101-1111utc
2015 | 10 | 9 | 282 | 1378 | HJ | 60 | * | 52 | 1650 | 1260 |  | 2 | Y |  | 600 | 1160 | 85 | [5 empty] | Y
2015 | 10 | 10 | 283 | 1379 | HJ, JG, NB | 60 | Added1L | 52 | 1640 | 1260 |  | 2 | y |  | 600 | 1160 | 90 | [5 empty] | y |  |  |  |  | NB/JG climbed tower 1315-1355utc. Removed snow from inlet.
2015 | 10 | 11 | 284 | 1380 | HJ, JG, NB | 60 | * | 52 | 1640 | 1260 |  | 2 | Y |  | 600 | 1150 | 90 | [5 empty] | Y |  |  |  |  | NB climbed tower 1322-1340z and cleaned frost off inlet.
2015 | 10 | 12 | 285 | 1381 | NS, JG, NB | 60 | * | 52 | 1640 | 1260 |  | 2 | Y |  | 600 | 1160 | 90 | [5 empty] | Y
2015 | 10 | 13 | 286 | 1382 | NB,JG | 60 | * | 52 | 1640 | 1250 |  | 2 | Y |  | 600 | 1150 | 90 | [5 empty] | Y
2015 | 10 | 14 | 287 | 1383 | NB,JG | 60 | * | 52 | 1640 | 1240 |  | 2 | Y |  | 600 | 1160 | 90 | [5 empty] | Y
2015 | 10 | 15 | 288 | 1384 | NB,JG | 60 | * | 52 | 1640 | 1240 |  | 2 | Y |  | 600 | 1150 | 90 | [5 empty] | Y |  |  |  |  | JG climbed tower 1203-1229z and cleaned frost off inlet.
2015 | 10 | 16 | 289 | 1385 | NB,JG | 60 | * | 52 | 1640 | 1240 |  | 2 | Y |  | 600 | 1150 | 90 | [5 empty] | Y
2015 | 10 | 17 | 290 | 1386 | NB,JG | 60 | * | 51 | 1640 | 1230 |  | 2 | Y |  | 600 | 1150 | 90 | [5 empty] | Y
2015 | 10 | 18 | 291 | 1387 | NB | 60 | * | 53 | 1640 | 1230 |  | 2 | Y |  | 600 | 1150 | 90 | [5 empty] | Y
2015 | 10 | 19 | 292 | 1388 | NB,JG | 60 | * | 52 | 1640 | 1230 |  | 2 | Y |  | 600 | 1150 | 90 | [5 empty] | Y
2015 | 10 | 20 | 293 | 1389 | NB,JG | 60 | * | 53 | 1640 | 1225 |  | 2 | Y |  | 600 | 1150 | 90 | [5 empty] | Y |  |  |  |  | Ran room air test 1157z to 1627z. NB/JG
2015 | 10 | 21 | 294 | 1390 | NB | 60 | * | 52 | 1640 | 1210 |  | 2 | Y |  | 590 | 1150 | 90 | [5 empty] | Y |  |  |  |  | NB climbed tower 1305-1320z and cleaned frost off inlet.
2015 | 10 | 22 | 295 | 1391 | NB | 60 | * | 52 | 1640 | 1210 |  | 2 | Y |  | 590 | 1160 | 90 | [5 empty] | Y
2015 | 10 | 23 | 296 | 1392 | NB | 60 | * | 52 | 1640 | 1210 |  | 2 | Y |  | 590 | 1160 | 90 | [5 empty] | Y |  |  |  |  | JG climbed tower 1125-1210z and cleaned frost off inlet.
2015 | 10 | 24 | 297 | 1393 | NB | 60 | * | 52 | 1640 | 1210 |  | 2 | Y |  | 590 | 1150 | 90 | [5 empty] | Y
2015 | 10 | 25 | 298 | 1394 | JG | 60 | Added1L | 54 | 1630 | 1210 |  | 2 | Y |  | 590 | 1150 | 90 | [5 empty] | Y
2015 | 10 | 26 | 299 | 1395 | NB | 60 | * | 53 | 1630 | 1200 |  | 2 | Y |  | 590 | 1150 | 90 | [5 empty] | Y
2015 | 10 | 27 | 300 | 1396 | NB | 60 | * | 52 | 1630 | 1200 |  | 2 | Y |  | 590 | 1150 | 90 | [5 empty] | Y
2015 | 10 | 28 | 301 | 1397 | NB | 60 | * | 53 | 1630 | 1200 |  | 2 | Y |  | 590 | 1160 | 90 | [5 empty] | Y
2015 | 10 | 29 | 302 | 1398 | JG | 60 |  | 53 | 1630 | 1200 |  | 2 | Y |  | 590 | 1160 | 90 | [5 empty] | Y
2015 | 10 | 30 | 303 | 1399 | NB | 60 | * | 53 | 1630 | 1190 |  | 2 | Y |  | 590 | 1160 | 90 | [5 empty] | Y
2015 | 10 | 31 | 304 | 1400 | JG | 60 | * | 53 | 1630 | 1190 |  | 2 | Y |  | 590 | 1160 | 90 | [5 empty] | Y
2015 | 11 | 1 | 305 | 1401 | NB | 60 | * | 52 | 1630 | 1170 |  | 2 | Y |  | 590 | 1160 | 90 | [5 empty] | Y
2015 | 11 | 2 | 306 | 1402 | JG | 60 |  | 52 | 1630 | 1170 |  | 2 | Y |  | 590 | 1160 | 90 | [5 empty] | Y
2015 | 11 | 3 | 307 | 1403
2015 | 11 | 4 | 308 | 1404
2015 | 11 | 5 | 309 | 1405 | NB | 60 | * | 52 | 1630 | 1160 |  | 2 | Y |  | 590 | 1150 | 90 | [5 empty] | Y
2015 | 11 | 6 | 310 | 1406 | JG | 60 | * | 53 | 1630 | 1160 |  | 2 | Y |  | 590 | 1150 | 90 | [5 empty] | Y
2015 | 11 | 7 | 311 | 1407 | NB | 60 | Added1L | 52 | 1630 | 1160 |  | 2 | Y |  | 590 | 1150 | 90 | [5 empty] | Y
2015 | 11 | 8 | 312 | 1408 | JG | 60 | * | 53 | 1630 | 1150 |  | 2 | Y |  | 580 | 1150 | 90 | [5 empty] | Y
2015 | 11 | 9 | 313 | 1409 | NB | 60 | * | 53 | 1630 | 1150 |  | 2 | Y |  | 580 | 1150 | 90 | [5 empty] | Y |  |  |  |  | Turned off Methane STD 2 at the tank per request of Jacques via Brendan.
2015 | 11 | 10 | 314 | 1410 | JG | 60 | * | 53 | 1630 | off |  | 2 | Y |  | 580 | 1150 | 90 | [5 empty] | Y
2015 | 11 | 11 | 315 | 1411 | NB | 60 | * | 52 | 1630 | off |  | 2 | Y |  | 580 | 1150 | 90 | [5 empty] | Y |  |  |  |  | JG climbed tower 1334-1400z and cleaned frost off inlets.
2015 | 11 | 12 | 316 | 1412 | JG | 60 | * | 52 | off | 1140 |  | 2 | Y |  | 580 | 1150 | 90 | [5 empty] | Y |  |  |  |  | Turned off Methane STD 1 off and STD 2 back on at 1220 UTC.
2015 | 11 | 13 | 317 | 1413 | NB | 60 | * | 53 | off | 1140 |  | 2 | Y |  | 580 | 1140 | 90 | [5 empty] | Y
2015 | 11 | 14 | 318 | 1414 | JG | 60 | * | 52 | off | 1140 |  | 2 | Y |  | 580 | 1140 | 90 | [5 empty] | Y |  |  |  |  | NB climbed tower 1305-1320z and cleaned frost from inlet.
2015 | 11 | 15 | 319 | 1415 | NB | 60 | * | 53 | off | 1130 |  | 2 | Y |  | 580 | 1140 | 90 | [5 empty] | Y
2015 | 11 | 16 | 320 | 1416 | JG | 60 | * | 52 | off | 1130 |  | 2 | Y |  | 580 | 1140 | 90 | [5 empty] | Y
2015 | 11 | 17 | 321 | 1417 | NB | 60 | * | 53 | off | 1130 |  | 2 | Y |  | 570 | 1140 | 90 | [5 empty] | Y |  |  |  |  | JG climbed tower 1237-1308z and removed frost from inlet.
2015 | 11 | 18 | 322 | 1418 | JG | 60 | * | 54 | off | 1120 |  | 2 | Y |  | 570 | 1140 | 90 | [5 empty] | Y
2015 | 11 | 19 | 323 | 1419 | NB | 60 | * | 53 | off | 1120 |  | 2 | Y |  | 570 | 1130 | 90 | [5 empty] | Y
2015 | 11 | 20 | 324 | 1420 | JG | 60 | * | 53 | off | 1120 |  | 2 | Y |  | 570 | 1130 | 90 | [5 empty] | Y
2015 | 11 | 21 | 325 | 1421 | NB | 60 | * | 53 | off | 1120 |  | 2 | Y |  | 570 | 1130 | 90 | [5 empty] | Y
2015 | 11 | 22 | 326 | 1422 | JG | 60 | * | 53 | off | 1120 |  | 2 | Y |  | 570 | 1130 | 90 | [5 empty] | Y
2015 | 11 | 23 | 327 | 1423 | NB | 60 | Added1L | 52 | off | 1120 |  | 2 | Y |  | 560 | 1120 | 90 | [5 empty] | Y
2015 | 11 | 24 | 328 | 1424 | JG | 60 | * | 52 | off | 1110 |  | 2 | Y |  | 560 | 1120 | 90 | [5 empty] | Y
2015 | 11 | 25 | 329 | 1425 | NB | 60 | * | 52 | off | 1100 |  | 2 | Y |  | 560 | 1120 | 90 | [5 empty] | Y
2015 | 11 | 26 | 330 | 1426 | JG | 60 | * | 52 | off | 1090 |  | 2 | Y |  | 560 | 1120 | 90 | [5 empty] | Y
2015 | 11 | 27 | 331 | 1427 | NB | 60 | * | 52 | off | 1090 |  | 2 | Y |  | 560 | 1120 | 90 | [5 empty] | Y
2015 | 11 | 28 | 332 | 1428 | JG | 60 | * | 53 | off | 1080 |  | 2 | Y |  | 560 | 1120 | 90 | [5 empty] | Y
2015 | 11 | 29 | 333 | 1429 | NB | 60 | * | 52 | off | 1080 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 11 | 30 | 334 | 1430 | JG | 60 | * | 53 | off | 1070 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 1 | 335 | 1431 | NB | 60 | * | 52 | off | 1070 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 2 | 336 | 1432 | JG | 60 |  | 53 | off | 1060 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 3 | 337 | 1433 | NB | 60 | * | 53 | off | 1060 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 4 | 338 | 1434 | JG | 60 | * | 52 | off | 1060 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 5 | 339 | 1435 | NB | 60 | * | 50 | off | 1060 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 6 | 340 | 1436 | JG | 60 | * | 52 | off | 1060 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 7 | 341 | 1437 | NB | 60 | Added1L | 52 | off | 1050 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y |  |  |  |  | NB climbed tower 1327-1342z and removed frost from inlet.
2015 | 12 | 8 | 342 | 1438 | JG | 60 | * | 52 | off | 1050 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 9 | 343 | 1439 | NB | 60 | * | 52 | off | 1040 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 10 | 344 | 1440 | JG | 60 | * | 52 | off | 1040 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 11 | 345 | 1441 | NB | 60 | * | 52 | off | 1040 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 12 | 346 | 1442 | JG | 60 | * | 52 | off | 1030 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 13 | 347 | 1443 | NB | 60 | * | 52 | off | 1030 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 14 | 348 | 1444 | JG | 60 | * | 52 | off | 1020 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 15 | 349 | 1445 | NB | 60 | * | 52 | off | 1020 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 16 | 350 | 1446 | JG | 60 | * | 52 | off | 1010 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 17 | 351 | 1447 | NB | 60 | * | 51 | off | 1010 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y |  |  |  |  | JG climbed tower 1235-1304z and removed frost from inlet.
2015 | 12 | 18 | 352 | 1448 | JG | 60 | * | 52 | off | 1000 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 19 | 353 | 1449 | NB | 60 | * | 52 | off | 1000 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 20 | 354 | 1450 | JG | 60 | * | 52 | off | 1000 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 21 | 355 | 1451 | NB | 60 | * | 51 | off | 1000 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 22 | 356 | 1452 | JG | 60 | * | 52 | off | 990 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 23 | 357 | 1453 | NB | 60 | Added1L | 52 | off | 980 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 24 | 358 | 1454 | JG | 60 | * | 52 | off | 980 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 25 | 359 | 1455 | NB | 60 | * | 52 | off | 980 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y |  |  |  |  | Hydrogen generator is calling for it 16000 hour/24 month service. This is a complete service requiring service kit MKH2PEM-24M. The unit is currently operating with no problems.
2015 | 12 | 26 | 360 | 1456 | JG | 60 | * | 52 | off | 970 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 27 | 361 | 1457 | NB | 60 | * | 51 | off | 970 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 28 | 362 | 1458 | JG | 60 | * | 52 | off | 970 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y
2015 | 12 | 29 | 363 | 1459 | NB | 60 | * | 52 | off | 970 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y |  |  |  |  | Switched to H2 tank and shut down H2 generator at ~1815z. H2 tank pressure 2300psi/line pressure 54psi. Found desiccant in H2 line filter to be spent. Replaced H2 line filter at ~1820z.
2015 | 12 | 30 | 364 | 1460 | JG | off | off | off | off | 970 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y |  |  |  |  | H2 tank pres 2300 psi; line press 50 psi.
2015 | 12 | 31 | 365 | 1461 | NB | off | off | off | off | 960 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y |  |  |  |  | H2 tank pres 2300 psi; line press 49 psi.
2016 | 1 | 1 | 1 | 1462 | JG | off | off | off | off | 960 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y |  |  |  |  | H2 tank pres 2300 psi; line pres 49 psi.
2016 | 1 | 2 | 2 | 1463 | NB | off | off | off | off | 960 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y |  |  |  |  | H2 tank pres 2300 psi; line pres 49 psi.
2016 | 1 | 3 | 3 | 1464 | JG | off | off | off | off | 950 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y |  |  |  |  | H2 tank pres 2250 psi; line pres 50 psi.
2016 | 1 | 4 | 4 | 1465 | NB | off | off | off | off | 950 |  | 2 | Y |  | 560 | 1110 | 90 | [5 empty] | Y |  | * |  |  | H2 tank pres 2250 psi; line pres 50 psi.
2016 | 1 | 5 | 5 | 1466 | JG | off | off | off | off | 950 |  | 2 | Y |  | 560 | 1110 | 95 | [5 empty] | Y |  |  |  |  | H2 tank pres 2250 psi; line pres 50 psi.
2016 | 1 | 6 | 6 | 1467 | NB | off | off | off | off | 940 |  | 2 | Y |  | 560 | 1110 | 95 | [5 empty] | Y |  |  |  |  | H2 tank pres 2240 psi; line pres 49 psi.
2016 | 1 | 7 | 7 | 1468 | JG | off | off | off | off | 930 |  | 2 | Y |  | 550 | 1110 | 95 | [5 empty] | Y |  |  |  |  | H2 tank pres 2200 psi; line pres 50 psi.
2016 | 1 | 8 | 8 | 1469 | NB | off | off | off | off | 930 |  | 2 | Y |  | 550 | 1110 | 95 | [5 empty] | Y |  |  |  |  | H2 tank pres 2150 psi; line pres 51 psi.
2016 | 1 | 9 | 9 | 1470 | JG | off | off | off | off | 930 |  | 2 | Y |  | 550 | 1110 | 95 | [5 empty] | Y |  |  |  |  | H2 tank pres 2070 psi; line pres 52 psi.
2016 | 1 | 10 | 10 | 1471 | NB | off | off | off | off | 920 |  | 2 | Y |  | 550 | 1110 | 95 | [5 empty] | Y |  |  |  |  | H2 tank pres 2050 psi; line pres 51 psi. JG climbed tower 1513-1530z and removed frost from inlet.
2016 | 1 | 11 | 11 | 1472 | JG | off | off | off | off | 910 |  | 2 | Y |  | 550 | 1110 | 95 | [5 empty] | Y |  | * |  |  | H2 tank pres 2070 psi; line pres 52 psi.
2016 | 1 | 12 | 12 | 1473 | NB | off | off | off | off | 910 |  | 2 | Y |  | 550 | 1110 | 95 | [5 empty] | Y |  |  |  |  | H2 tank pres 2060 psi; line pres 50 psi. JG climbed tower 1243-1307z and removed frost from inlet. Installed spare H2 generator at ~1615z. Shut off H2 tank - tank pres 2060psi.
2016 | 1 | 13 | 13 | 1474 | JG | 60 | * | 52 | off | 900 |  | 2 | Y |  | 550 | 1110 | 98 | [5 empty] | Y |  |  |  |  | Planned power outage for e-gen test 1307-1309 and 1321-1324 UTC.
2016 | 1 | 14 | 14 | 1475 | NB | 60 | added 1L | 53 | off | 900 |  | 2 | Y |  | 550 | 1110 | 97 | [5 empty] | Y
2016 | 1 | 15 | 15 | 1476 | JG | 60 | * | 54 | off | 900 |  | 2 | Y |  | 550 | 1110 | 98 | [5 empty] | Y |  |  |  |  | JG climbed tower 1327-1347 and cleaned frost from inlet.
2016 | 1 | 16 | 16 | 1477 | NB | 60 | * | 52 | off | 890 |  | 2 | Y |  | 550 | 1110 | 98 | [5 empty] | Y
2016 | 1 | 17 | 17 | 1478 | JG | 60 | * | 55 | off | 880 |  | 2 | Y |  | 550 | 1110 | 98 | [5 empty] | Y
2016 | 1 | 18 | 18 | 1479 | NB | 60 | * | 53 | off | 880 |  | 2 | Y |  | 550 | 1110 | 97 | [5 empty] | Y | * | *
2016 | 1 | 19 | 19 | 1480 | JG | 60 | * | 53 | off | 880 |  | 2 | Y |  | 550 | 1110 | 98 | [5 empty] | Y
2016 | 1 | 20 | 20 | 1481 | NB | 60 | * | 55 | off | 880 |  | 2 | Y |  | 550 | 1110 | 98 | [5 empty] | Y
2016 | 1 | 21 | 21 | 1482 | JG | 60 | * | 54 | off | 870 |  | 2 | Y |  | 550 | 1110 | 98 | [5 empty] | Y
2016 | 1 | 22 | 22 | 1483 | NB | 60 | * | 53 | off | 870 |  | 2 | Y |  | 550 | 1110 | 98 | [5 empty] | Y
2016 | 1 | 23 | 23 | 1484 | JG | 60 | * | 54 | off | 860 |  | 2 | Y |  | 550 | 1110 | 98 | [5 empty] | Y |  |  |  |  | NB climbed tower and removed frost from inlets 1217-1234z.
2016 | 1 | 24 | 24 | 1485 | NB | 60 | * | 53 | off | 860 |  | 2 | Y |  | 550 | 1100 | 98 | [5 empty] | Y
2016 | 1 | 25 | 25 | 1486 | JG | 60 | * | 53 | off | 860 |  | 2 | Y |  | 550 | 1100 | 98 | [5 empty] | Y | * | *
2016 | 1 | 26 | 26 | 1487 | NB | 60 | * | 54 | off | 850 |  | 2 | Y |  | 550 | 1100 | 98 | [5 empty] | Y
2016 | 1 | 27 | 27 | 1488 | JG | 60 | * | 54 | off | 840 |  | 2 | Y |  | 550 | 1100 | 98 | [5 empty] | Y
2016 | 1 | 28 | 28 | 1489 | NB | 60 | * | 54 | off | 840 |  | 2 | Y |  | 550 | 1100 | 98 | [5 empty] | Y
2016 | 1 | 29 | 29 | 1490 | JG | 60 | added 1L | 54 | off | 840 |  | 2 | Y |  | 550 | 1100 | 98 | [5 empty] | Y
2016 | 1 | 30 | 30 | 1491 | NB | 60 | * | 54 | off | 840 |  | 2 | Y |  | 550 | 1100 | 98 | [5 empty] | Y
2016 | 1 | 31 | 31 | 1492 | JG | 60 | * | 55 | off | 830 |  | 2 | Y |  | 550 | 1100 | 98 | [5 empty] | Y
2016 | 2 | 1 | 32 | 1493 | NB | 60 | * | 54 | off | 820 |  | 2 | Y |  | 550 | 1100 | 98 | [5 empty] | Y | * | * |  |  | JG climbed tower 1310-1327z and cleaned frost from inlet.
2016 | 2 | 2 | 33 | 1494 | JG | 60 | * | 55 | off | 810 |  | 2 | Y |  | 550 | 1100 | 100 | [5 empty] | Y
2016 | 2 | 3 | 34 | 1495 | NB | 60 | * | 56 | off | 810 |  | 2 | Y |  | 550 | 1100 | 100 | [5 empty] | Y
2016 | 2 | 4 | 35 | 1496 | JG | 60 | * | 55 | off | 810 |  | 2 | Y |  | 550 | 1100 | 100 | [5 empty] | Y |  |  |  |  | NB climbed tower 1224-1239z and removed frost from inlet.
2016 | 2 | 5 | 36 | 1497 | NB | 60 | * | 54 | off | 810 |  | 2 | Y |  | 550 | 1100 | 100 | [5 empty] | Y |  |  |  |  | JG climbed tower 1342-1359z and removed frost from inlet.
2016 | 2 | 6 | 37 | 1498 | JG | 60 | * | 54 | off | 800 |  | 2 | Y |  | 550 | 1100 | 100 | [5 empty] | Y |  |  |  |  | NB climbed tower 1407-1421z and removed frost from inlet.
2016 | 2 | 7 | 38 | 1499 | JJ/MB | 60 | * | 54 | off | 800 |  | 2 | Y |  | 530 | 1100 | 100 | [5 empty] | Y
2016 | 2 | 8 | 39 | 1500 | JJ/MB/SD | 60 | * | 55 | off | 800 |  | 2 | Y |  | 530 | 1100 | 100 | [5 empty] | Y | * | *
2016 | 2 | 9 | 40 | 1501 | JJ/MB/NB | 60 | * | 55 | off | 790 |  | 2 | Y |  | 530 | 1100 | 100 | [5 empty] | Y
2016 | 2 | 10 | 41 | 1502 | SWD/JG/NB | 60 | * | 54 | off | 780 |  | 2 | Y |  | 530 | 1100 | 100 | [5 empty] | Y
2016 | 2 | 11 | 42 | 1503 | JJ/MB | 60 | * | 54 | off | 770 |  | 2 | Y |  | 530 | 1100 | 100 | [5 empty] | Y
2016 | 2 | 12 | 43 | 1504 | JJ/MB | 60 | * | 55 | off | 770 |  | 2 | y |  | 530 | 1080 | 100 | [5 empty] | Y
2016 | 2 | 13 | 44 | 1505 | JJ/MB | 60 | * | 55 | off | 770 |  | 2 | Y |  | 530 | 1080 | 100 | [5 empty] | Y
2016 | 2 | 14 | 45 | 1506 | JJ/MB | 60 | added 1L | 55 | off | 760 |  | 2 | y |  | 530 | 1080 | 100 | [5 empty] | Y |  |  |  |  | Added 1L DI water to H2 generator. JJ climbed tower to clean instruments from 1737-1833z.
2016 | 2 | 15 | 46 | 1507 | JJ | 60 | * | 54 | off | 760 |  | 2 | Y |  | 525 | 1070 | 100 | [5 empty] | Y | * | *
2016 | 2 | 16 | 47 | 1508 | JJ/MB | 60 | * | 54 | off | 760 |  | 2 | y |  | 530 | 1080 | 100 | [5 empty] | y
2016 | 2 | 17 | 48 | 1509 | MB/JJ | 60 | * | 53 | off | 750 |  | 2 | Y |  | 530 | 1080 | 100 | [5 empty] | Y
2016 | 2 | 18 | 49 | 1510 | MB/JJ | 60 | * | 53 | off | 740 |  | 2 | Y |  | 520 | 1080 | 100 | [5 empty] | Y
2016 | 2 | 19 | 50 | 1511 | JJ/MB | 60 | * | 55 | off | 730 |  | 2 | Y |  | 510 | 1080 | 85 | [5 empty] | Y
2016 | 2 | 20 | 51 | 1512 | JJ/MB | 60 | * | 54 | off | 730 |  | 2 | Y |  | 530 | 1080 | 100 | [5 empty] | Y
2016 | 2 | 21 | 52 | 1513 | JJ/MB | 60 | * | 55 | off | 720 |  | 2 | Y |  | 520 | 1080 | 100 | [5 empty] | Y
2016 | 2 | 22 | 53 | 1514 | JJ/MB | 60 | * | 55 | off | 720 |  | 2 | Y |  | 510 | 1080 | 100 | [5 empty] | Y | * | *
2016 | 2 | 23 | 54 | 1515 | MB | 60 | * | 55 | off | 730 |  | 2 | Y |  | 520 | 1080 | 100 | [5 empty] | Y
2016 | 2 | 24 | 55 | 1516 | JJ/MB | 60 | * | 54 | off | 710 |  | 2 | Y |  | 510 | 1070 | 100 | [5 empty] | Y
2016 | 2 | 25 | 56 | 1517 | jj/mb | 60 | * | 55 | off | 710 |  | 2 | Y |  | 520 | 1070 | 100 | [5 empty] | Y
2016 | 2 | 26 | 57 | 1518 | JJ/MB | 60 | * | 55 | off | 710 |  | 2 | Y |  | 510 | 1070 | 100 | [5 empty] | Y
2016 | 2 | 27 | 58 | 1519 | MB | 60 | * | 55 | off | 700 |  | 2 | Y |  | 520 | 1080 | 100 | [5 empty] | Y
2016 | 2 | 28 | 59 | 1520 | JJ | 60 | * | 55 | off | 700 |  | 2 | y |  | 510 | 1070 | 100 | [5 empty] | Y
2016 | 2 | 29 | 60 | 1521 | JJ/MB | 60 | * | 57 | off | 700 |  | 2 | y |  | 510 | 1070 | 100 | [5 empty] | Y | * | *
2016 | 3 | 1 | 61 | 1522 | JJ/MB | 60 | * | 55 | off | 680 |  | 2 | Y |  | 510 | 1060 | 100 | [5 empty] | Y
2016 | 3 | 2 | 62 | 1523 | MB/JJ | 60 | added 1L | 56 | off | 680 |  | 2 | Y |  | 510 | 1060 | 100 | [5 empty] | Y |  |  |  |  | Added 1L DI water to H2 generator. JJ climbed tower to clean instruments from 1737-1833z.
2016 | 3 | 3 | 63 | 1524 | MB/JJ | 60 | * | 55 | off | 670 |  | 2 | Y |  | 500 | 1060 | 100 | [5 empty] | Y
2016 | 3 | 4 | 64 | 1525 | JJ/MB | 60 | * | 55 | off | 660 |  | 2 | Y |  | 510 | 1060 | 100 | [5 empty] | Y
2016 | 3 | 5 | 65 | 1526 | jj/mb | 60 | * | 56 | off | 660 |  | 2 | Y |  | 510 | 1070 | 100 | [5 empty] | Y
2016 | 3 | 6 | 66 | 1527 | MB | 60 | * | 55 | off | 660 |  | 2 | Y |  | 510 | 1070 | 100 | [5 empty] | Y
2016 | 3 | 7 | 67 | 1528 | jj/mb | 60 | * | 56 | off | 660 |  | 2 | Y |  | 510 | 1060 | 100 | [5 empty] | Y | * | *
2016 | 3 | 8 | 68 | 1529 | jj/mb | 60 | * | 56 | off | 650 |  | 2 | Y |  | 510 | 1060 | 100 | [5 empty] | Y
2016 | 3 | 9 | 69 | 1530 | JJ | 60 | * | 55 | off | 650 |  | 2 | y |  | 510 | 1070 | 100 | [5 empty] | Y
2016 | 3 | 10 | 70 | 1531 | [23 empty] | White out conditions (Condition 1) did not travel to TAWO this day.
2016 | 3 | 11 | 71 | 1532 | JJ | 60 | * | 55 | off | 630 |  | 2 | y |  | 510 | 1060 | 100 | [5 empty] | Y
2016 | 3 | 12 | 72 | 1533 | MB | 60 | * | 55 | off | 620 |  | 2 | Y |  | 510 | 1060 | 100 | [5 empty] | Y
2016 | 3 | 13 | 73 | 1534 | JJ | 60 | * | 55 | off | 620 |  | 2 | y |  | 510 | 1060 | 100 | [5 empty] | Y |  |  |  |  | unexpected power outage today from 13:16 to 13:23 UTC
2016 | 3 | 14 | 74 | 1535 | JJ/MB | 60 | * | 55 | off | 620 |  | 2 | y |  | 510 | 1070 | 100 | [5 empty] | Y | * | *
2016 | 3 | 15 | 75 | 1536 | JJ/MB | 60 | * | 57 | off | 610 |  | 2 | y |  | 500 | 1060 | 100 | [5 empty] | Y
2016 | 3 | 16 | 76 | 1537 | MB | 60 | * | 57 | off | 600 |  | 2 | Y |  | 500 | 1050 | 100 | [5 empty] | Y
2016 | 3 | 17 | 77 | 1538 | JJ/MB | 60 | * | 57 | off | 600 |  | 2 | Y |  | 500 | 1060 | 100 | [5 empty] | Y
2016 | 3 | 18 | 78 | 1539 | MB/JJ | 60 | * | 56 | off | 590 |  | 2 | Y |  | 500 | 1060 | 100 | [5 empty] | Y
2016 | 3 | 19 | 79 | 1540 | MB/JJ | 60 | * | 55 | off | 590 |  | 2 | Y |  | 500 | 1060 | 100 | [5 empty] | Y
2016 | 3 | 20 | 80 | 1541 | MB | 60 | * | 57 | off | 580 |  | 2 | Y |  | 500 | 1060 | 100 | [5 empty] | Y
2016 | 3 | 21 | 81 | 1542 | JJ/MB | 60 | added 1L | 55 | off | 580 |  | 2 | Y |  | 500 | 1060 | 100 | [5 empty] | Y | * | *
2016 | 3 | 22 | 82 | 1543 | MB | 60 | * | 56 | off | 570 |  | 2 | Y |  | 500 | 1060 | 100 | [5 empty] | Y
2016 | 3 | 23 | 83 | 1544 | MB/JJ | 60 | * | 55 | off | 560 |  | 2 | Y |  | 500 | 1060 | 100 | [5 empty] | Y
2016 | 3 | 24 | 84 | 1545 | MB/JJ | 60 | * | 56 | off | 570 |  | 2 | Y |  | 500 | 1060 | 100 | [5 empty] | Y
2016 | 3 | 25 | 85 | 1546 | JJ | 60 | * | 56 | off | 560 |  | 2 | y |  | 500 | 1050 | 95 | [5 empty] | Y
2016 | 3 | 26 | 86 | 1547 | MB | 60 | * | 56 | off | 550 |  | 2 | Y |  | 500 | 1050 | 95 | [5 empty] | Y
2016 | 3 | 27 | 87 | 1548 | JJ | 60 | * | 55 | off | 550 |  | 2 | Y |  | 500 | 1060 | 95 | [5 empty] | Y
2016 | 3 | 28 | 88 | 1549 | MB | 60 | * | 55 | off | 540 |  | 2 | Y |  | 500 | 1050 | 100 | [5 empty] | Y | * | *
2016 | 3 | 29 | 89 | 1550 | JJ | 60 | * | 56 | off | 540 |  | 2 | Y |  | 500 | 1050 | 85 | [5 empty] | Y
2016 | 3 | 30 | 90 | 1551 | MB | 60 | * | 55 | off | 540 |  | 2 | Y |  | 500 | 1050 | 100 | [5 empty] | Y
2016 | 3 | 31 | 91 | 1552 | JJ/MB | 60 | * | 55 | off | 540 |  | 2 | Y |  | 500 | 1050 | 95 | [5 empty] | Y
2016 | 4 | 1 | 92 | 1553 | JJ | 60 | * | 56 | off | 530 |  | 2 | Y |  | 500 | 1050 | 95 | [5 empty] | Y
2016 | 4 | 2 | 93 | 1554 | JJ/MB | 60 | * | 53 | off | 520 |  | 2 | Y |  | 500 | 1050 | 100 | [5 empty] | Y
2016 | 4 | 3 | 94 | 1555 | MB | 60 | * | 55 | off | 520 |  | 2 | Y |  | 500 | 1050 | 100 | [5 empty] | Y
2016 | 4 | 4 | 95 | 1556 | JJ/MB | 60 | * | 55 | off | 510 |  | 2 | Y |  | 500 | 1050 | 100 | [5 empty] | Y | * | *
2016 | 4 | 5 | 96 | 1557 | MB/JJ | 60 | * | 56 | off | 500 |  | 2 | Y |  | 500 | 1050 | 100 | [5 empty] | Y
2016 | 4 | 6 | 97 | 1558 | JJ | 60 | added 1L | 56 | off | 500 |  | 2 | Y |  | 500 | 1050 | 100 | [5 empty] | Y
2016 | 4 | 7 | 98 | 1559 | MB/JJ | 60 | * | 56 | off | 500 |  | 2 | Y |  | 500 | 1050 | 100 | [5 empty] | Y
2016 | 4 | 8 | 99 | 1560 | MB/JJ | 60 | * | 56 | off | 500 |  | 2 | Y |  | 500 | 1050 | 100 | [5 empty] | Y |  |  |  |  | SWD climbed the tower from 1715-1800z to measure 10m and 16m windbird alignment. SWD reattached turnbuckle to strength cable from CUGC exhaust line to tower.
2016 | 4 | 9 | 100 | 1561 | MB/JJ | 60 | * | 57 | off | 480 |  | 2 | Y |  | 500 | 1050 | 100 | [5 empty] | Y
2016 | 4 | 10 | 101 | 1562 | JJ | 60 | * | 56 | off | 480 |  | 2 | Y |  | 490 | 1050 | 100 | [5 empty] | Y
2016 | 4 | 11 | 102 | 1563 | MB/JJ | 60 | * | 56 | off | 470 |  | 2 | Y |  | 500 | 1050 | 100 | [5 empty] | Y | * | *
2016 | 4 | 12 | 103 | 1564 | MB/JJ | 60 | * | 55 | off | 470 |  | 2 | Y |  | 500 | 1050 | 100 | [5 empty] | Y
2016 | 4 | 13 | 104 | 1565 | MB/JJ | 60 | * | 57 | off | 460 |  | 2 | Y |  | 500 | 1040 | 100 | [5 empty] | Y
2016 | 4 | 14 | 105 | 1566 | MB/JJ | 60 | * | 57 | off | 460 |  | 2 | Y |  | 490 | 1040 | 100 | [5 empty] | Y |  |  |  |  | JJ climbed tower from 1356-1422z to work on surface ozone inlet and clean instruments.
2016 | 4 | 15 | 106 | 1567 | MB/JJ/SWD | 60 | * | 56 | off | 445 |  | 2 | Y |  | 480 | 1040 | 100 | [5 empty] | Y
2016 | 4 | 16 | 107 | 1568 | JJ/MB | 60 | * | 55 | off | 450 |  | 2 | Y |  | 480 | 1040 | 100 | [5 empty] | Y
2016 | 4 | 17 | 108 | 1569 | MB | 60 | * | 56 | off | 440 |  | 2 | Y |  | 480 | 1040 | 100 | [5 empty] | Y
2016 | 4 | 18 | 109 | 1570 | MB | 60 | * | 57 | off | 440 |  | 2 | Y |  | 480 | 1040 | 100 | [5 empty] | Y | * | *
2016 | 4 | 19 | 110 | 1571 | MB/JJ | 60 | * | 57 | off | 430 |  | 2 | Y |  | 500 | 1040 | 100 | [5 empty] | Y
2016 | 4 | 20 | 111 | 1572 | MB/JJ | 60 | * | 56 | off | 440 |  | 2 | Y |  | 480 | 1030 | 100 | [5 empty] | Y
2016 | 4 | 21 | 112 | 1573 | MB/JJ/SWD | 60 | * | 56 | off | 410 |  | 2 | Y |  | 480 | 1040 | 100 | [5 empty] | Y
2016 | 4 | 22 | 113 | 1574 | JJ/MB | 60 | * | 56 | off | 400 |  | 2 | Y |  | 480 | 1030 | 100 | [5 empty] | Y
2016 | 4 | 23 | 114 | 1575 | [23 empty] | White out conditions (Condition 1) did not travel to TAWO this day.
2016 | 4 | 24 | 115 | 1576 | JJ/MB | 60 | added 1L | 57 | off | 400 |  | 2 | Y |  | 480 | 1030 | 100 | [5 empty] | Y
2016 | 4 | 25 | 116 | 1577 | JJ/MB | 60 | * | 56 | off | 390 |  | 2 | Y |  | 480 | 1030 | 100 | [5 empty] | Y | * | *
2016 | 4 | 26 | 117 | 1578 | JJ | 60 | * | 57 | off | 390 |  | 2 | Y |  | 480 | 1030 | 100 | [5 empty] | Y
2016 | 4 | 27 | 118 | 1579 | JJ/MB | 60 | * | 74 | 1650 | 390 |  | 2 | Y |  | 490 | 1030 | 130 | [5 empty] | High H2 flow |  |  |  |  | Detlev working on the GC.
2016 | 4 | 28 | 119 | 1580 | MB | 60 | * | 55 | 1580 | 380 |  | 2 | Y |  | 480 | 1030 | 100 | [5 empty] | Y |  |  |  |  | Detlev working on the GC. Trap housing >250, reset to 100 ml/min. 1255-1305z JJ climbed the tower to clean instruments.
2016 | 4 | 29 | 120 | 1581 | MB/JJ | 60 | * | 55 | 1450 | 370 |  | 2 | Y |  | 480 | 1040 | 110 | [5 empty] | Y
2016 | 4 | 30 | 121 | 1582 | MB/JJ | 60 | * | 53 | 1530 | 370 |  | 2 | Y |  | 470 | 1050 | 110 | [5 empty] | Y |  |  |  |  | Detlev working with tanks.
2016 | 5 | 1 | 122 | 1583 | MB/JJ | 60 | * | 50 | 1620 | 370 |  | 2 | Y |  | 470 | 480 | 100 | [5 empty] | switched tanks
2016 | 5 | 2 | 123 | 1584 | MB/JJ | 60 | * | 50 | 1540 | 360 |  | 2 | Y |  | 450 | 480 | 105 | [5 empty] | Y | * | *
2016 | 5 | 3 | 124 | 1585 | MB/JJ | 60 | * | 65 | 1650 | 350 |  | 2 | Y |  | 490 | 640 | 110 | [5 empty] | High H2 flow; switched tanks |  |  |  |  | 1153-1217z JJ climbed tower to clean instruments and attach ties to CUGC intake hose.
2016 | 5 | 4 | 125 | 1586 | MB/JJ | 60 | * | 64 | 1650 | 350 |  | 2 | Y |  | 490 | 1040 | 100 | [5 empty] | Y |  |  |  |  | Adjusted trap housing purge to 100 ml/min.
2016 | 5 | 5 | 126 | 1587 | MB | 60 | * | 62 | 1630 | 350 |  | 2 | Y |  | 480 | 1040 | 105 | [5 empty] | Y |  |  |  |  | Adjusted trap housing purge to 100 ml/min.
2016 | 5 | 6 | 127 | 1588 | JJ/MB | 60 | * | 65 | 1620 | 350 |  | 2 | Y |  | 480 | 1040 | 110 | [5 empty] | Y |  |  |  |  | Adjusted trap housing purge to 100 ml/min.
2016 | 5 | 7 | 128 | 1589 | JJ/MB | 60 | * | 65 | 1620 | 350 |  | 2 | Y |  | 480 | 1030 | 90 | [5 empty] | Y |  |  |  |  | Adjusted trap housing purge to 100 ml/min.
2016 | 5 | 8 | 129 | 1590 | JJ | 60 | * | 62 | 1620 | 350 |  | 2 | Y |  | 470 | 1040 | 90 | [5 empty] | Y |  |  |  |  | Adjusted trap housing purge to 100 ml/min.
2016 | 5 | 9 | 130 | 1591 | JJ/MB | 60 | * | 65 | 1640 | 350 |  | 2 | Y |  | 470 | 1050 | 100 | [5 empty] | Y | * | *
2016 | 5 | 10 | 131 | 1592 | JJ/MB | 60 | * | 66 | 1640 | 350 |  | 2 | Y |  | 470 | 1030 | 100 | [5 empty] | Y
2016 | 5 | 11 | 132 | 1593 | MB/JJ | 60 | * | 65 | 1610 | 340 |  | 2 | Y |  | 470 | 1030 | 100 | [5 empty] | Y
2016 | 5 | 12 | 133 | 1594 | MB/JJ | 60 | added 1L | 65 | 1620 | 340 |  | 2 | Y |  | 480 | 1030 | 100 | [5 empty] | Y
2016 | 5 | 13 | 134 | 1595 | MB/JJ | 60 | * | 64 | 1620 | 340 |  | 2 | Y |  | 480 | 1030 | 100 | [5 empty] | Y
2016 | 5 | 14 | 135 | 1596 | JJ/MB | 60 | * | 64 | 1620 | 340 |  | 2 | Y |  | 480 | 1030 | 100 | [5 empty] | Y
2016 | 5 | 15 | 136 | 1597 | MB | 60 | * | 65 | 1620 | 340 |  | 2 | Y |  | 480 | 1030 | 100 | [5 empty] | Y
2016 | 5 | 16 | 137 | 1598 | JJ/MB | 60 | * | 65 | 1620 | 340 |  | 2 | Y |  | 480 | 1030 | 100 | [5 empty] | Y | * | *
2016 | 5 | 17 | 138 | 1599 | JJ/MB | 60 | * | 66 | 1610 | 330 |  | 2 | Y |  | 470 | 1030 | 100 | [5 empty] | Y
2016 | 5 | 18 | 139 | 1600 | JJ/MB | 60 | * | 64 | 1620 | 330 |  | 2 | Y |  | 470 | 1030 | 100 | [5 empty] | Y
2016 | 5 | 19 | 140 | 1601 | JJ | 60 | * | 64 | 1610 | 330 |  | 2 | Y |  | 470 | 1030 | 95 | [5 empty] | Y |  |  |  |  | Adjusted trap housing purge to 100 ml/min
2016 | 5 | 20 | 141 | 1602 | JJ/MB | 60 | * | 65 | 1620 | 330 |  | 2 | Y |  | 470 | 1030 | 95 | [5 empty] | Y
2016 | 5 | 21 | 142 | 1603 | JJ/MB | 60 | * | 65 | 1610 | 330 |  | 2 | Y |  | 470 | 1030 | 95 | [5 empty] | Y
2016 | 5 | 22 | 143 | 1604 | JJ | 60 | * | 66 | 1620 | 330 |  | 2 | Y |  | 460 | 1030 | 95 | [5 empty] | Y
2016 | 5 | 23 | 144 | 1605 | JJ | 60 | * | 65 | 1620 | 330 |  | 2 | Y |  | 460 | 1030 | 95 | [5 empty] | Y | * | *
2016 | 5 | 24 | 145 | 1606 | JJ/MB | 60 | * | 65 | 1620 | 320 |  | 2 | Y |  | 470 | 1030 | 95 | [5 empty] | Y
2016 | 5 | 25 | 146 | 1607 | JJ | 60 | * | 62 | 1620 | 330 |  | 2 | Y |  | 450 | 1030 | 95 | [5 empty] | Y
2016 | 5 | 26 | 147 | 1608 | JJ/MB | 60 | * | 65 | 1610 | 320 |  | 2 | Y |  | 460 | 1030 | 90 | [5 empty] | Y |  |  |  |  | Adjusted trap housing purge to 100 ml/min
2016 | 5 | 27 | 148 | 1609 | JJ/MB | 60 | * | 65 | 1610 | 320 |  | 2 | Y |  | 460 | 1030 | 100 | [5 empty] | Y
2016 | 5 | 28 | 149 | 1610 | MB | 60 | added 1L | 65 | 1610 | 310 |  | 2 | Y |  | 460 | 1030 | 100 | [5 empty] | Y |  |  |  |  | Added 1L of DI water.
2016 | 5 | 29 | 150 | 1611 | MB | 60 | * | 65 | 1610 | 310 |  | 2 | Y |  | 460 | 1030 | 100 | [5 empty] | Y
2016 | 5 | 30 | 151 | 1612 | JJ | 60 | * | 65 | 1630 | 320 |  | 2 | Y |  | 460 | 1030 | 100 | [5 empty] | Y | * | * |  |  | Planned power outage today from 2130 to 2230
2016 | 5 | 31 | 152 | 1613 | JJ | 60 | * | 65 | 1620 | 320 |  | 2 | Y |  | 450 | 1030 | 100 | [5 empty] | Y
2016 | 6 | 1 | 153 | 1614 | JJ | 60 | * | 66 | 1620 | 320 |  | 2 | Y |  | 450 | 1020 | 105 | [5 empty] | Y
2016 | 6 | 2 | 154 | 1615 | MB/JJ | 60 | * | 65 | 1620 | 310 |  | 2 | Y |  | 460 | 1030 | 105 | [5 empty] | Y
2016 | 6 | 3 | 155 | 1616 | MB/JJ/HJ | 60 | * | 54 | 1660 | 310 |  | 2 | Y |  | 460 | 1030 | 105 | [5 empty] | Y |  |  |  |  | H2 line filter desiccant pale yellow. Turned off H2 generator and replaced with H2 backup tank. Switched valve to receive H2 tank at 2012z. H2 tank pressure 2100 psi.
2016 | 6 | 4 | 156 | 1617 | MB | offline | offline | offline | 1610 | 300 |  | 2 | Y |  | 460 | 1030 | 95 | [5 empty] | Y |  |  |  |  | H2 tank pressure 2060psi.
2016 | 6 | 5 | 157 | 1618 | JJ/TG/EB | offline | offline | offline | 1620 | 310 |  | 2 | Y |  | 450 | 1020 | 100 | [5 empty] | Y
2016 | 6 | 6 | 158 | 1619 | MB/EB/JJ/TG | offline | offline | offline | 1620 | 300 |  | 2 | Y |  | 450 | 1040 | 105 | [5 empty] | Y | * | * |  |  | H2 Line filter desiccant very pale/clear yellow; H2 tank pressure 2100 psi.
2016 | 6 | 7 | 159 | 1620 | TG/EB/JJ | offline | offline | offline | 1610 | 300 |  | 2 | Y |  | 450 | 1020 | 104 | [5 empty] | Y |  |  |  |  | H2 tank pressure 2060psi.
2016 | 6 | 8 | 160 | 1621 | TG/EB/JJ/MB | offline | offline | offline | 1620 | 300 |  | 2 | Y |  | 450 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 2000psi.
2016 | 6 | 9 | 161 | 1622 | TG/EB | offline | offline | offline | 1630 | 300 |  | 2 | Y |  | 420 | 1020 | 105 | [5 empty] | Y
2016 | 6 | 10 | 162 | 1623 | TG/EB | offline | offline | offline | 1630 | 300 |  | 2 | Y |  | 450 | 1030 | 105 | [5 empty] | Y
2016 | 6 | 11 | 163 | 1624 | TG/EB | offline | offline | offline | 1620 | 300 |  | 2 | Y |  | 440 | 1040 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1950psi.
2016 | 6 | 12 | 164 | 1625 | TG/EB | offline | offline | offline | 1630 | 300 |  | 2 | Y |  | 450 | 1040 | 105 | [5 empty] | Y
2016 | 6 | 13 | 165 | 1626 | TG/EB | offline | offline | offline | 1620 | 300 |  | 2 | Y |  | 450 | 1020 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1900psi
2016 | 6 | 14 | 166 | 1627 | TG/EB | offline | offline | offline | 1620 | 300 |  | 2 | Y |  | 450 | 1030 | 103 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1900psi
2016 | 6 | 15 | 167 | 1628 | TG/EB | offline | offline | offline | 1610 | 300 |  | 2 | Y |  | 440 | 1020 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1895psi
2016 | 6 | 16 | 168 | 1629 | TG | offline | offline | offline | 1610 | 295 |  | 2 | Y |  | 430 | 1010 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1870psi
2016 | 6 | 17 | 169 | 1630 | TG | offline | offline | offline | 1620 | 295 |  | 2 | Y |  | 420 | 1010 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1860psi
2016 | 6 | 18 | 170 | 1631 | TG/EB | offline | offline | offline | 1625 | 290 |  | 2 | Y |  | 450 | 1020 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1825psi
2016 | 6 | 19 | 171 | 1632 | EB | offline | offline | offline | 1620 | 300 |  | 2 | Y |  | 450 | 1020 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1795psi
2016 | 6 | 20 | 172 | 1633 | TG/EB | offline | offline | offline | 1625 | 290 |  | 2 | Y |  | 450 | 1010 | 105 | [5 empty] | Y | * | * |  |  | H2 tank pressure 1740
2016 | 6 | 21 | 173 | 1634 | TG/EB | offline | offline | offline | 1620 | 290 |  | 2 | Y |  | 440 | 1010 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1750psi.
2016 | 6 | 22 | 174 | 1635 | TG/EB | offline | offline | offline | 1620 | 300 |  | 2 | Y |  | 450 | 1020 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1700psi.
2016 | 6 | 23 | 175 | 1636 | EB | offline | offline | offline | 1520 | 290 |  | 2 | Y |  | 450 | 1025 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1680psi.
2016 | 6 | 24 | 176 | 1637 | TG | offline | offline | offline | 1620 | 280 |  | 2 | Y |  | 440 | 1010 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1640psi
2016 | 6 | 25 | 177 | 1638 | TG/EB | offline | offline | offline | 1640 | 290 |  | 2 | Y |  | 440 | 1010 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1650psi
2016 | 6 | 26 | 178 | 1639 | TG | offline | offline | offline | 1610 | 270 |  | 2 | Y |  | 440 | 1010 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1650psi
2016 | 6 | 27 | 179 | 1640 | EB/TG | offline | offline | offline | 1610 | 290 |  | 2 | Y |  | 440 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1610psi.
2016 | 6 | 28 | 180 | 1641 | TG/EB | offline | offline | offline | 1610 | 290 |  | 2 | Y |  | 450 | 1010 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1600 psi.
2016 | 6 | 29 | 181 | 1642 | TG/EB | offline | offline | offline | 1610 | 270 |  | 2 | Y |  | 420 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1590psi.
2016 | 6 | 30 | 182 | 1643 | TG/EB | offline | offline | offline | 1620 | 275 |  | 2 | Y |  | 450 | 1020 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1550psi
2016 | 7 | 1 | 183 | 1644 | TG/EB | offline | offline | offline | 1620 | 260 |  | 2 | Y |  | 430 | 1010 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1510psi
2016 | 7 | 2 | 184 | 1645 | TG/EB | offline | offline | offline | 1610 | 260 |  | 2 | Y |  | 420 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1380psi
2016 | 7 | 3 | 185 | 1646 | EB | offline | offline | offline | 1620 | 260 |  | 2 | Y |  | 450 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1410psi
2016 | 7 | 4 | 186 | 1647 | TG | offline | offline | offline | 1620 | 260 |  | 2 | y |  | 420 | 1010 | 105 | [5 empty] | y |  |  |  |  | H2 tank pressure 1410psi
2016 | 7 | 5 | 187 | 1648 | TG/EB | offline | offline | offline | 1620 | 260 |  | 2 | Y |  | 450 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1380psi
2016 | 7 | 6 | 188 | 1649 | TG | offline | offline | offline | 1620 | 260 |  | 2 | Y |  | 420 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1390psi
2016 | 7 | 7 | 189 | 1650 | TG/EB | offline | offline | offline | 1620 | 260 |  | 2 | Y |  | 420 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1360psi
2016 | 7 | 8 | 190 | 1651 | TG/EB | offline | offline | offline | 1620 | 260 |  | 2 | Y |  | 450 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1290psi
2016 | 7 | 9 | 191 | 1652 | TG | offline | offline | offline | 1620 | 260 |  | 2 | Y |  | 420 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1260psi
2016 | 7 | 10 | 192 | 1653 | TG | offline | offline | offline | 1620 | 250 |  | 2 | Y |  | 420 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1250psi
2016 | 7 | 11 | 193 | 1654 | EB | offline | offline | offline | 1620 | 250 |  | 2 | Y |  | 450 | 1010 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1210psi
2016 | 7 | 12 | 194 | 1655 | TG/EB | offline | offline | offline | 1620 | 250 |  | 2 | Y |  | 420 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1220psi
2016 | 7 | 13 | 195 | 1656 | TG | offline | offline | offline | 1620 | 250 |  | 3 | Y |  | 420 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1200psi
2016 | 7 | 14 | 196 | 1657 | TG/EB | offline | offline | offline | 1610 | 250 |  | 2 | Y |  | 440 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1160psi
2016 | 7 | 15 | 197 | 1658 | TG/EB | offline | offline | offline | 1620 | 250 |  | 2 | Y |  | 410 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1150psi
2016 | 7 | 16 | 198 | 1659 | TG/EB | offline | offline | offline | 1610 | 245 |  | 2 | Y |  | 440 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1110psi
2016 | 7 | 17 | 199 | 1660 | EB | offline | offline | offline | 1610 | 240 |  | 2 | Y |  | 440 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1070psi
2016 | 7 | 18 | 200 | 1661 | TG | offline | offline | offline | 1610 | 240 |  | 2 | Y |  | 410 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1060psi
2016 | 7 | 19 | 201 | 1662 | TG/EB | offline | offline | offline | 1620 | 240 |  | 2 | Y |  | 420 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1010 psi
2016 | 7 | 20 | 202 | 1663 | TG/EB | offline | offline | offline | 1610 | 240 |  | 2 | Y |  | 420 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1000 psi
2016 | 7 | 21 | 203 | 1664 | TG | offline | offline | offline | 1620 | 230 |  | 2 | Y |  | 420 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 1000 psi
2016 | 7 | 22 | 204 | 1665 | TG | offline | offline | offline | 1610 | 230 |  | 2 | Y |  | 420 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 995 psi
2016 | 7 | 23 | 205 | 1666 | TG | offline | offline | offline | 1620 | 230 |  | 2 | Y |  | 420 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 990 psi
2016 | 7 | 24 | 206 | 1667 | TG/EB | offline | offline | offline | 1620 | 220 |  | 2 | Y |  | 420 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 920psi
2016 | 7 | 25 | 207 | 1668 | TG/EB | offline | offline | offline | 1610 | 220 |  | 2 | Y |  | 420 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 900psi
2016 | 7 | 26 | 208 | 1669 | TG/EB | offline | offline | offline | 1610 | 220 |  | 2 | Y |  | 420 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 900psi
2016 | 7 | 27 | 209 | 1670
2016 | 7 | 28 | 210 | 1671 | TG/EB | offline | offline | offline | 1610 | 220 |  | 2 | Y |  | 410 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 850psi
2016 | 7 | 29 | 211 | 1672 | TG/EB | offline | offline | offline | 1610 | 220 |  | 2 | Y |  | 420 | 1000 | 105 | [5 empty] | Y
2016 | 7 | 30 | 212 | 1673 | TG/EB | offline | offline | offline | 1620 | 210 |  | 2 | Y |  | 410 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 800psi
2016 | 7 | 31 | 213 | 1674 | EB | offline | offline | offline | 1610 | 210 |  | 2 | Y |  | 420 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 790psi
2016 | 8 | 1 | 214 | 1675 | EB | offline | offline | offline | 1610 | 220 |  | 2 | Y |  | 410 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 760psi
2016 | 8 | 2 | 215 | 1676 | EB | offline | offline | offline | 1610 | 210 |  | 2 | Y |  | 420 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 740psi
2016 | 8 | 3 | 216 | 1677 | TG | offline | offline | offline | 1610 | 200 |  | 2 | Y |  | 410 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 710psi
2016 | 8 | 4 | 217 | 1678 | EB | offline | offline | offline | 1610 | 200 |  | 2 | Y |  | 410 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 680psi
2016 | 8 | 5 | 218 | 1679 | TG/EB | offline | offline | offline | 1610 | 200 |  | 2 | Y |  | 410 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 670psi
2016 | 8 | 6 | 219 | 1680 | EB | offline | offline | offline | 1610 | 200 |  | 2 | Y |  | 410 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2 tank pressure 610psi
2016 | 8 | 7 | 220 | 1681 | TG | offline | offline | offline | 1620 | 200 |  | 2 | Y |  | 410 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2: 600psi
2016 | 8 | 8 | 221 | 1682 | TG | offline | offline | offline | 1620 | 200 |  | 2 | Y |  | 400 | 1000 | 100 | [5 empty] | Y |  |  |  |  | H2: 580psi
2016 | 8 | 9 | 222 | 1683 | TG | offline | offline | offline | 1620 | 190 |  | 2 | Y |  | 410 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2: 560psi
2016 | 8 | 10 | 223 | 1684 | EB | offline | offline | offline | 1620 | 200 |  | 2 | Y |  | 410 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2: 540psi
2016 | 8 | 11 | 224 | 1685 | EB | offline | offline | offline | 1600 | 200 |  | 2 | Y |  | 405 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2: 510psi
2016 | 8 | 12 | 225 | 1686 | TG | offline | offline | offline | 1610 | 190 |  | 2 | Y |  | 410 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2: 500psi Power outage 1609-1611z. 1642-1644z
2016 | 8 | 13 | 226 | 1687 | EB | offline | offline | offline | 1600 | 200 |  | 2 | Y |  | 410 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2: 480psi
2016 | 8 | 14 | 227 | 1688 | EB | offline | offline | offline | 1600 | 200 |  | 2 | Y |  | 410 | 1000 | 105 | [5 empty] | Y |  |  |  |  | H2:450psi
2016 | 8 | 15 | 228 | 1689 | TG/EB | 60 | * | 47 | 1850 | 190 |  | 2 | Y |  | 410 | 990 | 105 | [5 empty] | Y |  |  |  |  | H2:450psi
2016 | 8 | 16 | 229 | 1690 | TG | 60 | * | 45 | 400 | 1800 |  | 2 | Y |  | 400 | 990 | 105 | [5 empty] | Y | * | * | * |  | Looks like Jacque switched CH4 std 1 tank from the large tank to the small yellow tank that has much less pressure
2016 | 8 | 17 | 230 | 1691 | TG | 60 | * | 43 | 400 | 1800 |  | 2 | y |  | 400 | 990 | 105 | [5 empty] | Y | * | *
2016 | 8 | 18 | 231 | 1692 | TG | 60 | * | 43 | 400 | 1800 |  | 2 | y |  | 400 | 990 | 105 | [5 empty] | Y | * | * | *
2016 | 8 | 19 | 232 | 1693 | EB/TG | 60 | * | 37 | 400 | 1800 |  | 2 | Y |  | 400 | 990 | 105 | [5 empty] | Y | * | * | *
2016 | 8 | 20 | 233 | 1694 | TG | 60 | * | 45 | 400 | 1790 |  | 2 | Y |  | 400 | 990 | 105 | [5 empty] | Y | * | * | *
2016 | 8 | 21 | 234 | 1695 | EB | 60 | * | 47 | 400 | 1790 |  | 2 | Y |  | 400 | 990 | 95 | [5 empty] | Y | * | * | *
2016 | 8 | 22 | 235 | 1696 | TG | 60 | * | 45 | 400 | 1790 |  | 2 | Y |  | 400 | 990 | 105 | [5 empty] | Y | * | * | *
2016 | 8 | 23 | 236 | 1697 | EB | 60 | * | 44 | 400 | 1780 |  | 2 | Y |  | 410 | 1000 | 105 | [5 empty] | Y | * | * | *
2016 | 8 | 24 | 237 | 1698 | TG | 60 | * | 46 | 400 | 1790 |  | 2 | Y |  | 400 | 990 | 105 | [5 empty] | Y | * | * | *
2016 | 8 | 25 | 238 | 1699 | EB | 60 | * | 45 | 400 | 1780 |  | 2 | Y |  | 410 | 980 | 105 | [5 empty] | Y | * | * | *
2016 | 8 | 26 | 239 | 1700 | TG | 60 | * | 43 | 400 | 1790 |  | 2 |  |  | 410 | 970 | 105 | [5 empty] | y | * | * | *
2016 | 8 | 27 | 240 | 1701 | EB | 60 | * | 46 | 400 | 1790 |  | 2 | Y |  | 400 | 960 | 105 | [5 empty] | Y | * | * | *
2016 | 8 | 28 | 241 | 1702 | TG | 60 | * | 44 | 390 | 1790 |  | 2 | Y |  | 400 | 960 | 105 | [5 empty] | Y | * | * | * |  | Is NMHC dropping faster than it should? Didn't see anything with a snoop test
2016 | 8 | 29 | 242 | 1703 | EB | 60 | * | 41 | 400 | 1770 |  | 2 | Y |  | 400 | 960 | 90 | [5 empty] | Y | * | * | *
2016 | 8 | 30 | 243 | 1704 | TG | 60 | * | 45 | 400 | 1780 |  | 2 | Y |  | 400 | 960 | 105 | [5 empty] | Y | * | * | *
2016 | 8 | 31 | 244 | 1705 | EB | 60 | * | 41 | 400 | 1780 |  | 2 | Y |  | 400 | 955 | 105 | [5 empty] | Y | * | * | *
2016 | 9 | 1 | 245 | 1706 | TG | 60 | * | 45 | 400 | 1780 |  | 2 | Y |  | 400 | 960 | 105 | [5 empty] | Y | * | * | *
2016 | 9 | 2 | 246 | 1707 | EB | 60 | * | 43 | 400 | 1760 |  | 2 | Y |  | 400 | 950 | 100 | [5 empty] | Y | * | * | *
2016 | 9 | 3 | 247 | 1708 | TG | 60 | * | 43 | 400 | 1760 |  | 2 | Y |  | 390 | 950 | 100 | [5 empty] | Y | * | * | *
2016 | 9 | 4 | 248 | 1709 | TG | 60 | * | 45 | 400 | 1780 |  | 2 | Y |  | 390 | 950 | 105 | [5 empty] | Y | * | * | *
2016 | 9 | 5 | 249 | 1710 | TG | 60 | * | 45 | 400 | 1780 |  | 2 | Y |  | 400 | 950 | 105 | [5 empty] | y | * | * | *
2016 | 9 | 6 | 250 | 1711
2016 | 9 | 7 | 251 | 1712 | TG | 60 | * | 45 | 400 | 1780 |  | 2 | Y |  | 400 | 950 | 105 | [5 empty] | y | * | * | *
2016 | 9 | 8 | 252 | 1713
2016 | 9 | 9 | 253 | 1714 | TG | 60 | * | 45 | 400 | 177 |  | 2 | Y |  | 400 | 940 | 105 | [5 empty] | Y | * | * | *
2016 | 9 | 10 | 254 | 1715
2016 | 9 | 11 | 255 | 1716 | TG | 60 | * | 48 | 400 | 1770 |  | 2 | Y |  | 400 | 950 | 105 | [5 empty] | Y | * | * | *
2016 | 9 | 12 | 256 | 1717
2016 | 9 | 13 | 257 | 1718 | TG | 60 | * | 45 | 400 | 1770 |  | 2 | Y |  | 400 | 950 | 105 | [5 empty] | Y | * | * | *
2016 | 9 | 14 | 258 | 1719
2016 | 9 | 15 | 259 | 1720 | TG | 60 | * | 46 | 400 | 1760 |  | 2 | Y |  | 400 | 940 | 105 | [5 empty] | Y | * | * | *
2016 | 9 | 16 | 260 | 1721
2016 | 9 | 17 | 261 | 1722 | TG | 60 | * | 45 | 400 | 1760 |  | 2 | Y |  | 400 | 940 | 105 | [5 empty] | Y | * | * | *
2016 | 9 | 18 | 262 | 1723
2016 | 9 | 19 | 263 | 1724 | TG | 60 | * | 45 | 400 | 1760 |  | 2 | Y |  | 400 | 940 | 105 | [5 empty] | Y | * | * | *
2016 | 9 | 20 | 264 | 1725
2016 | 9 | 21 | 265 | 1726 | TG | 60 | * | 45 | 400 | 1750 |  | 2 | Y |  | 400 | 940 | 105 | [5 empty] | Y | * | * | *
2016 | 9 | 22 | 266 | 1727
2016 | 9 | 23 | 267 | 1728 | TG | 60 | * | 46 | 400 | 1750 |  | 2 | Y |  | 400 | 940 | 105 | [5 empty] | Y | * | * | *
2016 | 9 | 24 | 268 | 1729
2016 | 9 | 25 | 269 | 1730 | TG | 60 | * | 50 | 400 | 1740 |  | 2 | y |  | 400 | 940 | 105 | [5 empty] | Y | * | * | *
2016 | 9 | 26 | 270 | 1731
2016 | 9 | 27 | 271 | 1732 | TG | 60 | * | 45 | 400 | 1740 |  | 2 | Y |  | 400 | 940 | 105 | [5 empty] | Y | * | * | *
2016 | 9 | 28 | 272 | 1733
2016 | 9 | 29 | 273 | 1734 | TG | 60 | * | 44 | 400 | 1740 |  | 2 | Y |  | 400 | 940 | 105 | [5 empty] | Y | * | * | *
2016 | 9 | 30 | 274 | 1735
2016 | 10 | 1 | 275 | 1736 | TG | 60 | * | 43 | 400 | 1740 |  | 2 | Y |  | 400 | 940 | 105 | [5 empty] | Y | * | *
2016 | 10 | 2 | 276 | 1737
2016 | 10 | 3 | 277 | 1738 | TG | 60 | * | 44 | 400 | 1730 |  | 2 | Y |  | 400 | 940 | 105 | [5 empty] | Y | * | * | *
2016 | 10 | 4 | 278 | 1739
2016 | 10 | 5 | 279 | 1740 | TG | 60 | * | 44 | 400 | 1720 |  | 2 | Y |  | 400 | 940 | 105 | [5 empty] | Y | * | * | *
2016 | 10 | 6 | 280 | 1741
2016 | 10 | 7 | 281 | 1742 | TG | 60 | * | 42 | 400 | 1730 |  | 2 | Y |  | 400 | 930 | 105 | [5 empty] | Y | * | * | *
2016 | 10 | 8 | 282 | 1743
2016 | 10 | 9 | 283 | 1744 | TG | 60 | * | 44 | 400 | 1730 |  | 2 | Y |  | 400 | 930 | 105 | [5 empty] | Y | * | * | *
2016 | 10 | 10 | 284 | 1745
2016 | 10 | 11 | 285 | 1746 | TG | 60 | * | 44 | 400 | 1720 |  | 2 | y |  | 400 | 930 | 105 | [5 empty] | y | * | * | *
2016 | 10 | 12 | 286 | 1747
2016 | 10 | 13 | 287 | 1748 | TG | 60 | * | 43 | 400 | 1720 |  | 2 | Y |  | 400 | 930 | 105 | [5 empty] | Y | * | * | *
2016 | 10 | 14 | 288 | 1749
2016 | 10 | 15 | 289 | 1750 | TG | 60 | * | 44 | 400 | 1700 |  | 2 | Y |  | 400 | 930 | 105 | [5 empty] | Y | * | * | *
2016 | 10 | 16 | 290 | 1751
2016 | 10 | 17 | 291 | 1752 | TG | 60 | * | 43 | 400 | 1710 |  | 2 | Y |  | 400 | 920 | 105 | [5 empty] | Y | * | * | *
2016 | 10 | 18 | 292 | 1753
2016 | 10 | 19 | 293 | 1754 | TG | 60 | * | 43 | 400 | 1710 |  | 2 | Y |  | 390 | 940 | 405 | [5 empty] | Y | * | * | *
2016 | 10 | 20 | 294 | 1755
2016 | 10 | 21 | 295 | 1756 | TG | 60 | * | 35 | 400 | 1710 |  | 2 | Y |  | 390 | 930 | 105 | [5 empty] | Y | * | * | *
2016 | 10 | 22 | 296 | 1757
2016 | 10 | 23 | 297 | 1758 | TG | 60 | * | 44 | 400 | 1690 |  | 2 | Y |  | 400 | 930 | 105 | [5 empty] | Y | * | * | *
2016 | 10 | 24 | 298 | 1759
2016 | 10 | 25 | 299 | 1760 | MG/TG/NB | 60 | * | 43 | 400 | 1690 |  | 2 | Y |  | 400 | 930 | 95 | [5 empty] | Y | * | * | *
2016 | 10 | 26 | 300 | 1761 | MG/TG | 61 | * | 72 | 395 | 1690 |  | 2 | Y |  | 400 | 925 | 100 | [5 empty] | Y | * | * | *
2016 | 10 | 27 | 301 | 1762 | TG | 61 | * | 43 | 400 | 1680 |  | 2 | Y |  | 400 | 920 | 105 | [5 empty] | Y | * | * | * |  | Added 1 L of DI water to H2 Generator
2016 | 10 | 28 | 302 | 1763 | MG/NB | 61 | * | 35 | 400 | 1680 |  | 2 | Y |  | 400 | 920 | 100 | [5 empty] | Y | * | * | *
2016 | 10 | 29 | 303 | 1764 | MG/NB | 61 | * | 34 | 400 | 1680 |  | 2 | Y |  | 390 | 910 | 100 | [5 empty] | Y | * | * | *
2016 | 10 | 30 | 304 | 1765 | MG/NB | 61 | * | 35 | 390 | 1680 |  | 2 | Y |  | 390 | 910 | 105 | [5 empty] | Y
2016 | 10 | 31 | 305 | 1766 | NB | 61 | * | 44 | 390 | 1680 |  | 2 | Y |  | 390 | 910 | 100 | [5 empty] | Y | * | * | *
2016 | 11 | 1 | 306 | 1767 | NB | 61 | * | 44 | 390 | 1680 |  | 2 | Y |  | 390 | 910 | 105 | [5 empty] | Y | * | * | * |  | MG climbed tower 1455-1520z and removed frost from sample inlet.
2016 | 11 | 2 | 307 | 1768 | MG | 61 | * | 44 | 390 | 1670 |  | 2 | Y |  | 390 | 900 | 105 | [5 empty] | y
2016 | 11 | 3 | 308 | 1769 | NB | 61 | * | 44 | 390 | 1670 |  | 2 | Y |  | 390 | 910 | 105 | [5 empty] | Y | * | * | *
2016 | 11 | 4 | 309 | 1770 | MG | 61 | * | 44 | 390 | 1670 |  | 2 | Y |  | 380 | 900 | 105 | [5 empty] | Y | * | * | *
2016 | 11 | 5 | 310 | 1771 | MG | 61 | * | 64 | 390 | 1670 |  | 2 |  |  | 390 | 910 | 105 | [5 empty] | Y | * | * | *
2016 | 11 | 6 | 311 | 1772 | MG | 61 | * | 44 | 390 | 1670 |  | 2 | Y |  | 390 | 910 | 105 | [5 empty] | Y | * | * | *
2016 | 11 | 7 | 312 | 1773 | NB | 61 | * | 41 | 390 | 1660 |  | 2 | Y |  | 390 | 910 | 105 | [5 empty] | Y | * | * | *
2016 | 11 | 8 | 313 | 1774 | MG | 61 | * | 43 | 390 | 1660 |  | 2 | Y |  | 390 | 910 | 100 | [5 empty] | Y | * | * | *
2016 | 11 | 9 | 314 | 1775 | MG | 61 | * | 41 | 390 | 1660 |  | 2 | Y |  | 390 | 910 | 100 | [5 empty] | Y | * | * | *
2016 | 11 | 10 | 315 | 1776 | MG | 61 | * | 44 | 390 | 1660 |  | 2 | Y |  | 390 | 910 | 100 | [5 empty] | Y | * | * | *
2016 | 11 | 11 | 316 | 1777 | MG | 61 | * | 44 | 390 | 1660 |  | 2 | Y |  | 390 | 910 | 100 | [5 empty] | Y | * | * | *
2016 | 11 | 12 | 317 | 1778 | [23 empty] | Con 1 did not go to TAWO
2016 | 11 | 13 | 318 | 1779 | MG | 61 | * | 44 | 390 | 1570 |  | 2 | Y |  | 350 | 870 | 100 | [5 empty] | Y | * | * | * |  | Cold in TAWO today do to strong east wind (42F). Tank pressures low as a result.
2016 | 11 | 14 | 319 | 1780 | MG | 61 | * | 40 | 390 | 1660 |  | 2 | Y |  | 390 | 910 | 105 | [5 empty] | Y | * | * | *
2016 | 11 | 15 | 320 | 1781 | NB | 61 | * | 43 | 390 | 1660 |  | 2 | Y |  | 380 | 910 | 100 | [5 empty] | T | * | * | *
2016 | 11 | 16 | 321 | 1782 | MG | 61 | * | 41 | 390 | 1650 |  | 2 | Y |  | 380 | 910 | 100 | [5 empty] | Y | * | * | *
2016 | 11 | 17 | 322 | 1783 | MG/NB | 61 | * | 35 | 390 | 1670 |  | 2 | Y |  | 390 | 900 | 105 | [5 empty] | Y | * | * | * |  | Added 1 L of DI water to H2 Generator
2016 | 11 | 18 | 323 | 1784 | MG | 60 | * | 42 | 390 | 1640 |  | 2 | Y |  | 390 | 900 | 105 | [5 empty] | Y | * | * | *
2016 | 11 | 19 | 324 | 1785 | NB | 61 | * | 44 | 380 | 1620 |  | 2 | Y |  | 390 | 900 | 105 | [5 empty] | Y | * | * | * |  | MG climbed tower 1252 - 1336z and removed frost from sample inlet.
2016 | 11 | 20 | 325 | 1786 | NB | 61 | * | 43 | 390 | 1660 |  | 2 | Y |  | 390 | 900 | 105 | [5 empty] | Y | * | * | *
2016 | 11 | 21 | 326 | 1787 | NB | 61 | * | 43 | 390 | 1630 |  | 2 | Y |  | 380 | 900 | 105 | [5 empty] | Y | * | * | *
2016 | 11 | 22 | 327 | 1788 | MG | 61 | * | 44 | 390 | 1630 |  | 2 | Y |  | 380 | 900 | 105 | [5 empty] | Y | * | * | *
2016 | 11 | 23 | 328 | 1789 | NB | 61 | * | 44 | 390 | 1620 |  | 2 | Y |  | 380 | 900 | 105 | [5 empty] | Y | * | * | *
2016 | 11 | 24 | 329 | 1790 | NB | 61 | * | 43 | 390 | 1620 |  | 2 | Y |  | 380 | 890 | 100 | [5 empty] | Y | * | * | *
2016 | 11 | 25 | 330 | 1791 | MG | 61 | * | 44 | 390 | 1620 |  | 2 | Y |  | 380 | 900 | 105 | [5 empty] | Y | * | * | *
2016 | 11 | 26 | 331 | 1792 | NB | 61 | * | 43 | 390 | 1620 |  | 2 | Y |  | 380 | 900 | 105 | [5 empty] | Y | * | * | *
2016 | 11 | 27 | 332 | 1793 | MG | 61 | * | 43 | 390 | 1620 |  | 2 | Y |  | 380 | 900 | 105 | [5 empty] | Y | * | * | *
2016 | 11 | 28 | 333 | 1794 | NB | 61 | * | 44 | 390 | 1630 |  | 2 | Y |  | 380 | 900 | 105 | [5 empty] | Y | * | * | *
2016 | 11 | 29 | 334 | 1795 | NB | 61 | * | 43 | 390 | 1620 |  | 2 | Y |  | 380 | 900 | 105 | [5 empty] | Y | * | * | *
2016 | 11 | 30 | 335 | 1796 | MG | 61 | * | 48 | 390 | 1620 |  | 2 | Y |  | 370 | 900 | 105 | [5 empty] | Y | * | * | *
2016 | 12 | 1 | 336 | 1797 | NB | 61 | * | 44 | 390 | 1620 |  | 2 | Y |  | 370 | 900 | 105 | [5 empty] | Y | * | * | * |  | MG climbed tower 1302-1329z and removed frost from sample inlet.
2016 | 12 | 2 | 337 | 1798 | MG | 61 | * | 44 | 390 | 1620 |  | 2 | Y |  | 360 | 890 | 105 | [5 empty] | Y | * | * | *
2016 | 12 | 3 | 338 | 1799 | NB | 61 | * | 43 | 390 | 1620 |  | 2 | Y |  | 360 | 890 | 105 | [5 empty] | Y | * | * | *
2016 | 12 | 4 | 339 | 1800 | NB | 61 | * | 40 | 390 | 1620 |  | 2 | Y |  | 360 | 900 | 105 | [5 empty] | Y | * | * | *
2016 | 12 | 5 | 340 | 1801 | NB | 61 | * | 44 | 390 | 1620 |  | 2 | Y |  | 360 | 900 | 105 | [5 empty] | Y | * | * | *
2016 | 12 | 6 | 341 | 1802 | MG | 61 | * | 44 | 390 | 1620 |  | 2 | Y |  | 360 | 890 | 105 | [5 empty] | Y | * | * | * |  | NB climbed tower 1251-1320z and removed frost from sample inlet.
2016 | 12 | 7 | 342 | 1803 | MG | 61 | * | 43 | 390 | 1620 |  | 2 | Y |  | 360 | 890 | 105 | [5 empty] | Y | * | * | *
2016 | 12 | 8 | 343 | 1804 | MG/NB | 61 | * | 44 | 390 | 1610 |  | 2 | Y |  | 360 | 900 | 105 | [5 empty] | Y | * | * | * |  | Added 1 L of DI water to H2 Generator
2016 | 12 | 9 | 344 | 1805 | MG | 61 | * | 44 | 390 | 1610 |  | 2 | Y |  | 360 | 890 | 105 | [5 empty] | Y | * | * | *
2016 | 12 | 10 | 345 | 1806 | MG | 61 | * | 44 | 390 | 1610 |  | 2 | Y |  | 360 | 890 | 105 | [5 empty] | Y | * | * | *
2016 | 12 | 11 | 346 | 1807 | NB | 61 | * | 43 | 390 | 1610 |  | 2 | Y |  | 360 | 890 | 105 | [5 empty] | Y | * | * | *
2016 | 12 | 12 | 347 | 1808 | NB | 61 | * | 43 | 390 | 1600 |  | 2 | Y |  | 360 | 890 | 105 | [5 empty] | Y | * | * | *
2016 | 12 | 13 | 348 | 1809 | NB | 61 | * | 43 | 390 | 1600 |  | 2 | Y |  | 360 | 890 | 105 | [5 empty] | Y | * | * | *
2016 | 12 | 14 | 349 | 1810 | MG | 61 | * | 43 | 390 | 1590 |  | 2 | Y |  | 350 | 880 | 105 | [5 empty] | Y | * | * | *
2016 | 12 | 15 | 350 | 1811 | MG | 61 | * | 43 | 390 | 1600 |  | 2 | Y |  | 350 | 890 | 105 | [5 empty] | Y | * | * | *
2016 | 12 | 16 | 351 | 1812 | MG | 61 | * | 41 | 390 | 1590 |  | 2 | Y |  | 350 | 880 | 105 | [5 empty] | Y | * | * | *
2016 | 12 | 17 | 352 | 1813 | MG | 61 | * | 43 | 390 | 1590 |  | 2 | Y |  | 350 | 890 | 105 | [5 empty] | Y | * | * | *
2016 | 12 | 18 | 353 | 1814 | NB | 61 | * | 41 | 390 | 1580 |  | 2 | Y |  | 360 | 890 | 105 | [5 empty] | Y | * | * | *
2016 | 12 | 19 | 354 | 1815 | NB | 61 | * | 42 | 390 | 1580 |  | 2 | Y |  | 360 | 880 | 105 | [5 empty] | Y | * | * | *
2016 | 12 | 20 | 355 | 1816 | MG | 61 | * | 43 | 390 | 1570 |  | 2 | Y |  | 350 | 880 | 100 | [5 empty] | Y | * | * | *
2016 | 12 | 21 | 356 | 1817 | MG | 61 | * | 41 | 390 | 1570 |  | 2 | Y |  | 350 | 880 | 100 | [5 empty] | Y | * | * | *
2016 | 12 | 22 | 357 | 1818 | MG | 61 | * | 42 | 390 | 1540 |  | 2 | Y |  | 350 | 880 | 100 | [5 empty] | Y | * | * | *
2016 | 12 | 23 | 358 | 1819 | NB | 61 | * | 43 | 390 | 1570 |  | 2 | Y |  | 360 | 880 | 100 | [5 empty] | Y | * | * | *
2016 | 12 | 24 | 359 | 1820 | NB | 61 | * | 43 | 390 | 1570 |  | 2 | Y |  | 360 | 880 | 100 | [5 empty] | Y | * | * | *
2016 | 12 | 25 | 360 | 1821 | MG | 61 | * | 43 | 390 | 1560 |  | 2 | Y |  | 360 | 880 | 100 | [5 empty] | Y | * | * | *
2016 | 12 | 26 | 361 | 1822 | NB | 61 | * | 43 | 390 | 1560 |  | 2 | Y |  | 360 | 880 | 100 | [5 empty] | Y | * | * | * |  | Added 1 L of DI water to H2 Generator
2016 | 12 | 27 | 362 | 1823 | MG | 61 | * | 44 | 320 | 1510 |  | 2 | Y |  | 360 | 850 | 100 | [5 empty] | Y | * | * | *
2016 | 12 | 28 | 363 | 1824 | MG | 61 | * | 41 | 390 | 1560 |  | 2 | Y |  | 350 | 860 | 90 | [5 empty] | Y | * | * | *
2016 | 12 | 29 | 364 | 1825 | MG | 61 | * | 43 | 390 | 1560 |  | 2 | Y |  | 350 | 880 | 95 | [5 empty] | Y | * | * | *
2016 | 12 | 30 | 365 | 1826 | MG/NB | 61 | * | 41 | 390 | 1570 |  | 2 | Y |  | 350 | 880 | 105 | [5 empty] | Y | * | * | *
2016 | 12 | 31 | 366 | 1827 | NB | 61 | * | 43 | 390 | 1560 |  | 2 | Y |  | 350 | 870 | 105 | [5 empty] | Y | * | * | *
2017 | 1 | 1 | 1 | 1828 | NB | 61 | * | 42 | 390 | 1560 |  | 2 | Y |  | 350 | 870 | 105 | [5 empty] | Y | * | * | * |  | Happy New Year!
2017 | 1 | 2 | 2 | 1829 | NB | 61 | * | 39 | 390 | 1550 |  | 2 | Y |  | 350 | 860 | 105 | [5 empty] | Y | * | * | *
2017 | 1 | 3 | 3 | 1830 | NB | 61 | * | 42 | 390 | 1560 |  | 2 | Y |  | 350 | 870 | 105 | [5 empty] | Y | * | * | *
2017 | 1 | 4 | 4 | 1831 | NB | 61 | * | 43 | 390 | 1560 |  | 2 | Y |  | 350 | 870 | 105 | [5 empty] | Y | * | * | *
2017 | 1 | 5 | 5 | 1832 | MG | 61 | * | 43 | 390 | 1560 |  | 2 | Y |  | 250 | 870 | 100 | [5 empty] | Y | * | * | *
2017 | 1 | 6 | 6 | 1833 | NB | 61 | * | 43 | 390 | 1560 |  | 2 | Y |  | 350 | 870 | 100 | [5 empty] | Y | * | * | *
2017 | 1 | 7 | 7 | 1834 | MG | 61 | * | 43 | 390 | 1560 |  | 2 | Y |  | 350 | 870 | 100 | [5 empty] | Y | * | * | *
2017 | 1 | 8 | 8 | 1835 | MG | 61 | * | 43 | 390 | 1550 |  | 2 | Y |  | 350 | 870 | 100 | [5 empty] | Y | * | * | *
2017 | 1 | 9 | 9 | 1836 | NB | 61 | * | 43 | 390 | 1550 |  | 2 | Y |  | 340 | 870 | 100 | [5 empty] | Y | * | * | *
2017 | 1 | 10 | 10 | 1837 | MG | 61 | * | 43 | 390 | 1540 |  | 2 | Y |  | 340 | 860 | 100 | [5 empty] | Y | * | * | *
2017 | 1 | 11 | 11 | 1838 | MG | 61 | * | 42 | 390 | 1540 |  | 2 | Y |  | 340 | 860 | 100 | [5 empty] | Y | * | * | * |  | Added 1 liter to H2 gen
2017 | 1 | 12 | 12 | 1839 | MG | 61 | * | 41 | 390 | 1540 |  | 2 | Y |  | 340 | 860 | 100 | [5 empty] | Y | * | * | *
2017 | 1 | 13 | 13 | 1840 | MG | 61 | * | 43 | 390 | 1540 |  | 2 | Y |  | 340 | 860 | 100 | [5 empty] | Y | * | * | *
2017 | 1 | 14 | 14 | 1841 | MG | 61 | * | 43 | 390 | 1540 |  | 2 | Y |  | 340 | 860 | 100 | [5 empty] | Y | * | * | *
2017 | 1 | 15 | 15 | 1842 | NB | 61 | * | 42 | 390 | 1530 |  | 2 | Y |  | 340 | 860 | 100 | [5 empty] | Y | * | * | *
2017 | 1 | 16 | 16 | 1843 | MG | 61 | * | 43 | 390 | 1530 |  | 2 | Y |  | 340 | 850 | 100 | [5 empty] | Y | * | * | *
2017 | 1 | 17 | 17 | 1844 | MG | 61 | * | 41 | 390 | 1520 |  | 2 | Y |  | 340 | 850 | 100 | [5 empty] | Y | * | * | *
2017 | 1 | 18 | 18 | 1845 | MG | 61 | * | 43 | 390 | 1510 |  | 2 | Y |  | 320 | 840 | 100 | [5 empty] | Y | * | * | *
2017 | 1 | 19 | 19 | 1846 | MG | 61 | * | 44 | 390 | 1510 |  | 2 | Y |  | 320 | 840 | 100 | [5 empty] | Y | * | * | *
2017 | 1 | 20 | 20 | 1847 | NB | 61 | * | 43 | 390 | 1520 |  | 2 | Y |  | 340 | 860 | 100 | [5 empty] | Y | * | * | *
2017 | 1 | 21 | 21 | 1848 | MG | 61 | * | 43 | 390 | 1520 |  | 2 | Y |  | 340 | 860 | 100 | [5 empty] | Y | * | * | *
2017 | 1 | 22 | 22 | 1849 | MG | 61 | * | 43 | 390 | 1520 |  | 2 | Y |  | 340 | 860 | 100 | [5 empty] | Y | * | * | *
2017 | 1 | 23 | 23 | 1850 | NB | 61 | * | 43 | 390 | 1520 |  | 2 | Y |  | 330 | 860 | 100 | [5 empty] | Y | * | * | *
2017 | 1 | 24 | 24 | 1851 | MG | 61 | * | 43 | 390 | 1520 |  | 2 | Y |  | 330 | 860 | 100 | [5 empty] | Y | * | * | *
2017 | 1 | 25 | 25 | 1852 | NB | 61 | * | 43 | 390 | 1520 |  | 2 | Y |  | 330 | 860 | 100 | [5 empty] | Y | * | * | *
2017 | 1 | 26 | 26 | 1853 | MG | 61 | * | 43 | 390 | 1520 |  | 2 | Y |  | 320 | 850 | 100 | [5 empty] | Y | * | * | *
2017 | 1 | 27 | 27 | 1854 | MG | 61 | * | 43 | 390 | 1510 |  | 2 | Y |  | 330 | 850 | 100 | [5 empty] | Y | * | * | *
2017 | 1 | 28 | 28 | 1855 | MG | 61 | * | 35 | 390 | 1510 |  | 2 | Y |  | 330 | 860 | 100 | [5 empty] | Y | * | * | *
2017 | 1 | 29 | 29 | 1856 | NB | 61 | * | 36 | 390 | 1510 |  | 2 | Y |  | 330 | 860 | 100 | [5 empty] | Y | * | * | *
2017 | 1 | 30 | 30 | 1857 | NB | 61 | added 1 liter | 43 | 390 | 1510 |  | 2 | Y |  | 330 | 860 | 100 | [5 empty] | Y | * | * | * |  | MG climbed tower 1320-1408z and removed frost from sample inlet. Discovered a loose hose clamp near the end of the inlet, but did not have the tools to tighten. Will bring tools and tighten when tower is next climbed.
2017 | 1 | 31 | 31 | 1858 | MG | 61 | * | 43 | 390 | 1510 |  | 2 | Y |  | 330 | 860 | 100 | [5 empty] | Y | * | * | *
2017 | 2 | 1 | 32 | 1859 | MG | 61 | * | 43 | 390 | 1510 |  | 2 | Y |  | 330 | 860 | 100 | [5 empty] | Y | * | * | * |  | NB climbed the tower 1655-1732 UTC and removed frost from the inlet and fixed the loose clamp
2017 | 2 | 2 | 33 | 1860 | MG | 61 | * | 44 | 390 | 1510 |  | 2 | Y |  | 330 | 860 | 100 | [5 empty] | Y | * | * | *
2017 | 2 | 3 | 34 | 1861 | MG | 61 | * | 43 | 360 | 1510 |  | 2 | Y |  | 330 | 860 | 100 | [5 empty] | Y | * | * | *
2017 | 2 | 4 | 35 | 1862 | MG | 61 | * | 43 | 360 | 1510 |  | 2 | Y |  | 330 | 860 | 100 | [5 empty] | Y | * | * | *
2017 | 2 | 5 | 36 | 1863 | MG | 61 | * | 43 | 360 | 1510 |  | 2 | Y |  | 320 | 860 | 100 | [5 empty] | Y | * | * | *
2017 | 2 | 6 | 37 | 1864 | NB | 61 | * | 42 | 390 | 1500 |  | 2 | Y |  | 330 | 860 | 100 | [5 empty] | Y | * | * | *
2017 | 2 | 7 | 38 | 1865 | MG | 61 | * | 43 | 390 | 1500 |  | 2 | Y |  | 330 | 860 | 100 | [5 empty] | Y | * | * | *
2017 | 2 | 8 | 39 | 1866 | NB | 61 | * | 43 | 390 | 1460 |  | 2 | Y |  | 310 | 840 | 100 | [5 empty] | Y | * | * | * |  | A bit colder than normal in TAWO today due to high east winds, resulting in lower cylinder pressures.
2017 | 2 | 9 | 40 | 1867 | NB | 61 | * | 43 | 390 | 1470 |  | 2 | Y |  | 310 | 840 | 105 | [5 empty] | Y | * | * | *
2017 | 2 | 10 | 41 | 1868 | NB | 61 | * | 35 | 390 | 1490 |  | 2 | Y |  | 310 | 860 | 105 | [5 empty] | Y | * | * | * |  | MG made multiple trips up and down tower 1810-1919 z to address an issue with 10m windbird direction potentiometer. She also cleaned the frost from the CUGC sample inlet.
2017 | 2 | 11 | 42 | 1869 | [23 empty] | Did not go to TAWO due to whiteout conditions.
2017 | 2 | 12 | 43 | 1870 | NB | 61 | * | 43 | 390 | 1480 |  | 2 | Y |  | 310 | 860 | 105 | [5 empty] | Y | * | * | * |  | E-gen swap for unplanned service on one of the main generators occurred from 1337-1340z. No interruption in instrument operation was noted. A second power outage (switching back to the main generators) occured from 1936-1938z. Again, no interruption in instrument operation was noted.
2017 | 2 | 13 | 44 | 1871 | NB | 61 | * | 43 | 390 | 1470 |  | 2 | Y |  | 310 | 860 | 100 | [5 empty] | Y | * | * | *
2017 | 2 | 14 | 45 | 1872 | MG | 61 | * | 44 | 390 | 1480 |  | 2 | Y |  | 300 | 850 | 100 | [5 empty] | Y | * | * | *
2017 | 2 | 15 | 46 | 1873 | MG | 61 | * | 43 | 390 | 1490 |  | 2 | Y |  | 300 | 850 | 100 | [5 empty] | Y | * | * | *
2017 | 2 | 16 | 47 | 1874 | MG | 61 | * | 42 | 390 | 1480 |  | 2 | Y |  | 300 | 850 | 100 | [5 empty] | Y | * | * | *
2017 | 2 | 17 | 48 | 1875 | NB | 61 | * | 43 | 390 | 1470 |  | 2 | Y |  | 320 | 860 | 100 | [5 empty] | Y | * | * | *
2017 | 2 | 18 | 49 | 1876 | NB/AM/AR | 61 | * | 42 | 390 | 1460 |  | 2 | Y |  | 310 | 840 | 105 | [5 empty] | Y | * | * | * |  | Brief power outage from 1809-1813z for e-gen swap. No impacts were noted.
2017 | 2 | 19 | 50 | 1877 | MG/AR/AM | 61 | added 1 liter | 43 | 390 | 1460 |  | 2 | Y |  | 310 | 850 | 105 | [5 empty] | Y | * | * | *
2017 | 2 | 20 | 51 | 1878 | AM/AR/NB | 61 | * | 41 | 390 | 1470 |  | 2 | Y |  | 300 | 850 | 100 | [5 empty] | Y | * | * | *
2017 | 2 | 21 | 52 | 1879 | NB/AM/AR | 61 | * | 42 | 390 | 1460 |  | 2 | Y |  | 310 | 860 | 100 | [5 empty] | Y | * | * | *
2017 | 2 | 22 | 53 | 1880 | AM/AR | 61 | * | 44 | 375 | 1410 |  | 2 | Y |  | 300 | 840 | 100 | [5 empty] | Y | * | * | * |  | low temperatures inside tawo likely cause of low gas pressure readings
2017 | 2 | 23 | 54 | 1881 | AM/AR | 61 | * | 42 | 400 | 1460 |  | 2 | Y |  | 300 | 850 | 100 | [5 empty] | Y | * | * | * |  | Increased tank pressure due to warmer TAWO temps and high atmospheric pressure
2017 | 2 | 24 | 55 | 1882 | AM/AR | 61 | * | 43 | 400 | 1460 |  | 2 | Y |  | 300 | 850 | 100 | [5 empty] | Y | * | * | *
2017 | 2 | 25 | 56 | 1883 | AM | 61 | * | 45 | 375 | 1360 |  | 2 | Y |  | 290 | 800 | 100 | [5 empty] | Y | * | * | * |  | extreme winds prevented checking inlets
2017 | 2 | 26 | 57 | 1884 | AR/AM | 61 | * | 41 | 380 | 1410 |  | 2 | Y |  | 300 | 840 | 100 | [5 empty] | Y | * | * | *
2017 | 2 | 27 | 58 | 1885 | AR/AM | 61 | * | 41 | 390 | 1460 |  | 2 | Y |  | 300 | 900 | 100 | [5 empty] | Y | * | * | *
2017 | 2 | 28 | 59 | 1886 | AM | 61 | * | 43 | 400 | 1460 |  | 2 | Y |  | 300 | 860 | 100 | [5 empty] | Y | * | * | *
2017 | 3 | 1 | 60 | 1887 | AR | 61 | * | 43 | 390 | 1450 |  | 2 | Y |  | 310 | 850 | 100 | [5 empty] | Y | * | * | *
2017 | 3 | 2 | 61 | 1888 | AM | 61 | * | 43 | 400 | 1450 |  | 2 | Y |  | 300 | 850 | 105 | [5 empty] | Y | * | * | *
2017 | 3 | 3 | 62 | 1889 | AM/AR | 61 | * | 43 | 390 | 1450 |  | 2 | Y |  | 300 | 850 | 100 | [5 empty] | Y | * | * | *
2017 | 3 | 4 | 63 | 1890 | AR | 61 | * | 41 | 390 | 1440 |  | 2 | Y |  | 300 | 840 | 100 | [5 empty] | Y | * | * | *
2017 | 3 | 5 | 64 | 1891 | AM | 61 | added 1 liter | 43 | 400 | 1440 |  | 2 | Y |  | 300 | 850 | 105 | [5 empty] | Y | * | * | *
2017 | 3 | 6 | 65 | 1892 | AR | 61 | * | 41 | 390 | 1430 |  | 2 | Y |  | 300 | 850 | 105 | [5 empty] | Y | * | * | *
2017 | 3 | 7 | 66 | 1893 | AM | 61 | * | 43 | 400 | 1440 |  | 2 | Y |  | 300 | 840 | 105 | [5 empty] | Y | * | * | *
2017 | 3 | 8 | 67 | 1894 | AR | 61 | * | 43 | 380 | 1420 |  | 2 | Y |  | 300 | 840 | 105 | [5 empty] | Y | * | * | *
2017 | 3 | 9 | 68 | 1895 | AR/AM | 61 | * | 45 | 390 | 1410 |  | 2 | Y |  | 300 | 840 | 105 | [5 empty] | Y | * | * | *
2017 | 3 | 10 | 69 | 1896 | AM | 61 | * | 41 | 390 | 1420 |  | 2 | Y |  | 300 | 830 | 105 | [5 empty] | Y | * | * | *
2017 | 3 | 11 | 70 | 1897 | AR/HJ | 61 | * | 43 | 380 | 1420 |  | 2 | Y |  | 300 | 820 | 105 | [5 empty] | Y | * | * | *
2017 | 3 | 12 | 71 | 1898 | AM | 61 | * | 56 | 400 | 1440 |  | 2 | Y |  | 300 | 820 | 105 | [5 empty] | Y | * | * | * |  | accidentially reset hydrogen generator - appears to be workig normally
2017 | 3 | 13 | 72 | 1899 | AR/AM | 61 | * | 47 | 380 | 1420 |  | 2 | Y |  | 300 | 820 | 105 | [5 empty] | Y | * | * | *
2017 | 3 | 14 | 73 | 1900 | AM | 61 | * | 43 | 390 | 1410 |  | 2 | Y |  | 300 | 820 | 105 | [5 empty] | Y | * | * | *
2017 | 3 | 15 | 74 | 1901 | AR | 61 | * | 43 | 380 | 1400 |  | 2 | Y |  | 300 | 840 | 105 | [5 empty] | Y | * | * | *
2017 | 3 | 16 | 75 | 1902 | AM | 61 | * | 42 | 390 | 1410 |  | 2 | Y |  | 300 | 820 | 100 | [5 empty] | Y | * | * | *
2017 | 3 | 17 | 76 | 1903 | AM | 61 | * | 43 | 400 | 1410 |  | 2 | Y |  | 290 | 840 | 100 | [5 empty] | Y | * | * | *
2017 | 3 | 18 | 77 | 1904 | AM | 61 | * | 42 | 400 | 1400 |  | 2 | Y |  | 300 | 820 | 100 | [5 empty] | Y | * | * | *
2017 | 3 | 19 | 78 | 1905 | AR | 61 | * | 43 | 370 | 1400 |  | 2 | Y |  | 300 | 840 | 100 | [5 empty] | Y | * | * | *
2017 | 3 | 20 | 79 | 1906 | AR/AM | 61 | * | 43 | 390 | 1400 |  | 2 | Y |  | 290 | 830 | 100 | [5 empty] | Y | * | * | *
2017 | 3 | 21 | 80 | 1907 | AR | 61 | * | 41 | 370 | 1380 |  | 2 | Y |  | 300 | 800 | 100 | [5 empty] | Y | * | * | *
2017 | 3 | 22 | 81 | 1908 | AM | 61 | * | 42 | 390 | 1400 |  | 2 | Y |  | 300 | 820 | 105 | [5 empty] | Y | * | * | *
2017 | 3 | 23 | 82 | 1909 | AR | 61 | * | 41 | 380 | 1380 |  | 2 | Y |  | 300 | 830 | 100 | [5 empty] | Y | * | * | *
2017 | 3 | 24 | 83 | 1910 | AM | 61 | added 1 liter | 42 | 400 | 1410 |  | 2 | Y |  | 290 | 840 | 100 | [5 empty] | Y | * | * | * |  | H2 line dessicant becoming clear - emailed PI
2017 | 3 | 25 | 84 | 1911 | AR | _ | _ | _ | 370 | 1380 |  | 2 | y |  | 300 | 840 | 100 | [5 empty] | Y |  |  | * | 400 | H2 generator offline. Using cylinder in vestibule.
2017 | 3 | 26 | 85 | 1912 | AM | _ | _ | _ | 400 | 1400 |  | 2 | Y |  | 300 | 820 | 100 | [5 empty] | Y | _ | _ | _ | 350
2017 | 3 | 27 | 86 | 1913 | AM/AR | _ | _ | _ | 390 | 1400 |  | 2 | Y |  | 300 | 820 | 105 | [5 empty] | Y | _ | _ | _ | 300
2017 | 3 | 28 | 87 | 1914 | AM | _ | _ | _ | 390 | 1400 |  | 2 | Y |  | 300 | 820 | 105 | [5 empty] | Y | _ | _ | _ | 300
2017 | 3 | 29 | 88 | 1915 | AR | _ | _ | _ | 370 | 1380 |  | 2 | Y |  | 300 | 800 | 105 | [5 empty] | Y | _ | _ | _ | 300
2017 | 3 | 30 | 89 | 1916 | AM | _ | _ | _ | 390 | 1400 |  | 2 | Y |  | 300 | 820 | 105 | [5 empty] | Y | _ | _ | _ | 270
2017 | 3 | 31 | 90 | 1917 | AR | _ | _ | _ | 380 | 1380 |  | 2 | Y |  | 300 | 820 | 105 | [5 empty] | Y | _ | _ | _ | 280
2017 | 4 | 1 | 91 | 1918 | AM | _ | _ | _ | 390 | 1390 |  | 2 | Y |  | 300 | 820 | 105 | [5 empty] | Y | _ | _ | _ | 200 | per PI plan to swich H2 carrier gas cylinder on Monday 4/3
2017 | 4 | 2 | 92 | 1919 | AM | _ | _ | _ | 400 | 1400 |  | 2 | Y |  | 290 | 820 | 105 | [5 empty] | Y | _ | _ | _ | 190
2017 | 4 | 3 | 93 | 1920 | AM/AR | _ | _ | _ | 400 | 1380 |  | 2 | Y |  | 290 | 810 | 105 | [5 empty] | Y | _ | _ | _ | 190 | Swapped out H2 cylinder
2017 | 4 | 4 | 94 | 1921 | AR | _ | _ | _ | 370 | 1380 |  | 2 | Y |  | 300 | 820 | 105 | [5 empty] | Y | _ | _ | _ | 2300
2017 | 4 | 5 | 95 | 1922 | AM | _ | _ | _ | 400 | 1390 |  | 2 | y |  | 290 | 820 | 105 | [5 empty] | y | _ | _ | _ | 2280
2017 | 4 | 6 | 96 | 1923 | AR | _ | _ | _ | 380 | 1370 |  | 2 | y |  | 300 | 820 | 105 | [5 empty] | Y | _ | _ | _ | 2220
2017 | 4 | 7 | 97 | 1924 | AM | _ | _ | _ | 400 | 1390 |  | 2 | Y |  | 290 | 820 | 95 | [5 empty] | Y | _ | _ | _ | 2200
2017 | 4 | 8 | 98 | 1925 | AR | _ | _ | _ | 380 | 1360 |  | 2 | Y |  | 290 | 820 | 95 | [5 empty] | Y | _ | _ | _ | 2170
2017 | 4 | 9 | 99 | 1926 | AM | _ | _ | _ | 400 | 1360 |  | 2 | Y |  | 300 | 820 | 90 | [5 empty] | Y | _ | _ | _ | 2160
2017 | 4 | 10 | 100 | 1927 | AR/AM | _ | _ | _ | 400 | 1390 |  | 2 | Y |  | 290 | 820 | 90 | [5 empty] | Y | _ | _ | _ | 2190
2017 | 4 | 11 | 101 | 1928 | AR | _ | _ | _ | 380 | 1360 |  | 2 | Y |  | 290 | 820 | 90 | [5 empty] | Y | _ | _ | _ | 2140
2017 | 4 | 12 | 102 | 1929 | AM | _ | _ | _ | 400 | 1390 |  | 2 | Y |  | 290 | 820 | 90 | [5 empty] | Y | _ | _ | _ | 2140
2017 | 4 | 13 | 103 | 1930 | AM/AR | _ | _ | _ | 380 | 1360 |  | 2 | Y |  | 290 | 820 | 90 | [5 empty] | Y | _ | _ | _ | 2100
2017 | 4 | 14 | 104 | 1931 | AR | _ | _ | _ | 380 | 1360 |  | 2 | Y |  | 290 | 820 | 90 | [5 empty] | Y | _ | _ | _ | 2090
2017 | 4 | 15 | 105 | 1932 | AM | _ | _ | _ | 400 | 1360 |  | 2 | Y` |  | 300 | 820 | 90 | [5 empty] | Y | _ | _ | _ | 2040
2017 | 4 | 16 | 106 | 1933 | AR | _ | _ | _ | 380 | 1350 |  | 2 | Y` |  | 300 | 800 | 95 | [5 empty] | Y | _ | _ | _ | 1920
2017 | 4 | 17 | 107 | 1934 | AM/AR | _ | _ | _ | 380 | 1350 |  | 2 | Y |  | 290 | 810 | 95 | [5 empty] | Y | _ | _ | _ | 2050
2017 | 4 | 18 | 108 | 1935 | AM | _ | _ | _ | 380 | 1360 |  | 2 | Y |  | 290 | 810 | 90 | [5 empty] | _ | _ | _ | _ | 1970
2017 | 4 | 19 | 109 | 1936 | AR | _ | _ | _ | 380 | 1350 |  | 2 | Y |  | 290 | 810 | 95 | [5 empty] | _ | _ | _ | _ | 1910
2017 | 4 | 20 | 110 | 1937 | AM | _ | _ | _ | 390 | 1350 |  | 2 | Y |  | 290 | 810 | 95 | [5 empty] | _ | _ | _ |  | 1830
2017 | 4 | 21 | 111 | 1938 | AR | _ | _ | _ | 380 | 1340 |  | 2 | Y |  | 290 | 820 | 95 | [5 empty] | _ | _ | _ |  | 1810
2017 | 4 | 22 | 112 | 1939 | AM | _ | _ | _ | 390 | 1350 |  | 2 | Y |  | 290 | 810 | 90 | [5 empty] | _ | _ | _ | _ | 1790
2017 | 4 | 23 | 113 | 1940 | AR/AM | _ | _ | _ | 380 | 1340 |  | 2 | Y |  | 290 | 820 | 95 | [5 empty] | _ | _ | _ | _ | 1750
2017 | 4 | 24 | 114 | 1941 | AR | _ | _ | _ | 360 | 1330 |  | 2 | Y |  | 290 | 820 | 95 | [5 empty] | _ | _ | _ | _ | 1750
2017 | 4 | 25 | 115 | 1942 | AM | _ | _ | _ | 400 | 1350 |  | 2 | Y |  | 260 | 810 | 95 | [5 empty] | _ | _ | _ | _ | 1800
2017 | 4 | 26 | 116 | 1943 | AM | _ | _ | _ | 400 | 1350 |  | 2 | Y |  | 260 | 810 | 95 | [5 empty] | _ | _ | _ | _ | 1830
2017 | 4 | 27 | 117 | 1944 | AM | _ | _ | _ | 400 | 1350 |  | 2 | Y |  | 280 | 810 | 95 | [5 empty] | _ | _ | _ | _ | 1800
2017 | 4 | 28 | 118 | 1945 | AR | 61 | * | 45 | 380 | 1320 |  | 2 | Y |  | 290 | 810 | 95 | [5 empty] | _ | * | _ | _ | 1550 | Resumed using H2 generator April 28 2017 @ 1120 UTC
2017 | 4 | 29 | 119 | 1946 | AR | 61 | * | 43 | 380 | 1340 |  | 2 | Y |  | 280 | 820 | 95 | [5 empty] | _ | * | _ | _ | 1410
2017 | 4 | 30 | 120 | 1947 | [23 empty] | spreadsheed calendar off one day - April 30 was checked
2017 | 5 | 1 | 121 | 1948 | AM | 61 | * | 67 | 400 | 1350 |  | 2 | Y |  | 290 | 810 | 95 | [5 empty] | _ | * | replace | ok | _
2017 | 5 | 2 | 122 | 1949 | AR | 61 | * | 39 | 380 | 1310 |  | 2 | Y |  | 290 | 800 | 95 | [5 empty] | _ | * | replace | ok | _
2017 | 5 | 3 | 123 | 1950 | AM | 61 | * | 42 | 400 | 1340 |  | 2 | Y |  | 260 | 810 | 95 | [5 empty] | _ | * | replace | ok | _
2017 | 5 | 4 | 124 | 1951 | AR | 61 | * | 46 | 380 | 1310 |  | 2 | Y |  | 280 | 800 | 95 | [5 empty] | _ | * | replace | ok | _
2017 | 5 | 5 | 125 | 1952 | AM | 61 | * | 45 | 390 | 1340 |  | 2 | Y |  | 260 | 810 | 100 | [5 empty] | _ | * | replace | ok | _
2017 | 5 | 6 | 126 | 1953 | AR | 61 | * | 45 | 380 | 1300 |  | 2 | y |  | 280 | 800 | 100 | [5 empty] | _ | * | replace | ok | _
2017 | 5 | 7 | 127 | 1954 | AM | 61 | * | 41 | 400 | 1340 |  | 2 | y |  | 260 | 800 | 90 | [5 empty] | _ | * | replace | ok | - | -
2017 | 5 | 8 | 128 | 1955 | AM/AR | 61 | * | 45 | 380 | 1310 |  | 2 | Y |  | 280 | 800 | 95 | [5 empty] | - | * | replace | ok | 0
2017 | 5 | 9 | 129 | 1956 | AR/AM | 61 | * | 46 | 400 | 1310 |  | 2 | Y |  | 260 | 800 | 95 | [5 empty] | _ | * | replace | ok
2017 | 5 | 10 | 130 | 1957 | AR | 61 | * | 43 | 380 | 1310 |  | 2 | Y |  | 260 | 800 | 95 | [5 empty] | _ | ok | ok | ok | _
2017 | 5 | 11 | 131 | 1958 | AM | 61 | * | 44 | 400 | 1320 |  | 2 | Y |  | 250 | 800 | 95 | [5 empty] | - | * | * | *
2017 | 5 | 12 | 132 | 1959 | AR | 61 | * | 40 | 380 | 1310 |  | 2 | Y |  | 250 | 800 | 95 | [5 empty] | _ | * | * | * | _
2017 | 5 | 13 | 133 | 1960 | AM | 61 | * | 43 | 400 | 1310 |  | 2 | Y |  | 250 | 800 | 95 | [5 empty] | _ | * | * | * | _
2017 | 5 | 14 | 134 | 1961 | AR | 61 | * | 43 | 380 | 1300 |  | 2 | Y |  | 250 | 800 | 95 | [5 empty] | _ | * | * | * | _
2017 | 5 | 15 | 135 | 1962 | AM | 61 | * | 43 | 400 | 1310 |  | 2 | Y |  | 240 | 800 | 95 | [5 empty] | _ | * | * | * | _
2017 | 5 | 16 | 136 | 1963 | AR/AM | 61 | * | 44 | 400 | 1300 |  | 2 | Y |  | 250 | 800 | 95 | [5 empty] | _ | * | * | * | _
2017 | 5 | 17 | 137 | 1964 | AR/AM | 61 | * | 40 | 400 | 1310 |  | 2 | Y |  | 250 | 800 | 95 | [5 empty] | _ | * | * | * | _
2017 | 5 | 18 | 138 | 1965 | AM | 61 | * | 63 | 400 | 1310 |  | 2 | Y |  | 250 | 800 | 95 | [5 empty] | _ | * | * | * | _
2017 | 5 | 19 | 139 | 1966 | AM/HM | 61 | * | 43 | 400 | 1300 |  | 2 | Y |  | 250 | 800 | 95 | [5 empty] | _ | * | * | * | _
2017 | 5 | 20 | 140 | 1967 | AR/AM | 61 | * | 43 | 400 | 900 |  | 2 | Y |  | 250 | 800 | 95 | [5 empty] | _ | * | * | * | _
2017 | 5 | 21 | 141 | 1968 | HM/MG | 61 | * | 43 | 400 | 1300 |  | 2 | Y |  | 250 | 800 | 95 | [6 empty] | * | * | * | _
2017 | 5 | 22 | 142 | 1969 | HM/MG | 61 | * | 41 | 400 | 1300 |  | 2 | Y |  | 250 | 800 | 95 | [5 empty] | Y | * | * | * | _
2017 | 5 | 23 | 143 | 1970 | HM/MG | 61 | * | 43 | 400 | 1300 |  | 2 | Y |  | 250 | 800 | 95 | [5 empty] | Y | * | * | * | _
2017 | 5 | 24 | 144 | 1971 | HM | 61 | * | 41 | 400 | 1300 |  | 2 | Y |  | 250 | 800 | 95 | [5 empty] | Y | * | * | * | _
2017 | 5 | 25 | 145 | 1972 | HM | 61 | * | 41 | 400 | 1290 |  | 2 | Y |  | 250 | 800 | 95 | [5 empty] | Y | * | * | * | _
2017 | 5 | 26 | 146 | 1973 | HM | 61 | * | 41 | 400 | 1280 |  | 2 | Y |  | 250 | 800 | 95 | [5 empty] | Y | * | * | * | _
2017 | 5 | 27 | 147 | 1974 | HM | 61 | * | 43 | 400 | 1280 |  | 2 | Y |  | 250 | 800 | 95 | [5 empty] | Y | * | * | * | _
2017 | 5 | 28 | 148 | 1975 | HM | 61 | * | 43 | 400 | 1280 |  | 2 | Y |  | 250 | 800 | 95 | [5 empty] | Y | * | * | * | _
2017 | 5 | 29 | 149 | 1976 | MG | 61 | * | 43 | 400 | 1280 |  | 2 | Y |  | 250 | 790 | 95 | [5 empty] | Y | * | * | * | -
2017 | 5 | 30 | 150 | 1977 | MG | 61 | * | 44 | 400 | 1280 |  | 2 | Y |  | 240 | 790 | 95 | [5 empty] | Y | * | * | * | -
2017 | 5 | 31 | 151 | 1978 | MG | 61 | * | 44 | 400 | 1280 |  | 2 | Y |  | 240 | 790 | 95 | [5 empty] | Y | * | * | * | -
2017 | 6 | 1 | 152 | 1979 | HM | 61 | * | 43 | 400 | 1280 |  | 2 | Y |  | 240 | 790 | 95 | [5 empty] | Y | * | * | * | _
2017 | 6 | 2 | 153 | 1980 | MG | 61 | * | 42 | 390 | 1270 |  | 2 | Y |  | 240 | 790 | 95 | [5 empty] | Y | * | * | * | -
2017 | 6 | 3 | 154 | 1981 | HM | 61 | * | 41 | 390 | 1260 |  | 2 | Y |  | 240 | 790 | 95 | [5 empty] | Y | * | * | * | -
2017 | 6 | 4 | 155 | 1982 | MG | 61 | * | 39 | 390 | 1260 |  | 2 | Y |  | 240 | 790 | 95 | [5 empty] | Y | * | * | * | - | added 1 liter to H2 gen
2017 | 6 | 5 | 156 | 1983 | MG | 61 | * | 43 | 390 | 1270 |  | 2 | Y |  | 240 | 790 | 95 | [5 empty] | Y | * | * | * | *
2017 | 6 | 6 | 157 | 1984 | HM | 61 | * | 43 | 390 | 1250 |  | 2 | Y |  | 240 | 790 | 95 | [5 empty] | Y | * | * | * | *
2017 | 6 | 7 | 158 | 1985 | MG | 61 | * | 43 | 390 | 1250 |  | 2 | Y |  | 230 | 790 | 95 | [5 empty] | Y | * | * | * | *
2017 | 6 | 8 | 159 | 1986 | HM | 61 | * | 43 | 390 | 1250 |  | 2 | Y |  | 230 | 790 | 95 | [5 empty] | Y | * | * | * | *
2017 | 6 | 9 | 160 | 1987 | HM | 61 | * | 43 | 380 | 1250 |  | 2 | Y |  | 230 | 790 | 95 | [5 empty] | Y | * | * | * | *
2017 | 6 | 10 | 161 | 1988 | MG | 61 | * | 42 | 390 | 1360 |  | 2 | Y |  | 230 | 780 | 95 | [5 empty] | Y | * | * | * | *
2017 | 6 | 11 | 162 | 1989 | HM | 61 | * | 41 | 390 | 1250 |  | 2 | Y |  | 230 | 780 | 95 | [5 empty] | Y | * | * | * | -
2017 | 6 | 12 | 163 | 1990 | HM/MG | 61 | * | 44 | 390 | 1250 |  | 2 | Y |  | 230 | 770 | 95 | [5 empty] | Y | * | * | * | *
2017 | 6 | 13 | 164 | 1991 | HM/MG | 61 | * | 41 | 390 | 1260 |  | 2 | Y |  | 230 | 770 | 95 | [5 empty] | Y | * | * | * | *
2017 | 6 | 14 | 165 | 1992 | HM | 61 | * | 43 | 390 | 1240 |  | 2 | Y |  | 220 | 770 | 95 | [5 empty] | Y | * | * | * | *
2017 | 6 | 15 | 166 | 1993 | MG | 61 | * | 43 | 390 | 1240 |  | 2 | Y |  | 230 | 780 | 95 | [5 empty] | Y | * | * | * | *
2017 | 6 | 16 | 167 | 1994 | MG | 61 | * | 42 | 390 | 1240 |  | 2 | Y |  | 230 | 770 | 95 | [5 empty] | Y | * | * | * | *
2017 | 6 | 17 | 168 | 1995 | HM | 61 | * | 43 | 390 | 1240 |  | 2 | Y |  | 230 | 770 | 95 | [5 empty] | Y | * | * | * | *
2017 | 6 | 18 | 169 | 1996 | MG | 61 | * | 43 | 390 | 1240 |  | 2 | Y |  | 220 | 770 | 95 | [5 empty] | Y | * | * | * | *
2017 | 6 | 19 | 170 | 1997 | MG | 61 | * | 43 | 390 | 1240 |  | 2 | Y |  | 220 | 770 | 95 | [5 empty] | Y | * | * | * | * | added 1 liter to H2 gen
2017 | 6 | 20 | 171 | 1998 | HM | 61 | * | 43 | 390 | 1240 |  | 2 | Y |  | 220 | 770 | 95 | [5 empty] | Y | * | * | * | *
2017 | 6 | 21 | 172 | 1999 | HM | 61 | * | 43 | 390 | 1240 |  | 2 | Y |  | 220 | 770 | 95 | [5 empty] | Y | * | * | * | * | carps on roof
2017 | 6 | 22 | 173 | 2000 | MG | 61 | * | 43 | 390 | 1230 |  | 2 | Y |  | 220 | 770 | 95 | [5 empty] | Y | * | * | * | * | Carpenters still on roof
2017 | 6 | 23 | 174 | 2001 | MG | 61 | * | 44 | 390 | 1230 |  | 2 | Y |  | 220 | 770 | 95 | [5 empty] | Y | * | * | * | *
2017 | 6 | 24 | 175 | 2002 | MG | 61 | * | 38 | 390 | 1230 |  | 2 | Y |  | 220 | 770 | 95 | [5 empty] | Y | * | * | * | *
2017 | 6 | 25 | 176 | 2003 | HM | 61 | * | 43 | 390 | 1220 |  | 2 | Y |  | 220 | 770 | 95 | [5 empty] | Y | * | * | * | *
2017 | 6 | 26 | 177 | 2004 | MG | 61 | * | 43 | 390 | 1220 |  | 2 | Y |  | 210 | 760 | 95 | [5 empty] | Y | * | * | * | *
2017 | 6 | 27 | 178 | 2005 | MG | 61 | * | 42 | 390 | 1220 |  | 2 | Y |  | 220 | 760 | 95 | [5 empty] | Y | * | * | * | *
2017 | 6 | 28 | 179 | 2006 | MG | 61 | * | 42 | 390 | 1220 |  | 2 | Y |  | 220 | 760 | 95 | [5 empty] | Y | * | * | * | *
2017 | 6 | 29 | 180 | 2007 | MG | 61 | * | 43 | 390 | 1220 |  | 2 | Y |  | 220 | 760 | 95 | [5 empty] | Y | * | * | * | -
2017 | 6 | 30 | 181 | 2008 | HM | 61 | * | 43 | 390 | 1220 |  | 2 | Y |  | 220 | 760 | 95 | [5 empty] | Y | * | * | * | -
2017 | 7 | 1 | 182 | 2009 | MG | 61 | * | 43 | 390 | 1220 |  | 2 | Y |  | 220 | 760 | 95 | [5 empty] | Y | * | * | * | -
2017 | 7 | 2 | 183 | 2010 | MG | 61 | * | 43 | 390 | 1220 |  | 2 | Y |  | 220 | 760 | 95 | [5 empty] | Y | * | * | * | -
2017 | 7 | 3 | 184 | 2011 | MG | 61 | * | 42 | 390 | 1210 |  | 2 | Y |  | 220 | 760 | 95 | [5 empty] | Y | * | * | * | -
2017 | 7 | 4 | 185 | 2012 | HM | 61 | * | 42 | 390 | 1210 |  | 2 | Y |  | 220 | 760 | 95 | [5 empty] | Y | * | * | * | -
2017 | 7 | 5 | 186 | 2013 | MG | 61 | * | 43 | 390 | 1210 |  | 2 | Y |  | 220 | 760 | 95 | [5 empty] | Y | * | * | * | -
2017 | 7 | 6 | 187 | 2014 | HM | 61 | * | 35 | 390 | 1200 |  | 2 | Y |  | 220 | 760 | 95 | [5 empty] | Y | * | * | * | -
2017 | 7 | 7 | 188 | 2015 | MG | 61 | * | 43 | 390 | 1210 |  | 2 | Y |  | 220 | 760 | 95 | [5 empty] | Y | * | * | * | -
2017 | 7 | 8 | 189 | 2016 | HM | 61 | * | 43 | 390 | 1200 |  | 2 | Y |  | 200 | 760 | 95 | [5 empty] | Y | * | * | * | -
2017 | 7 | 9 | 190 | 2017 | HM | 61 | * | 38 | 400 | 1190 |  | 2 | y |  | 200 | 750 | 95 | [5 empty] | Y | * | * | * | -
2017 | 7 | 10 | 191 | 2018 | MG | 61 | * | 43 | 390 | 1190 |  | 2 | y |  | 210 | 760 | 95 | [5 empty] | Y | * | * | * | -
2017 | 7 | 11 | 192 | 2019 | MG/HM | 61 | * | 43 | 390 | 1190 |  | 2 | y |  | 210 | 760 | 95 | [5 empty] | Y | * | * | * | -
2017 | 7 | 12 | 193 | 2020 | MG | 61 | * | 42 | 390 | 1180 |  | 2 | y |  | 210 | 760 | 95 | [5 empty] | y | * | * | * | -
2017 | 7 | 13 | 194 | 2021 | HM | 61 | * | 40 | 390 | 1180 |  | 2 | Y |  | 200 | 750 | 92 | [5 empty] | Y | * | * | * | -
2017 | 7 | 14 | 195 | 2022 | HM | 61 | * | 43 | 390 | 1180 |  | 2 | Y |  | 200 | 750 | 102 | [5 empty] | Y | * | * | * | -
2017 | 7 | 15 | 196 | 2023 | MG | 61 | * | 42 | 390 | 1170 |  | 2 | Y |  | 210 | 770 | 105 | [5 empty] | Y | * | * | * | -
2017 | 7 | 16 | 197 | 2024 | MG | 61 | * | 40 | 390 | 1170 |  | 2 | Y |  | 210 | 770 | 105 | [5 empty] | Y | * | * | * | -
2017 | 7 | 17 | 198 | 2025 | MG/HM | 61 | * | 34 | 390 | 1160 |  | 2 | Y |  | 210 | 770 | 105 | [5 empty] | Y | * | * | *
2017 | 7 | 18 | 199 | 2026 | MG | 61 | * | 43 | 390 | 1160 |  | 2 | Y |  | 210 | 760 | 105 | [5 empty] | Y | * | * | * | -
2017 | 7 | 19 | 200 | 2027 | HM | 61 | * | 39 | 390 | 1160 |  | 2 | Y |  | 200 | 750 | 95 | [5 empty] | Y | * | * | * | -
2017 | 7 | 20 | 201 | 2028 | HM/MG | 61 | * | 42 | 390 | 1170 |  | 2 | Y |  | 200 | 770 | 105 | [5 empty] | Y | * | * | * | -
2017 | 7 | 21 | 202 | 2029 | MG | 61 | * | 43 | 390 | 1160 |  | 2 | Y |  | 210 | 760 | 105 | [5 empty] | Y | * | * | * | -
2017 | 7 | 22 | 203 | 2030 | MG | 61 | * | 43 | 390 | 1160 |  | 2 | Y |  | 200 | 760 | 105 | [5 empty] | Y | * | * | * | -
2017 | 7 | 23 | 204 | 2031 | HM | 61 | * | 43 | 390 | 1160 |  | 2 | Y |  | 200 | 760 | 105 | [5 empty] | Y | * | * | * | -
2017 | 7 | 24 | 205 | 2032 | HM/MG | 61 | * | 43 | 390 | 1160 |  | 2 | Y |  | 210 | 760 | 105 | [5 empty] | Y | * | * | * | -
2017 | 7 | 25 | 206 | 2033 | MG | 61 | * | 43 | 390 | 1160 |  | 2 | Y |  | 210 | 760 | 105 | [5 empty] | Y | * | * | * | -
2017 | 7 | 26 | 207 | 2034 | HM | 61 | * | 43 | 390 | 1150 |  | 2 | Y |  | 210 | 750 | 105 | [5 empty] | y | * | * | * | -
2017 | 7 | 27 | 208 | 2035 | HM | 61 | * | 43 | 390 | 1150 |  | 2 | Y |  | 200 | 750 | 105 | [5 empty] | Y | * | * | * | -
2017 | 7 | 28 | 209 | 2036 | HM | 61 | * | 43 | 390 | 1150 |  | 2 | Y |  | 200 | 750 | 105 | [5 empty] | Y | * | * | * | - | Scheduled power outage at 1130 UTC
2017 | 7 | 29 | 210 | 2037 | MG | 61 | * | 44 | 390 | 1150 |  | 2 | Y |  | 200 | 750 | 105 | [5 empty] | Y | * | * | * | -
2017 | 7 | 30 | 211 | 2038 | MG | 61 | * | 44 | 390 | 1150 |  | 2 | Y |  | 200 | 750 | 105 | [5 empty] | Y | * | * | * | -
2017 | 7 | 31 | 212 | 2039 | HM | 61 | * | 42 | 390 | 1150 |  | 2 | Y |  | 200 | 750 | 105 | [5 empty] | Y | * | * | * | -
2017 | 8 | 1 | 213 | 2040 | HM | 61 | * | 43 | 390 | 1150 |  | 2 | Y |  | 200 | 750 | 105 | [5 empty] | Y | * | * | * | -
2017 | 8 | 2 | 214 | 2041 | MG | 61 | * | 43 | 390 | 1140 |  | 2 | Y |  | 200 | 750 | 105 | [5 empty] | Y | * | * | * | -
2017 | 8 | 3 | 215 | 2042 | HM | 61 | * | 42 | 390 | 1140 |  | 2 | Y |  | 200 | 750 | 105 | [5 empty] | Y | * | * | * | -
2017 | 8 | 4 | 216 | 2043 | MG | 61 | * | 43 | 390 | 1140 |  | 2 | Y |  | 190 | 750 | 105 | [5 empty] | Y | * | * | * | -
2017 | 8 | 5 | 217 | 2044 | MG | 61 | * | 43 | 390 | 1140 |  | 2 | Y |  | 190 | 750 | 105 | [5 empty] | Y | * | *
2017 | 8 | 6 | 218 | 2045 | HM | 61 | * | 43 | 390 | 1140 |  | 2 | Y |  | 190 | 750 | 105 | [5 empty] | Y | * | * | * | -
2017 | 8 | 7 | 219 | 2046 | HM | 61 | * | 43 | 390 | 1140 |  | 2 | Y |  | 190 | 750 | 105 | [5 empty] | Y | * | * | * | -
2017 | 8 | 8 | 220 | 2047 | HM | 61 | * | 43 | 390 | 1130 |  | 2 | Y |  | 190 | 750 | 105 | [5 empty] | Y | * | * | * | -
2017 | 8 | 9 | 221 | 2048 | MG | 61 | * | 43 | 390 | 1130 |  | 2 | Y |  | 190 | 750 | 105 | [5 empty] | Y | * | * | * | -
2017 | 8 | 10 | 222 | 2049 | HM | 61 | * | 43 | 390 | 1130 |  | 2 | Y |  | 190 | 750 | 105 | [5 empty] | Y | * | * | * | -
2017 | 8 | 11 | 223 | 2050 | HM | 61 | * | 43 | 390 | 1130 |  | 2 | Y |  | 190 | 750 | 105 | [5 empty] | Y | * | * | * | - | Scheduled Power Outage 1030-1033 UTC
2017 | 8 | 12 | 224 | 2051 | MG | 61 | * | 43 | 390 | 1120 |  | 2 | Y |  | 190 | 750 | 105 | [5 empty] | Y | * | * | * | -
2017 | 8 | 13 | 225 | 2052 | MG | 61 | * | 43 | 390 | 1120 |  | 2 | Y |  | 190 | 750 | 105 | [5 empty] | Y | * | * | * | -
2017 | 8 | 14 | 226 | 2053 | MG | 61 | * | 43 | 390 | 1120 |  | 2 | Y |  | 190 | 750 | 100 | [5 empty] | Y | * | * | * | -
2017 | 8 | 15 | 227 | 2054 | HM | 61 | * | 42 | 390 | 1120 |  | 2 | Y |  | 190 | 750 | 105 | [5 empty] | Y | * | * | * | -
2017 | 8 | 16 | 228 | 2055 | HM | 61 | * | 41 | 390 | 1120 |  | 2 | Y |  | 190 | 750 | 105 | [5 empty] | Y | * | * | * | -
2017 | 8 | 17 | 229 | 2056 | MG | 61 | * | 43 | 390 | 1120 |  | 2 | Y |  | 190 | 750 | 100 | [5 empty] | Y | * | * | * | - | Added 1 liter of water
2017 | 8 | 18 | 230 | 2057 | MG | 61 | * | 43 | 390 | 1120 |  | 2 | Y |  | 160 | 750 | 100 | [5 empty] | Y | * | * | * | - | Rainer and Jacques arrived for performing the VOC audit
2017 | 8 | 19 | 231 | 2058 | HM | 61 | * | 42 | 390 | 1110 |  | 2 | Y |  | 0 | 750 | 105 | [5 empty] | Y | * | * | * | - | J. Hueber on site performing audit
2017 | 8 | 20 | 232 | 2059 | HM | 61 | * | 43 | 390 | 1110 |  | 2 | Y |  | 0 | 750 | 100 | [5 empty] | Y | * | * | * | - | J. Hueber on site performing audit
2017 | 8 | 21 | 233 | 2060 | HM | 61 | * | 42 | 390 | 1100 |  | 2 | Y |  | 0 | 750 | 100 | [5 empty] | Y | * | * | * | - | J. Hueber on site performing audit; Scheduled power outage 1200 UTC and 1400 UTC.
2017 | 8 | 22 | 234 | 2061 | HM | 61 | * | 43 | 100 | 1100 |  | 2 | Y |  | 0 | 750 | 100 | [5 empty] | Y | * | * | * | -
2017 | 8 | 23 | 235 | 2062 | HM | 61 | * | 43 | 400 | 1100 |  | 2 | Y |  | 0 | 750 | 100 | [5 empty] | Y | * | * | * | - | J. Hueber on site performing audit
2017 | 8 | 24 | 236 | 2063 | HM | 61 | * | 43 | 400 | 1100 |  | 2 | Y |  | 200 | 750 | 100 | [5 empty] | Y | * | * | * | - | J. Hueber on site performing audit
2017 | 8 | 25 | 237 | 2064 | MG | - | - | - | 400 | 1100 |  | 2 | Y |  | 180 | 750 | 100 | [5 empty] | Y | - | * | * | 1820 | power outage 1515-1519, J. Hueber on site performing audit, drained Hydrogen generator, CU GC now hooked up to Hydrogen tank in vestibule.
2017 | 8 | 26 | 238 | 2065 | HM | - | - | - | 300 | 1500 |  | - | - |  | 420 | 700 | - | [5 empty] | - | - | - | - | 1800 | Methane Std 1 swithced to large tank; CU GC turned off 8/25/17 for winter ops.
2017 | 8 | 27 | 239 | 2066 | HM | - | - | - | 0 | 1500 |  | - | - |  | 420 | 520 | - | [5 empty] | - | - | - | - | 1790
2017 | 8 | 28 | 240 | 2067 | HM | - | - | - | 0 | 1500 |  | - | - |  | 420 | 520 | - | [5 empty] | - | - | - | - | 1790
2017 | 8 | 29 | 241 | 2068 | HM | - | - | - | 0 | 1500 |  | - | - |  | 400 | 150 | - | [5 empty] | - | - | - | - | 1700
2017 | 8 | 30 | 242 | 2069 | HM | - | - | - | 0 | 1500 |  | - | - |  | 400 | 150 | - | [5 empty] | - | - | - | - | 1700 | Checked all tank valves- all were closed.
2017 | 8 | 31 | 243 | 2070 | HM | - | - | - | 0 | 1500 |  | - | - |  | 400 | 150 | - | [5 empty] | - | - | - | - | 1700
2017 | 9 | 1 | 244 | 2071
2017 | 9 | 2 | 245 | 2072
2017 | 9 | 3 | 246 | 2073
2017 | 9 | 4 | 247 | 2074
2017 | 9 | 5 | 248 | 2075
2017 | 9 | 6 | 249 | 2076
2017 | 9 | 7 | 250 | 2077
2017 | 9 | 8 | 251 | 2078
2017 | 9 | 9 | 252 | 2079
2017 | 9 | 10 | 253 | 2080
2017 | 9 | 11 | 254 | 2081
2017 | 9 | 12 | 255 | 2082
2017 | 9 | 13 | 256 | 2083
2017 | 9 | 14 | 257 | 2084
2017 | 9 | 15 | 258 | 2085
2017 | 9 | 16 | 259 | 2086
2017 | 9 | 17 | 260 | 2087
2017 | 9 | 18 | 261 | 2088
2017 | 9 | 19 | 262 | 2089
2017 | 9 | 20 | 263 | 2090
2017 | 9 | 21 | 264 | 2091
2017 | 9 | 22 | 265 | 2092
2017 | 9 | 23 | 266 | 2093
2017 | 9 | 24 | 267 | 2094
2017 | 9 | 25 | 268 | 2095
2017 | 9 | 26 | 269 | 2096
2017 | 9 | 27 | 270 | 2097
2017 | 9 | 28 | 271 | 2098
2017 | 9 | 29 | 272 | 2099
2017 | 9 | 30 | 273 | 2100
2017 | 10 | 1 | 274 | 2101
2017 | 10 | 2 | 275 | 2102
2017 | 10 | 3 | 276 | 2103
2017 | 10 | 4 | 277 | 2104
2017 | 10 | 5 | 278 | 2105
2017 | 10 | 6 | 279 | 2106
2017 | 10 | 7 | 280 | 2107
2017 | 10 | 8 | 281 | 2108
2017 | 10 | 9 | 282 | 2109
2017 | 10 | 10 | 283 | 2110
2017 | 10 | 11 | 284 | 2111
2017 | 10 | 12 | 285 | 2112
2017 | 10 | 13 | 286 | 2113
2017 | 10 | 14 | 287 | 2114
2017 | 10 | 15 | 288 | 2115
2017 | 10 | 16 | 289 | 2116
2017 | 10 | 17 | 290 | 2117
2017 | 10 | 18 | 291 | 2118
2017 | 10 | 19 | 292 | 2119
2017 | 10 | 20 | 293 | 2120 | HJ/HA | 61 | * | 0 | 1500 | 1100 |  | 1.75 | 1 |  | 400 | 700 | 0 | [6 empty] | * | * | * |  | 1830 UTC, CUCG turned back on with Jacques on phone. Changed dessicant and refilled H2 generator with DI water. Running on cylinders until enough H2 is generated. 1 of 4 pumps running, notified Jacques. Did not read flow correctly
2017 | 10 | 21 | 294 | 2121
2017 | 10 | 22 | 295 | 2122
2017 | 10 | 23 | 296 | 2123 | HA | 1.2 | * | 0 | 1500 | 1100 |  | 2 | Y |  | 450 | 750 | 110 | [6 empty] | * | * | * | 1600 | H2 Generator had alarm going off; needed to reset psi- called Jacque on the phone and restarted generator and successfully reset. Did not read flow correctly
2017 | 10 | 24 | 297 | 2124
2017 | 10 | 25 | 298 | 2125
2017 | 10 | 26 | 299 | 2126
2017 | 10 | 27 | 300 | 2127
2017 | 10 | 28 | 301 | 2128
2017 | 10 | 29 | 302 | 2129
2017 | 10 | 30 | 303 | 2130 | HA | 61 | * | 41 | 1500 | 1100 |  | 1.75 | Y |  | 425 | 750 | 110 | [6 empty] | * | * | * | 1625 | H2 Generator alarm going off indicating low pressure. Called Jacque on phone and worked through procedure to correct float chamber error by draining water from generator, wiping probes within float chamber dry, reassembling system, and resetting. Switched valve to source H2 from generator and not H2 tank.
2017 | 10 | 31 | 304 | 2131
2017 | 11 | 1 | 305 | 2132
2017 | 11 | 2 | 306 | 2133
2017 | 11 | 3 | 307 | 2134
2017 | 11 | 4 | 308 | 2135
2017 | 11 | 5 | 309 | 2136
2017 | 11 | 6 | 310 | 2137 | HA | 61 | * | 42 | 1500 | 1100 |  | 1.75 | Y |  | 460 | 715 | 100 | [6 empty] | * | * |  | 1525
2017 | 11 | 7 | 311 | 2138
2017 | 11 | 8 | 312 | 2139
2017 | 11 | 9 | 313 | 2140
2017 | 11 | 10 | 314 | 2141
2017 | 11 | 11 | 315 | 2142
2017 | 11 | 12 | 316 | 2143
2017 | 11 | 13 | 317 | 2144
2017 | 11 | 14 | 318 | 2145 | HJ | 61 | * | 43 | 1500 | 1050 |  | 1.75 | Y |  | 425 | 725 | 100 | [6 empty] | * | * |  | 1525 | 1700Z- climbed tower to clear inlet of snow- very clogged.
2017 | 11 | 15 | 319 | 2146
2017 | 11 | 16 | 320 | 2147
2017 | 11 | 17 | 321 | 2148
2017 | 11 | 18 | 322 | 2149 | HJ | 61 | * | 44 | 1500 | 1050 |  | 1.75 | Y |  | 425 | 725 | 100 | [6 empty] | * | * |  | 1525 | Scheduled power outage 1303-1305, 1306-1308 UTC
2017 | 11 | 19 | 323 | 2150
2017 | 11 | 20 | 324 | 2151 | HA | 61 | * | 44 | 1500 | 1000 |  | 1.75 | Y |  | 425 | 725 | 100 | [6 empty] | * | * |  | 1500 | scheduled power outage 1702-1704 UTC
2017 | 11 | 21 | 325 | 2152
2017 | 11 | 22 | 326 | 2153
2017 | 11 | 23 | 327 | 2154
2017 | 11 | 24 | 328 | 2155 | HA | 61 | * | 43 | 1500 | 1000 |  | 1.75 | Y |  | 425 | 700 | 100 | [6 empty] | * | * |  | 1500
2017 | 11 | 25 | 329 | 2156
2017 | 11 | 26 | 330 | 2157
2017 | 11 | 27 | 331 | 2158 | HA | 61 | * | 43 | 1500 | 1000 |  | 1.75 | Y |  | 425 | 725 | 100 | [6 empty] | * | * |  | 1600 | Scheduled power outage 17:01-17:03 UTC
2017 | 11 | 28 | 332 | 2159
2017 | 11 | 29 | 333 | 2160
2017 | 11 | 30 | 334 | 2161 | HA | 61 | * | 41 | 1500 | 1000 |  | 1.75 | Y |  | 425 | 725 | 100 | [6 empty] | * | * |  | 1600
2017 | 12 | 1 | 335 | 2162
2017 | 12 | 2 | 336 | 2163
2017 | 12 | 3 | 337 | 2164
2017 | 12 | 4 | 338 | 2165 | HA | 61 | * | 43 | 1500 | 1000 |  | 1.75 | Y |  | 425 | 725 | 100 | [6 empty] | * | * |  | 1550
2017 | 12 | 5 | 339 | 2166 | 0
2017 | 12 | 6 | 340 | 2167
2017 | 12 | 7 | 341 | 2168 | 0
2017 | 12 | 8 | 342 | 2169 | HA | 61 | * | 43 | 1500 | 1000 |  | 1.75 | Y |  | 425 | 700 | 100 | [6 empty] | * | * |  | 1550
2017 | 12 | 9 | 343 | 2170 | 0
2017 | 12 | 10 | 344 | 2171 | 0
2017 | 12 | 11 | 345 | 2172 | HA | 61 | * | 44 | 1500 | 1000 |  | 1.75 | Y |  | 425 | 725 | 100 | [6 empty] | * | * |  | 1600 | H2 generator float chamber error; dried probes on float chamber head twice. Called Jacques and reported high flow rate when H2 gen valve was closed, troubleshoot potential causes with Jaqcue. Returned to normal operation with flow returning to accepted level, but unsure as to why.
2017 | 12 | 12 | 346 | 2173 | HA | 61 | * | 44 | 1500 | 975 |  | 1.75 | Y |  | 425 | 700 | 100 | [6 empty] | * | * |  | 1300
2017 | 12 | 13 | 347 | 2174 | 0
2017 | 12 | 14 | 348 | 2175 | 0
2017 | 12 | 15 | 349 | 2176 | HA | 61 | * | 44 | 1500 | 975 |  | 1.75 | Y |  | 425 | 700 | 100 | [6 empty] | * | * |  | 1500
2017 | 12 | 16 | 350 | 2177 | 0
2017 | 12 | 17 | 351 | 2178 | 0
2017 | 12 | 18 | 352 | 2179 | HA | 61 | * | 43 | 1500 | 975 |  | 1.75 | Y |  | 425 | 700 | 100 | [6 empty] | * | * |  | 1525
2017 | 12 | 19 | 353 | 2180 | HA | 61 | * | 42 | 1500 | 975 |  | 1.75 | Y |  | 425 | 700 | 100 | [6 empty] | * | * |  | 1525
2017 | 12 | 20 | 354 | 2181 | 0
2017 | 12 | 21 | 355 | 2182 | 0
2017 | 12 | 22 | 356 | 2183 | HA | 61 | * | 43 | 1500 | 975 |  | 1.75 | Y |  | 425 | 700 | 100 | [6 empty] | * | * |  | 1500
2017 | 12 | 23 | 357 | 2184 | 0
2017 | 12 | 24 | 358 | 2185 | 0
2017 | 12 | 25 | 359 | 2186 | HA | 61 | * | 42 | 1500 | 950 |  | 1.75 | Y |  | 425 | 700 | 100 | [6 empty] | * | * |  | 1500
2017 | 12 | 26 | 360 | 2187
2017 | 12 | 27 | 361 | 2188 | HA | 61 | * | 44 | 1500 | 950 |  | 1.75 | Y |  | 425 | 700 | 100 | [6 empty] | * | * |  | 1490
2017 | 12 | 28 | 362 | 2189
2017 | 12 | 29 | 363 | 2190
2017 | 12 | 30 | 364 | 2191 | HA | 61 | * | 42 | 1500 | 950 |  | 1.75 | Y |  | 425 | 700 | 100 | [6 empty] | * | * |  | 1450 | added I liter of water to H2 generator
2017 | 12 | 31 | 365 | 2192
2018 | 1 | 1 | 1 | 2193 | HA | 61 | * | 42 | 1500 | 950 |  | 1.75 | Y |  | 425 | 700 | 100 | [6 empty] | * | * |  | 1450
2018 | 1 | 2 | 2 | 2194 | [23 empty] | Could not visit TAWO due to weather
2018 | 1 | 3 | 3 | 2195 | [23 empty] | Could not visit TAWO due to weather
2018 | 1 | 4 | 4 | 2196 | [23 empty] | Could not visit TAWO due to weather
2018 | 1 | 5 | 5 | 2197 | HA | 61 | * | 43 | 1500 | 900 |  | 1.75 | Y |  | 425 | 700 | 100 | [6 empty] | * | * |  | 1450
2018 | 1 | 6 | 6 | 2198
2018 | 1 | 7 | 7 | 2199
2018 | 1 | 8 | 8 | 2200 | HA | 61 | * | 42 | 1500 | 900 |  | 1.75 | Y |  | 425 | 700 | 100 | [6 empty] | * | * |  | 1450
2018 | 1 | 9 | 9 | 2201
2018 | 1 | 10 | 10 | 2202 | HA | 61 | * | 40 | 1500 | 900 |  | 1.75 | Y |  | 425 | 700 | 100 | [6 empty] | * | *
2018 | 1 | 11 | 11 | 2203
2018 | 1 | 12 | 12 | 2204 | HA | 61 | added 1 liter | 42 | 1500 | 900 |  | 1.75 | y |  | 425 | 700 | 100 | [6 empty] | * | * |  | 1425
2018 | 1 | 13 | 13 | 2205
2018 | 1 | 14 | 14 | 2206
2018 | 1 | 15 | 15 | 2207 | HA | 61 | * | 42 | 1500 | 900 |  | 1.75 | Y |  | 425 | 700 | 100 | [6 empty] | * | * |  | 1375
2018 | 1 | 16 | 16 | 2208 | 0
2018 | 1 | 17 | 17 | 2209 | 0
2018 | 1 | 18 | 18 | 2210 | HA | 61 | * | 43 | 1500 | 900 |  | 1.75 | Y |  | 425 | 700 | 100 | [6 empty] | * | * |  | 1400
2018 | 1 | 19 | 19 | 2211 | 0
2018 | 1 | 20 | 20 | 2212 | HA | 61 | * | 38 | 1500 | 900 |  | 1.75 | Y |  | 425 | 700 | 100 | [6 empty] | * | * |  | 1400
2018 | 1 | 21 | 21 | 2213 | 0
2018 | 1 | 22 | 22 | 2214 | HA | 61 | * | 37 | 1500 | 850 |  | 1.75 | Y |  | 425 | 700 | 90 | [6 empty] | * | * |  | 1400
2018 | 1 | 23 | 23 | 2215 | 0
2018 | 1 | 24 | 24 | 2216 | HA | 61 | * | 42 | 1500 | 850 |  | 1.75 | Y |  | 425 | 700 | 100 | [6 empty] | * | * |  | 1375
2018 | 1 | 25 | 25 | 2217 | 0
2018 | 1 | 26 | 26 | 2218 | 0
2018 | 1 | 27 | 27 | 2219 | 0
2018 | 1 | 28 | 28 | 2220 | 0
2018 | 1 | 29 | 29 | 2221 | HA | 61 | added 1/2 liter DI water | 42 | 1500 | 850 |  | 1.75 | Y |  | 425 | 700 | 100 | [6 empty] | * | * |  | 1400
2018 | 1 | 30 | 30 | 2222 | 0
2018 | 1 | 31 | 31 | 2223 | HA | 61 | * | 42 | 1500 | 850 |  | 1.75 | Y |  | 425 | 700 | 100 | [6 empty] | * | * |  | 1375
2018 | 2 | 1 | 32 | 2224 | 0
2018 | 2 | 2 | 33 | 2225 | 0
2018 | 2 | 3 | 34 | 2226 | HA | 61 | * | 42 | 1500 | 850 |  | 1.75 | Y |  | 425 | 700 | 100 | [6 empty] | * | * |  | 1350
2018 | 2 | 4 | 35 | 2227
2018 | 2 | 5 | 36 | 2228
2018 | 2 | 6 | 37 | 2229
2018 | 2 | 7 | 38 | 2230
2018 | 2 | 8 | 39 | 2231
2018 | 2 | 9 | 40 | 2232
2018 | 2 | 10 | 41 | 2233
2018 | 2 | 11 | 42 | 2234
2018 | 2 | 12 | 43 | 2235
2018 | 2 | 13 | 44 | 2236
2018 | 2 | 14 | 45 | 2237 | SWD | 61 | * | 41 | 1500 | 800 |  | 1.75 | Y |  |  | 670 | 95 | [6 empty] | * | * |  | 1320
2018 | 2 | 15 | 46 | 2238
2018 | 2 | 16 | 47 | 2239
2018 | 2 | 17 | 48 | 2240 | SWD | 61 | * | 42 | 1500 | 800 |  | 1.75 | Y |  |  | 670 | 100 | [6 empty] | * | * | * | 1320
2018 | 2 | 18 | 49 | 2241
2018 | 2 | 19 | 50 | 2242
2018 | 2 | 20 | 51 | 2243
2018 | 2 | 21 | 52 | 2244 | SWD | 61 | * | 41 | 1500 | 800 |  | 1.75 | Y |  |  | 670 | 95 | [6 empty] | * | * | * | 1320
2018 | 2 | 22 | 53 | 2245
2018 | 2 | 23 | 54 | 2246
2018 | 2 | 24 | 55 | 2247
2018 | 2 | 25 | 56 | 2248 | SWD | 61 | * | 59 | 1500 | 800 |  | 1.75 | Y |  |  | 670 | 95 | [6 empty] | * | * | * | 1320
2018 | 2 | 26 | 57 | 2249 | SWD | 61 | * | 42 | 1500 | 800 |  | 1.75 | Y |  |  | 670 | 95 | [6 empty] | * | * | * | 1320
2018 | 2 | 27 | 58 | 2250 | SWD | 61 | * | 42 | 1500 | 780 |  | 1.75 | Y |  |  | 670 | 95 | [6 empty] | * | * | * | 1320
2018 | 2 | 28 | 59 | 2251
2018 | 3 | 1 | 60 | 2252
2018 | 3 | 2 | 61 | 2253 | SWD | 61 | * | 42 | 1500 | 780 |  | 1.75 | Y |  |  | 670 | 100 | [6 empty] | * | * | * | 1320
2018 | 3 | 3 | 62 | 2254
2018 | 3 | 4 | 63 | 2255
2018 | 3 | 5 | 64 | 2256 | SWD | 61 | * | 42 | 1520 | 750 |  | 1.75 | Y |  |  | 670 | 100 | [6 empty] | * | * | * | 1320 | Added DI water to H2 generator at time of check.
2018 | 3 | 6 | 65 | 2257
2018 | 3 | 7 | 66 | 2258 | SWD | 61 | * | 40 | 1520 | 750 |  | 1.75 | Y |  | 380 | 670 | 95 | [6 empty] | * | * | * | 1320 | H2 Generator display unable to show full-precision 1 mL/min flow rate due to issue with button; instead the filled-square indicator is used (1 square = 20 mL/min).
2018 | 3 | 8 | 67 | 2259
2018 | 3 | 9 | 68 | 2260
2018 | 3 | 10 | 69 | 2261 | SWD | 61 | * | 40 | 1520 | 750 |  | 1.75 | Y |  | 380 | 660 | 95 | [6 empty] | * | * | * | 1320
2018 | 3 | 11 | 70 | 2262
2018 | 3 | 12 | 71 | 2263 | SWD | 61 | * | 40 | 1520 | 740 |  | 1.75 | Y |  | 380 | 660 | 95 | [6 empty] | * | * | * | 1320
2018 | 3 | 13 | 72 | 2264
2018 | 3 | 14 | 73 | 2265 | SWD | 61 | * | 40 | 1500 | 740 |  | 1.75 | Y |  | 380 | 660 | 95 | [6 empty] | * | * | * | 1320
2018 | 3 | 15 | 74 | 2266
2018 | 3 | 16 | 75 | 2267 | SWD | 61 | * | 40 | 1510 | 720 |  | 1.75 | Y |  | 380 | 660 | 95 | [6 empty] | * | * | * | 1320
2018 | 3 | 17 | 76 | 2268
2018 | 3 | 18 | 77 | 2269
2018 | 3 | 19 | 78 | 2270 | SWD | 61 | * | 40 | 1510 | 720 |  | 1.75 | Y |  | 380 | 660 | 100 | [6 empty] | * | * | * | 1320
2018 | 3 | 20 | 79 | 2271
2018 | 3 | 21 | 80 | 2272 | SWD | 61 | * | 40 | 1520 | 720 |  | 1.75 | Y |  | 380 | 660 | 95 | [6 empty] | * | * | * | 1320
2018 | 3 | 22 | 81 | 2273
2018 | 3 | 23 | 82 | 2274 | SWD/RN | 61 | * | 40 | 1530 | 690 |  | 1.75 | Y |  | 380 | 650 | 95 | [6 empty] | * | * | * | 1350
2018 | 3 | 24 | 83 | 2275
2018 | 3 | 25 | 84 | 2276 | SWD | 61 | * | 40 | 1530 | 690 |  | 1.75 | Y |  | 380 | 650 | 95 | [6 empty] | * | * | * | 1350
2018 | 3 | 26 | 85 | 2277
2018 | 3 | 27 | 86 | 2278
2018 | 3 | 28 | 87 | 2279 | NB | 61 | * | 40 | 1520 | 690 |  | 1.75 | Y |  | 380 | 650 | 100 | [6 empty] | * | * | * | 1340
2018 | 3 | 29 | 88 | 2280
2018 | 3 | 30 | 89 | 2281 | NB | 61 | * | 40 | 1510 | 690 |  | 1.75 | Y |  | 380 | 660 | 100 | [6 empty] | * | * | * | 1300
2018 | 3 | 31 | 90 | 2282
2018 | 4 | 1 | 91 | 2283 | NB | 61 | * | 40 | 1510 | 670 |  | 1.75 | Y |  | 380 | 660 | 100 | [6 empty] | * | * | * | 1300
2018 | 4 | 2 | 92 | 2284 | NB | 61 | * | 40 | 1510 | 680 |  | 1.75 | Y |  | 380 | 660 | 100 | [6 empty] | * | * | * | 1310
2018 | 4 | 3 | 93 | 2285
2018 | 4 | 4 | 94 | 2286
2018 | 4 | 5 | 95 | 2287 | NB | 61 | * | 40 | 1500 | 660 |  | 1.75 | Y |  | 370 | 650 | 100 | [6 empty] | * | * | * | 1260
2018 | 4 | 6 | 96 | 2288
2018 | 4 | 7 | 97 | 2289 | NB | 61 | * | 40 | 1500 | 650 |  | 1.75 | Y |  | 360 | 650 | 100 | [6 empty] | * | * | * | 1260
2018 | 4 | 8 | 98 | 2290
2018 | 4 | 9 | 99 | 2291 | NB | 61 | * | 40 | 1510 | 650 |  | 1.75 | Y |  | 370 | 650 | 100 | [6 empty] | * | * | * | 1310
2018 | 4 | 10 | 100 | 2292
2018 | 4 | 11 | 101 | 2293
2018 | 4 | 12 | 102 | 2294 | NB | 61 | * | 40 | 1510 | 650 |  | 1.75 | Y |  | 370 | 650 | 100 | [6 empty] | * | * | * | 1300
2018 | 4 | 13 | 103 | 2295
2018 | 4 | 14 | 104 | 2296 | NB | 61 | * | 40 | 1520 | 640 |  | 1.75 | Y |  | 370 | 650 | 100 | [6 empty] | * | * | * | 1350 | Added 1L DI water to H2 generator at 1322z.
2018 | 4 | 15 | 105 | 2297
2018 | 4 | 16 | 106 | 2298 | NB | 61 | * | 40 | 1510 | 640 |  | 1.75 | Y |  | 370 | 650 | 100 | [6 empty] | * | * | * | 1350
2018 | 4 | 17 | 107 | 2299
2018 | 4 | 18 | 108 | 2300 | NB | 61 | * | 40 | 1510 | 640 |  | 1.75 | Y |  | 360 | 650 | 100 | [6 empty] | * | * | * | 1350
2018 | 4 | 19 | 109 | 2301
2018 | 4 | 20 | 110 | 2302 | NB | 61 | * | 40 | 1510 | 630 |  | 1.75 | Y |  | 370 | 650 | 100 | [6 empty] | * | * | * | 1320
2018 | 4 | 21 | 111 | 2303
2018 | 4 | 22 | 112 | 2304 | NB | 61 | * | 40 | 1510 | 620 |  | 1.75 | Y |  | 360 | 640 | 100 | [6 empty] | * | * | * | 1310
2018 | 4 | 23 | 113 | 2305 | NB | 61 | * | 40 | 1510 | 620 |  | 1.75 | Y |  | 360 | 650 | 100 | [6 empty] | * | * | * | 1310
2018 | 4 | 24 | 114 | 2306
2018 | 4 | 25 | 115 | 2307
2018 | 4 | 26 | 116 | 2308 | NB | 61 | * | 40 | 1510 | 610 |  | 1.75 | Y |  | 360 | 640 | 98 | [6 empty] | * | * | * | 1280
2018 | 4 | 27 | 117 | 2309 | NB/MB | 61 | * | 40 | 1510 | 600 |  | 1.75 | Y |  | 370 | 630 | 100 | [6 empty] | * | * | * | 1290
2018 | 4 | 28 | 118 | 2310
2018 | 4 | 29 | 119 | 2311
2018 | 4 | 30 | 120 | 2312 | NB/MB | 61 | * | 40 | 1510 | 590 |  | 1.75 | Y |  | 360 | 630 | 85 | [6 empty] | * | * | * | 1260 | Trap housing purge flow increased to 98 ml/min when checked again ~1 hour later.
2018 | 5 | 1 | 121 | 2313 | MB | 61 | * | 40 | 1510 | 590 |  | 1.75 | Y |  | 360 | 630 | 98 | [6 empty] | * | * | * | 1280
2018 | 5 | 2 | 122 | 2314
2018 | 5 | 3 | 123 | 2315
2018 | 5 | 4 | 124 | 2316 | MB | 61 | * | 40 | 1510 | 580 |  | 1.75 | Y |  | 360 | 630 | 98 | [6 empty] | * | * | * | 1270 | Added 1L of DI water to H2 generator at 1248z.
2018 | 5 | 5 | 125 | 2317 | MB | 61 | * | 40 | 1510 | 570 |  | 1.75 | Y |  | 360 | 630 | 98 | [6 empty] | * | * | * | 1280
2018 | 5 | 6 | 126 | 2318
2018 | 5 | 7 | 127 | 2319 | MB | 61 | * | 40 | 1520 | 570 |  | 1.75 | Y |  | 360 | 630 | 98 | [6 empty] | * | * | * | 1250
2018 | 5 | 8 | 128 | 2320 | MB | 61 | * | 40 | 1520 | 570 |  | 1.7 | Y |  | 360 | 630 | 98 | [6 empty] | * | * | * | 1260
2018 | 5 | 9 | 129 | 2321
2018 | 5 | 10 | 130 | 2322 | MB | 61 | * | 40 | 1520 | 570 |  | 1.7 | Y |  | 360 | 630 | 98 | [6 empty] | * | * | * | 1260
2018 | 5 | 11 | 131 | 2323
2018 | 5 | 12 | 132 | 2324 | MB | 61 | * | 40 | 1520 | 570 |  | 1.75 | Y |  | 360 | 630 | 98 | [6 empty] | * | * | * | 1240
2018 | 5 | 13 | 133 | 2325
2018 | 5 | 14 | 134 | 2326 | MB | 61 | * | 40 | 1520 | 570 |  | 1.75 | Y |  | 360 | 630 | 98 | [6 empty] | * | * | * | 1260
2018 | 5 | 15 | 135 | 2327 | MB | 61 | * | 40 | 1520 | 570 |  | 1.75 | Y |  | 360 | 630 | 98 | [6 empty] | * | * | * | 1700 | Replaced H2 generator desiccant cartridge 1330-1407z. H2 generator acutal pressure 61 psi, flow 20 mL/min (1/5 squares). Flow increased to 40 mL/min (2/5 squares) in the following hour.
2018 | 5 | 16 | 136 | 2328
2018 | 5 | 17 | 137 | 2329 | MB | 61 | * | 40 | 1520 | 560 |  | 1.75 | Y |  | 360 | 630 | 98 | [6 empty] | * | * | * | 1690
2018 | 5 | 18 | 138 | 2330
2018 | 5 | 19 | 139 | 2331 | MB | 61 | * | 40 | 1520 | 560 |  | 1.75 | Y |  | 360 | 630 | 98 | [6 empty] | * | * | * | 1680 | Added 1L of DI water to H2 generator at 1226z.
2018 | 5 | 20 | 140 | 2332
2018 | 5 | 21 | 141 | 2333 | MB | 61 | * | 40 | 1520 | 560 |  | 1.75 | Y |  | 360 | 630 | 98 | [6 empty] | * | * | * | 1670
2018 | 5 | 22 | 142 | 2334 | MB | 61 | * | 40 | 1520 | 560 |  | 1.75 | Y |  | 360 | 630 | 98 | [6 empty] | * | * | * | 1650
2018 | 5 | 23 | 143 | 2335 | MB | 61 | * | 40 | 1520 | 560 |  | 1.75 | Y |  | 360 | 630 | 98 | [6 empty] | * | * | * | 1650
2018 | 5 | 24 | 144 | 2336 | MB | 61 | * | 40 | 1520 | 540 |  | 1.75 | Y |  | 360 | 630 | 98 | [6 empty] | * | * | * | 1650
2018 | 5 | 25 | 145 | 2337
2018 | 5 | 26 | 146 | 2338
2018 | 5 | 27 | 147 | 2339
2018 | 5 | 28 | 148 | 2340 | MB | 61 | * | 40 | 1520 | 540 |  | 1.7 | Y |  | 360 | 630 | 98 | [6 empty] | * | * | * | 1630
2018 | 5 | 29 | 149 | 2341 | MB | 61 | * | 40 | 1520 | 540 |  | 1.75 | Y |  | 360 | 630 | 90 | [6 empty] | * | * | * | 1610
2018 | 5 | 30 | 150 | 2342 | MB | 61 | * | 40 | 1520 | 540 |  | 1.75 | Y |  | 360 | 630 | 98 | [6 empty] | * | * | * | 1610
2018 | 5 | 31 | 151 | 2343
2018 | 6 | 1 | 152 | 2344 | MB | 61 | * | 40 | 1520 | 530 |  | 1.75 | Y |  | 360 | 620 | 98 | [6 empty] | * | * | * | 1600
2018 | 6 | 2 | 153 | 2345
2018 | 6 | 3 | 154 | 2346 | MB | 61 | * | 40 | 1520 | 530 |  | 1.75 | Y |  | 360 | 620 | 98 | [6 empty] | * | * | * | 1600
2018 | 6 | 4 | 155 | 2347
2018 | 6 | 5 | 156 | 2348 | MB/MM | 61 | * | 40 | 1520 | 530 |  | 1.75 | Y |  | 360 | 620 | 100 | [6 empty] | * | * | * | 1600
2018 | 6 | 6 | 157 | 2349 | MB/MM | 61 | * | 60 | 1520 | 500 |  | 1.75 | Y |  | 350 | 610 | 100 | [6 empty] | * | * | * | 1600
2018 | 6 | 7 | 158 | 2350 | MB/MM | 61 | * | 40 | 1520 | 500 |  | 1.75 | Y |  | 350 | 600 | 100 | [6 empty] | * | * | * | 1600 | Added 1L of DI water to H2 generator at 1213z.
2018 | 6 | 8 | 159 | 2351 | MM | 61 | * | 40 | 1520 | 500 |  | 1.75 | Y |  | 350 | 610 | 100 | [6 empty] | * | * | * | 1600
2018 | 6 | 9 | 160 | 2352 | MM | 61 | * | 40 | 1520 | 500 |  | 1.75 | Y |  | 350 | 610 | 100 | [6 empty] | * | * | * | 1600
2018 | 6 | 10 | 161 | 2353
2018 | 6 | 11 | 162 | 2354 | MM | 61 | * | 40 | 1520 | 500 |  | 1.75 | Y |  | 350 | 610 | 100 | [6 empty] | * | * | * | 1600 | Removed H2 Generator button panel at approximately 1810z
2018 | 6 | 12 | 163 | 2355 | MM | 62 | * | 40 | 1520 | 500 |  | 1.75 | Y |  | 350 | 610 | 100 | [6 empty] | * | * | * | 1590
2018 | 6 | 13 | 164 | 2356
2018 | 6 | 14 | 165 | 2357
2018 | 6 | 15 | 166 | 2358 | MM | 63 | * | 40 | 1510 | 500 |  | 1.75 | Y |  | 350 | 610 | 105 | [6 empty] | * | * | * | 1550 | Replaced the H2 Generator button panel at approximately 1730z. The GC was remotely powered down during a planned power outage and resumed operation at approximately 1930z
2018 | 6 | 16 | 167 | 2359 | MM | 63 | * | 40 | 1510 | 500 |  | 1.75 | Y |  | 350 | 600 | 105 | [6 empty] | * | * | * | 1550
2018 | 6 | 17 | 168 | 2360
2018 | 6 | 18 | 169 | 2361 | [23 empty] | The CU GC ran on the UPS during two brief planned power outages from 0115z-0130z and 0645z-0650z
2018 | 6 | 19 | 170 | 2362 | MM | 62 | * | 40 | 1510 | 500 |  | 1.75 | Y |  | 350 | 600 | 100 | [6 empty] | * | * | * | 1550
2018 | 6 | 20 | 171 | 2363 | MM | 62 | * | 40 | 1510 | 470 |  | 1.75 | Y |  | 350 | 600 | 100 | [6 empty] | * | * | * | 1525
2018 | 6 | 21 | 172 | 2364
2018 | 6 | 22 | 173 | 2365 | MM | 62 | * | 40 | 1510 | 450 |  | 1.75 | Y |  | 350 | 600 | 100 | [6 empty] | * | * | * | 1510
2018 | 6 | 23 | 174 | 2366
2018 | 6 | 24 | 175 | 2367
2018 | 6 | 25 | 176 | 2368 | MM | 62 | * | 40 | 1510 | 450 |  | 1.75 | Y |  | 350 | 600 | 100 | [6 empty] | * | * | * | 1510
2018 | 6 | 26 | 177 | 2369 | MM | 62 | * | 40 | 1510 | 450 |  | 1.75 | Y |  | 350 | 600 | 100 | [6 empty] | * | * | * | 1500
2018 | 6 | 27 | 178 | 2370 | MM | 62 | * | 40 | 1510 | 450 |  | 1.75 | Y |  | 350 | 600 | 100 | [6 empty] | * | * | * | 1500
2018 | 6 | 28 | 179 | 2371 | MM | 62 | * | 40 | 1510 | 450 |  | 1.75 | Y |  | 350 | 600 | 100 | [6 empty] | * | * | * | 1500 | Added 1.5L of DI water to H2 generator at 1540
2018 | 6 | 29 | 180 | 2372
2018 | 6 | 30 | 181 | 2373 | MM | 62 | * | 40 | 1510 | 450 |  | 1.75 | Y |  | 350 | 600 | 100 | [6 empty] | * | * | * | 1500
2018 | 7 | 1 | 182 | 2374
2018 | 7 | 2 | 183 | 2375
2018 | 7 | 3 | 184 | 2376 | MM | 62 | * | 40 | 1510 | 450 |  | 1.75 | Y |  | 350 | 600 | 100 | [6 empty] | * | * | * | 1500
2018 | 7 | 4 | 185 | 2377
2018 | 7 | 5 | 186 | 2378 | MM | 62 | * | 40 | 1510 | 450 |  | 1.75 | Y |  | 350 | 600 | 100 | [6 empty] | * | * | * | 1450
2018 | 7 | 6 | 187 | 2379
2018 | 7 | 7 | 188 | 2380
2018 | 7 | 8 | 189 | 2381
2018 | 7 | 9 | 190 | 2382
2018 | 7 | 10 | 191 | 2383 | MM | 62 | * | 40 | 1510 | 450 |  | 1.75 | Y |  | 350 | 600 | 100 | [6 empty] | * | * | * | 1450
2018 | 7 | 11 | 192 | 2384 | MM | 62 | * | 40 | 1510 | 450 |  | 1.75 | Y |  | 350 | 600 | 100 | [6 empty] | * | * | * | 1450
2018 | 7 | 12 | 193 | 2385
2018 | 7 | 13 | 194 | 2386
2018 | 7 | 14 | 195 | 2387
2018 | 7 | 15 | 196 | 2388
2018 | 7 | 16 | 197 | 2389 | MM | 62 | * | 20 | 1510 | 400 |  | 1.75 | Y |  | 350 | 600 | 100 | [6 empty] | * | * | * | 1450 | Removed rime from inlet at 2230, flow returned to 2/4 bars on hydrogen generator (40 lpm)
2018 | 7 | 17 | 198 | 2390
2018 | 7 | 18 | 199 | 2391 | MM | 62 | * | 40 | 1510 | 400 |  | 1.75 | Y |  | 350 | 600 | 100 | [6 empty] | * | * | * | 1400
2018 | 7 | 19 | 200 | 2392
2018 | 7 | 20 | 201 | 2393
2018 | 7 | 21 | 202 | 2394 | MM | 62 | * | 40 | 1510 | 400 |  | 1.75 | Y |  | 310 | 600 | 100 | [6 empty] | * | * | * | 1390
2018 | 7 | 22 | 203 | 2395
2018 | 7 | 23 | 204 | 2396 | [23 empty] | CUGC ran on UPS system during a planned power outage from 2153-2300z and again from 2307-2315z.
2018 | 7 | 24 | 205 | 2397 | MM | 62 | * | 40 | 1510 | 400 |  | 1.75 | Y |  | 310 | 600 | 90 | [6 empty] | * | * | * |  | Added 50ml of water at 21:35z to the H2 generator
2018 | 7 | 25 | 206 | 2398 | MM | 62 | * | 40 | 1510 | 350 |  | 1.75 | Y |  | 310 | 600 | 105 | [6 empty] | * | * | * | 1390
2018 | 7 | 26 | 207 | 2399
2018 | 7 | 27 | 208 | 2400 | MM | 62 | * | 40 | 1510 | 350 |  | 1.75 | Y |  | 310 | 600 | 105 | [6 empty] | * | * | * | 1350
2018 | 7 | 28 | 209 | 2401 | [23 empty] | TAWO planned power outage from 1218-1228z; Used simple green and rubbing alcohol in the TAWO vestibule
2018 | 7 | 29 | 210 | 2402 | MM | 62 | * | 40 | 1510 | 350 |  | 1.75 | Y |  | 310 | 580 | 105 | [6 empty] | * | * | * | 1350 | Added 1litre of DI water to H2 generator
2018 | 7 | 30 | 211 | 2403 | MM | 62 | * | 40 | 1510 | 350 |  | 1.75 | Y |  | 310 | 580 | 105 | [6 empty] | * | * | * | 1350
2018 | 7 | 31 | 212 | 2404
2018 | 8 | 1 | 213 | 2405
2018 | 8 | 2 | 214 | 2406 | MM | 62 | * | 40 | 1510 | 350 |  | 1.75 | Y |  | 310 | 580 | 100 | [6 empty] | * | * | * | 1350 | CUGC in standby mode during power outages
2018 | 8 | 3 | 215 | 2407 | MM | 62 | * | 40 | 1510 | 350 |  | 1.75 | Y |  | 310 | 575 | 105 | [6 empty] | * | * | * | 1325 | CUGC in standby mode during power outages
2018 | 8 | 4 | 216 | 2408 | [23 empty] | CUGC remotely shutdown during power outages
2018 | 8 | 5 | 217 | 2409
2018 | 8 | 6 | 218 | 2410
2018 | 8 | 7 | 219 | 2411 | MM | 62 | * | 40 | 1510 | 350 |  | 0 | N |  | 310 | 575 | 0 | [6 empty] | * | * | * | 1310 | CUGC remotely shutdown during power outages
2018 | 8 | 8 | 220 | 2412
2018 | 8 | 9 | 221 | 2413 | MM | 62 | * | 40 | 1510 | 350 |  | 1.75 | Y |  | 310 | 575 | 100 | [6 empty] | * | * | * | 1310
2018 | 8 | 10 | 222 | 2414
2018 | 8 | 11 | 223 | 2415
2018 | 8 | 12 | 224 | 2416
2018 | 8 | 13 | 225 | 2417 | MM | 62 | * | 40 | 1510 | 340 |  | 1.75 | Y |  | 310 | 575 | 100 | [6 empty] | * | * | * | 1300
2018 | 8 | 14 | 226 | 2418
2018 | 8 | 15 | 227 | 2419
2018 | 8 | 16 | 228 | 2420
2018 | 8 | 17 | 229 | 2421 | MM | 62 | * | 40 | 1510 | 340 |  | 1.75 | Y |  | 315 | 575 | 100 | [6 empty] | * | * | * | 1300
2018 | 8 | 18 | 230 | 2422
2018 | 8 | 19 | 231 | 2423 | MM | 62 | * | 40 | 1510 | 340 |  | 1.75 | Y |  | 315 | 575 | 100 | [6 empty] | * | * | * | 1300 | Added 1L of DI water to H2 generator at 2220z.
2018 | 8 | 20 | 232 | 2424
2018 | 8 | 21 | 233 | 2425 | MM | 62 | * | 40 | 1510 | 320 |  | 2 | Y |  | 310 | 590 | 90 | [6 empty] | * | * | * | 1300
2018 | 8 | 22 | 234 | 2426
2018 | 8 | 23 | 235 | 2427
2018 | 8 | 24 | 236 | 2428
2018 | 8 | 25 | 237 | 2429 | MM | 62 | * | 40 | 1510 | 250 |  | 1.75 | Y |  | 310 | 590 | 100 | [6 empty] | * | * | * | 1300 | Jacques onsite.
2018 | 8 | 26 | 238 | 2430 | [23 empty] | Jacques onsite. Climbed TAWO inlet tower at 1740 to change CUGC inlet filter
2018 | 8 | 27 | 239 | 2431 | [23 empty] | Jacques onsite.
2018 | 8 | 28 | 240 | 2432 | [23 empty] | Jacques onsite.
2018 | 8 | 29 | 241 | 2433
2018 | 8 | 30 | 242 | 2434
2018 | 8 | 31 | 243 | 2435 | MM | 62 | * | 40 | 1510 | 290 |  | 1.9 | Y |  | 310 | 590 | 115 | [6 empty] | * | * | * | 1700
2018 | 9 | 1 | 244 | 2436
2018 | 9 | 2 | 245 | 2437
2018 | 9 | 3 | 246 | 2438
2018 | 9 | 4 | 247 | 2439 | MM | 62 | * | 40 | 1510 | 290 |  | 2 | Y |  | 310 | 590 | 115 | [6 empty] | * | * | * | 1670
2018 | 9 | 5 | 248 | 2440 | MM | 62 | * | 40 | 1510 | 290 |  | 2 | Y |  | 310 | 560 | 120 | [6 empty] | * | * | * | 1690 | Added 1L of DI water to H2 generator at 1700z.
2018 | 9 | 6 | 249 | 2441
2018 | 9 | 7 | 250 | 2442
2018 | 9 | 8 | 251 | 2443
2018 | 9 | 9 | 252 | 2444 | MM | 62 | * | 40 | 1510 | 290 |  | 2 | Y |  | 310 | 560 | 120 | [6 empty] | * | * | * | 1640
2018 | 9 | 10 | 253 | 2445
2018 | 9 | 11 | 254 | 2446
2018 | 9 | 12 | 255 | 2447 | MM | 62 | * | 40 | 1510 | 290 |  | 2 | Y |  | 310 | 560 | 120 | [6 empty] | * | * | * | 1640
2018 | 9 | 13 | 256 | 2448
2018 | 9 | 14 | 257 | 2449
2018 | 9 | 15 | 258 | 2450 | MM | 62 | * | 40 | 1510 | 290 |  | 2 | Y |  | 310 | 560 | 105 | [6 empty] | * | * | * | 1640
2018 | 9 | 16 | 259 | 2451
2018 | 9 | 17 | 260 | 2452
2018 | 9 | 18 | 261 | 2453 | MM | 62 | * | 40 | 1510 | 275 |  | 2 | Y |  | 310 | 560 | 115 | [6 empty] | * | * | * | 1590
2018 | 9 | 19 | 262 | 2454
2018 | 9 | 20 | 263 | 2455
2018 | 9 | 21 | 264 | 2456 | MM | 62 | * | 40 | 1510 | 275 |  | 2 | Y |  | 310 | 560 | 118 | [6 empty] | * | * | * | 1570
2018 | 9 | 22 | 265 | 2457
2018 | 9 | 23 | 266 | 2458
2018 | 9 | 24 | 267 | 2459 | MM | 62 | * | 40 | 1510 | 250 |  | 2 | Y |  | 310 | 580 | 118 | [6 empty] | * | * | * | 1550
2018 | 9 | 25 | 268 | 2460
2018 | 9 | 26 | 269 | 2461
2018 | 9 | 27 | 270 | 2462 | MM | 62 | * | 40 | 1510 | 245 |  | 2 | Y |  | 310 | 575 | 113 | [6 empty] | * | * | * | 1540 | Added 1L of DI water to H2 generator at 2005z.
2018 | 9 | 28 | 271 | 2463 | MM | 62 | * | 40 | 1510 | 245 |  | 2 | Y |  | 310 | 560 | 118 | [6 empty] | * | * | * | 1530
2018 | 9 | 29 | 272 | 2464
2018 | 9 | 30 | 273 | 2465
2018 | 10 | 1 | 274 | 2466
2018 | 10 | 2 | 275 | 2467 | MM | 62 | * | 40 | 1510 | 200 |  | 2 | Y |  | 300 | 560 | 118 | [6 empty] | * | * | * | 1510
2018 | 10 | 3 | 276 | 2468
2018 | 10 | 4 | 277 | 2469
2018 | 10 | 5 | 278 | 2470 | MM | 62 | * | 40 | 1510 | 200 |  | 2 | Y |  | 300 | 560 | 115 | [6 empty] | * | * | * | 1500
2018 | 10 | 6 | 279 | 2471
2018 | 10 | 7 | 280 | 2472
2018 | 10 | 8 | 281 | 2473 | MM | 62 | * | 40 | 1510 | 200 |  | 2 | Y |  | 295 | 510 | 113 | [6 empty] | * | * | * | 1500
2018 | 10 | 9 | 282 | 2474
2018 | 10 | 10 | 283 | 2475 | [23 empty] | Replaced O3 Analyzer inlet filter on TAWO mast from 1631-1645z.
2018 | 10 | 11 | 284 | 2476
2018 | 10 | 12 | 285 | 2477 | MM | 62 | * | 40 | 1500 | 175 |  | 2 | Y |  | 300 | 550 | 118 | [6 empty] | * | * | * | 1475 | Added 1L of DI water to H2 generator at 1128z.
2018 | 10 | 13 | 286 | 2478
2018 | 10 | 14 | 287 | 2479
2018 | 10 | 15 | 288 | 2480 | MM | 62 | * | 40 | 1510 | 160 |  | 2 | Y |  | 300 | 550 | 118 | [6 empty] | * | * | * | 1490
2018 | 10 | 16 | 289 | 2481
2018 | 10 | 17 | 290 | 2482
2018 | 10 | 18 | 291 | 2483 | MM/TN | 61 | * | 40 | 1515 | 150 |  | 2 | Y |  | 295 | 550 | 110 | [6 empty] | * | * | * | 1660 | Changed the H2 gen desiccant cartridge 1617-1714z. (1632 powered H2 generator back on, 1643 began self-test)
2018 | 10 | 19 | 292 | 2484 | MM/TN
2018 | 10 | 20 | 293 | 2485 | 0
2018 | 10 | 21 | 294 | 2486 | TN | 61 | * | 40 | 1510 | 150 |  | 2 | Y |  | 295 | 545 | 100 | [6 empty] | * | * | * | 1620 | Added 1L of DI water to H2 generator at 1655z
2018 | 10 | 22 | 295 | 2487 | TN | 61 | * | 40 | 1505 | 150 |  | 2 | Y |  | 295 | 545 | 110 | [6 empty] | * | * | * | 1620
2018 | 10 | 23 | 296 | 2488 | TN | 61 | * | 40 | 1500 | 150 |  | 2 | Y |  | 300 | 545 | 115 | [6 empty] | * | * | * | 1650
2018 | 10 | 24 | 297 | 2489 | TN | 61 | * | 40 | 1500 | 150 |  | 2 | Y |  | 295 | 550 | 110 | [6 empty] | * | * | * | 1645
2018 | 10 | 25 | 298 | 2490 | TN | 61 | * | 40 | 1505 | 145 |  | 2 | Y |  | 295 | 545 | 115 | [6 empty] | * | * | * | 1625
2018 | 10 | 26 | 299 | 2491 | TN | 61 | * | 40 | 1500 | 150 |  | 2 | Y |  | 295 | 545 | 115 | [6 empty] | * | * | * | 1620
2018 | 10 | 27 | 300 | 2492 | TN | 61 | * | 40 | 1500 | 145 |  | 2 | Y |  | 295 | 545 | 110 | [6 empty] | * | * | * | 1600
2018 | 10 | 28 | 301 | 2493 | TN | 61 | * | 40 | 1505 | 145 |  | 2 | Y |  | 295 | 545 | 110 | [6 empty] | * | * | * | 1620
2018 | 10 | 29 | 302 | 2494 | TN | 61 | * | 40 | 1505 | 145 |  | 2 | Y |  | 295 | 545 | 110 | [6 empty] | * | * | * | 1620
2018 | 10 | 30 | 303 | 2495 | TN | 61 | * | 40 | 1510 | 140 |  | 2 | Y |  | 295 | 540 | 110 | [6 empty] | * | * | * | 1610
2018 | 10 | 31 | 304 | 2496 | TN | 61 | * | 40 | 1510 | 140 |  | 2 | Y |  | 295 | 545 | 110 | [6 empty] | * | * | * | 1600
2018 | 11 | 1 | 305 | 2497 | TN | 61 | * | 40 | 1500 | 140 |  | 2 | Y |  | 295 | 540 | 110 | [6 empty] | * | * | * | 1590
2018 | 11 | 2 | 306 | 2498 | TN | 61 | * | 40 | 1505 | 125 |  | 2 | Y |  | 295 | 540 | 110 | [6 empty] | * | * | * | 1600
2018 | 11 | 3 | 307 | 2499 | TN | 61 | * | 40 | 1505 | 120 |  | 2 | Y |  | 295 | 545 | 110 | [6 empty] | * | * | * | 1595 | H2 gen briefly dropped to 20 ml/min flow.
2018 | 11 | 4 | 308 | 2500 | TN | 61 | * | 40 | 1505 | 120 |  | 2 | Y |  | 295 | 540 | 110 | [6 empty] | * | * | * | 1580
2018 | 11 | 5 | 309 | 2501 | TN | 61 | * | 40 | 1505 | 115 |  | 2 | Y |  | 295 | 540 | 110 | [6 empty] | * | * | * | 1580
2018 | 11 | 6 | 310 | 2502 | TN | 61 | * | 40 | 1500 | 115 |  | 2 | Y |  | 295 | 540 | 110 | [6 empty] | * | * | * | 1580
2018 | 11 | 7 | 311 | 2503 | TN | 61 | * | 40 | 1500 | 110 |  | 2 | Y |  | 295 | 540 | 110 | [6 empty] | * | * | * | 1575
2018 | 11 | 8 | 312 | 2504 | TN | 61 | * | 40 | 1500 | 110 |  | 2 | Y |  | 295 | 540 | 110 | [6 empty] | * | * | * | 1580
2018 | 11 | 9 | 313 | 2505 | TN | 61 | * | 40 | 1500 | 110 |  | 2 | Y |  | 290 | 530 | 110 | [6 empty] | * | * | * | 1590
2018 | 11 | 10 | 314 | 2506 | TN | 61 | * | 40 | 1500 | 110 |  | 2 | Y |  | 295 | 540 | 110 | [6 empty] | * | * | * | 1590
2018 | 11 | 11 | 315 | 2507 | TN | 61 | * | 40 | 1500 | 110 |  | 2 | Y |  | 290 | 545 | 110 | [6 empty] | * | * | * | 1575
2018 | 11 | 12 | 316 | 2508 | TN | 61 | * | 40 | 1510 | 110 |  | 2 | Y |  | 295 | 545 | 110 | [6 empty] | * | * | * | 1575
2018 | 11 | 13 | 317 | 2509 | TN | 61 | * | 40 | 1505 | 110 |  | 2 | Y |  | 295 | 540 | 110 | [6 empty] | * | * | * | 1575 | Added 1L of DI water to H2 generator at 1351z
2018 | 11 | 14 | 318 | 2510 | TN | 61 | * | 40 | 1510 | 110 |  | 2 | Y |  | 290 | 540 | 110 | [6 empty] | * | * | * | 1575 | H2 gen flow briefly rose to 60%
2018 | 11 | 15 | 319 | 2511 | TN | 61 | * | 40 | 1515 | 110 |  | 2 | Y |  | 295 | 540 | 110 | [6 empty] | * | * | * | 1570 | Switched CH4 from Std 2 to Std1 tank at 1340 Z
2018 | 11 | 16 | 320 | 2512 | TN | 61 | * | 40 | 1515 | 110 |  | 2 | Y |  | 290 | 535 | 110 | [6 empty] | * | * | * | 1560 | After PI consultation, re-opened valve to Std2. Installed Teamviewer 14 on GC computer. Changed Std tank view GC computer with guidance from PI.
2018 | 11 | 17 | 321 | 2513 | TN | 61 | * | 40 | 1510 | 110 |  | 2 | Y |  | 290 | 530 | 110 | [6 empty] | * | * | * | 1550
2018 | 11 | 18 | 322 | 2514 | TN | 61 | * | 40 | 1510 | 90 |  | 2 | Y |  | 290 | 540 | 107 | [6 empty] | * | * | * | 1550
2018 | 11 | 19 | 323 | 2515 | TN | 61 | * | 40 | 1510 | 90 |  | 2 | Y |  | 295 | 540 | 107 | [6 empty] | * | * | * | 1550
2018 | 11 | 20 | 324 | 2516 | TN | 61 | * | 40 | 1500 | 90 |  | 2 | Y |  | 290 | 525 | 107 | [6 empty] | * | * | * | 1550
2018 | 11 | 21 | 325 | 2517 | TN | 61 | * | 40 | 1490 | 90 |  | 2 | Y |  | 290 | 525 | 107 | [6 empty] | * | * | * | 1550
2018 | 11 | 22 | 326 | 2518 | TN | 61 | * | 40 | 1500 | 90 |  | 2 | Y |  | 290 | 530 | 107 | [6 empty] | * | * | * | 1510
2018 | 11 | 23 | 327 | 2519 | TN | 61 | * | 40 | 1500 | 90 |  | 2 | Y |  | 295 | 525 | 107 | [6 empty] | * | * | * | 1510
2018 | 11 | 24 | 328 | 2520 | TN | 61 | * | 40 | 1500 | 90 |  | 2 | Y |  | 295 | 525 | 110 | [6 empty] | * | * | * | 1520
2018 | 11 | 25 | 329 | 2521 | TN | 61 | * | 40 | 1505 | 90 |  | 2 | Y |  | 295 | 540 | 110 | [6 empty] | * | * | * | 1515
2018 | 11 | 26 | 330 | 2522 | TN | 1 | * | 0 | 1505 | 90 |  | 2 | Y |  | 280 | 530 | 108 | [6 empty] | * | * | * | 1510 | alarm sounding on H2 generator. Float champer lock. Awaiting instructions from PI.
2018 | 11 | 27 | 331 | 2523 | TN | x | * | x | 1505 | 90 |  | 2 | Y |  | 275 | 525 | 108 | [6 empty] | * | * | * | 1645 | backup cylinder in use; H2 generator is off.
2018 | 11 | 28 | 332 | 2524 | TN | x | * | x | 1500 | 90 |  | 2 | y |  | 270 | 520 | 107 | [6 empty] | * | * | * | 1600 | backup cylinder in use; H2 generator is off.
2018 | 11 | 29 | 333 | 2525 | TN | 61 | * | 40 | 1495 | 90 |  | 2 | Y |  | 260 | 505 | 106 | [6 empty] | * | * | * | 1600 | H2 generator back online as of 1640Z
2018 | 11 | 30 | 334 | 2526 | TN | 61 | * | 40 | 1500 | 90 |  | 2 | Y |  | 260 | 520 | 106 | [6 empty] | * | * | * | 1600
2018 | 12 | 1 | 335 | 2527 | TN | 61 | * | 40 | 1495 | 90 |  | 2 | Y |  | 270 | 520 | 105 | [6 empty] | * | * | * | 1580 | added 1 L of DI water to H2 generator reservoir.
2018 | 12 | 2 | 336 | 2528 | TN | 61 | * | 40 | 1500 | 90 |  | 2 | Y |  | 280 | 530 | 108 | [6 empty] | * | * | * | 1590
2018 | 12 | 3 | 337 | 2529 | TN | 61 | * | 40 | 1505 | 90 |  | 2 | Y |  | 280 | 530 | 108 | [6 empty] | * | * | * | 1595
2018 | 12 | 4 | 338 | 2530 | TN | 61 | * | 40 | 1520 | 90 |  | 2 | Y |  | 280 | 530 | 108 | [6 empty] | * | * | * | 1610
2018 | 12 | 5 | 339 | 2531 | TN | 61 | * | 40 | 1490 | 90 |  | 2 | Y |  | 280 | 525 | 107 | [6 empty] | * | * | * | 1560
2018 | 12 | 6 | 340 | 2532 | TN | 61 | * | 40 | 1510 | 90 |  | 2 | Y |  | 275 | 525 | 110 | [6 empty] | * | * | * | 1580
2018 | 12 | 7 | 341 | 2533 | TN | 61 | * | 40 | 1510 | 90 |  | 2 | Y |  | 280 | 525 | 110 | [6 empty] | * | * | * | 1590
2018 | 12 | 8 | 342 | 2534 | TN | 61 | * | 40 | 1510 | 90 |  | 2 | y |  | 270 | 525 | 110 | [6 empty] | * | * | * | 1600
2018 | 12 | 9 | 343 | 2535 | TN | 61 | * | 20 | 1510 | 90 |  | 2 | Y |  | 270 | 520 | 108 | [6 empty] | * | * | * | 1560 | Hydrogen gen flow down to 20% on bar graph (35 lpm actual flow).
2018 | 12 | 10 | 344 | 2536 | TN | 61 | * | 40 | 1505 | 90 |  | 2 | Y |  | 270 | 520 | 108 | [6 empty] | * | * | * | 1550
2018 | 12 | 11 | 345 | 2537 | TN | 61 | * | 40 | 1510 | 90 |  | 2 | Y |  | 270 | 525 | 111 | [6 empty] | * | * | * | 1560 | added .5 L DI water to H2 generator
2018 | 12 | 12 | 346 | 2538 | TN | 61 | * | 40 | 1480 | 90 |  | 2 | Y |  | 270 | 520 | 105 | [6 empty] | * | * | * | 1550
2018 | 12 | 13 | 347 | 2539 | TN | 61 | * | 40 | 1500 | 85 |  | 2 | Y |  | 270 | 525 | 106 | [6 empty] | * | * | * | 1555
2018 | 12 | 14 | 348 | 2540 | TN | 61 | * | 40 | 1500 | 85 |  | 2 | Y |  | 280 | 525 | 106 | [6 empty] | * | * | * | 1560
2018 | 12 | 15 | 349 | 2541 | TN | 61 | * | 40 | 1500 | 85 |  | 2 | Y |  | 275 | 525 | 107 | [6 empty] | * | * | * | 1550
2018 | 12 | 16 | 350 | 2542 | TN | 61 | * | 40 | 1495 | 85 |  | 2 | Y |  | 275 | 525 | 107 | [6 empty] | * | * | * | 1550
2018 | 12 | 17 | 351 | 2543 | TN | 61 | * | 40 | 1495 | 85 |  | 2 | Y |  | 275 | 520 | 105 | [6 empty] | * | * | * | 1545
2018 | 12 | 18 | 352 | 2544 | TN | 61 | * | 40 | 1490 | 85 |  | 2 | Y |  | 275 | 525 | 106 | [6 empty] | * | * | * | 1550
2018 | 12 | 19 | 353 | 2545 | TN | 61 | * | 40 | 1490 | 85 |  | 2 | Y |  | 270 | 530 | 106 | [6 empty] | * | * | * | 1550
2018 | 12 | 20 | 354 | 2546 | TN | 61 | * | 40 | 1495 | 85 |  | 2 | Y |  | 275 | 525 | 107 | [6 empty] | * | * | * | 1550
2018 | 12 | 21 | 355 | 2547 | TN | 61 | * | 40 | 1495 | 85 |  | 2 | Y |  | 275 | 520 | 107 | [6 empty] | * | * | * | 1540
2018 | 12 | 22 | 356 | 2548
2018 | 12 | 23 | 357 | 2549
2018 | 12 | 24 | 358 | 2550 | TN | 61 | * | 40 | 1490 | 85 |  | 2 | Y |  | 275 | 525 | 106 | [6 empty] | * | * | * | 1500
2018 | 12 | 25 | 359 | 2551 | TN | 61 | * | 40 | 1490 | 85 |  | 2 | Y |  | 270 | 525 | 105 | [6 empty] | * | * | * | 1505
2018 | 12 | 26 | 360 | 2552 | TN | 61 | * | 40 | 1490 | 85 |  | 2 | Y |  | 270 | 520 | 106 | [6 empty] | * | * | * | 1520
2018 | 12 | 27 | 361 | 2553 | TN | 61 | * | 40 | 1490 | 85 |  | 2 | Y |  | 270 | 520 | 107 | [6 empty] | * | * | * | 1510
2018 | 12 | 28 | 362 | 2554
2018 | 12 | 29 | 363 | 2555 | TN | 61 | * | 40 | 1490 | 85 |  | 2 | y |  | 260 | 515 | 107 | [6 empty] | * | * | * | 1495
2018 | 12 | 30 | 364 | 2556
2018 | 12 | 31 | 365 | 2557 | TN | 61 | * | 40 | 1490 | 85 |  | 2 | Y |  | 270 | 520 | 107 | [6 empty] | * | * | * | 1470 | added 0.5 l DI water to H2 generator
2019 | 1 | 1 | 1 | 2558 | TN | 61 | * | 40 | 1480 | 85 |  | 2 | Y |  | 270 | 520 | 106 | [6 empty] | * | * | * | 1470
2019 | 1 | 2 | 2 | 2559
2019 | 1 | 3 | 3 | 2560 | TN | 61 | * | 40 | 1460 | 85 |  | 2 | Y |  | 260 | 510 | 106 | [6 empty] | * | * | * | 1460
2019 | 1 | 4 | 4 | 2561 | TN | 61 | * | 40 | 1470 | 85 |  | 2 | Y |  | 260 | 510 | 106 | [6 empty] | * | * | * | 1465
2019 | 1 | 5 | 5 | 2562
2019 | 1 | 6 | 6 | 2563
2019 | 1 | 7 | 7 | 2564 | TN | 61 | * | 40 | 1450 | 85 |  | 2 | Y |  | 260 | 510 | 106 | [6 empty] | * | * | * | 1445 | Added 1 l DI water to H2 generator
2019 | 1 | 8 | 8 | 2565 | TN | 61 | * | 40 | 1420 | 85 |  | 2 | Y |  | 255 | 510 | 105 | [6 empty] | * | * | * | 1440
2019 | 1 | 9 | 9 | 2566
2019 | 1 | 10 | 10 | 2567 | TN | 61 | * | 40 | 1440 | 85 |  | 2 | Y |  | 250 | 500 | 105 | [6 empty] | * | * | * | 1375
2019 | 1 | 11 | 11 | 2568 | TN | 61 | * | 40 | 1420 | 85 |  | 2 | Y |  | 250 | 500 | 104 | [6 empty] | * | * | * | 1395
2019 | 1 | 12 | 12 | 2569 | TN | 61 | * | 40 | 1420 | 85 |  | 2 | Y |  | 250 | 500 | 104 | [6 empty] | * | * | * | 1400
2019 | 1 | 13 | 13 | 2570
2019 | 1 | 14 | 14 | 2571 | TN | 61 | * | 40 | 1410 | 85 |  | 2 | Y |  | 250 | 505 | 104 | [6 empty] | * | * | * | 1400
2019 | 1 | 15 | 15 | 2572 | TN | 61 | * | 40 | 1400 | 85 |  | 2 | Y |  | 250 | 500 | 103 | [6 empty] | * | * | * | 1400
2019 | 1 | 16 | 16 | 2573
2019 | 1 | 17 | 17 | 2574 | TN | 61 | * | 40 | 1405 | 85 |  | 2 | Y |  | 245 | 505 | 104 | [6 empty] | * | * | * | 1390 | H2 generator has float lock issue. Changed over to H2 cylindar at 1802 Z
2019 | 1 | 18 | 18 | 2575 | TN | * |  | * | 1405 | 85 |  | 2 | Y |  | 240 | 510 | 104 | [6 empty] | * | * | * | 1560 | H2 generator down, still on cylinder
2019 | 1 | 19 | 19 | 2576 | TN | * |  | * | 1400 | 85 |  | 2 | Y |  | 245 | 500 | 104 | [6 empty] | * | * | * | 1560 | H2 generator down, still on cylinder
2019 | 1 | 20 | 20 | 2577 |  | * |  | * | [19 empty] | H2 generator down, still on cylinder
2019 | 1 | 21 | 21 | 2578 | TN | * |  | * | 1395 | 85 |  | 2 | Y |  | 245 | 500 | 103 | [6 empty] | * | * | * | 1540 | H2 generator down, still on cylinder
2019 | 1 | 22 | 22 | 2579 | TN | * |  | * | 1360 | 85 |  | 2 | Y |  | 245 | 500 | 103 | [6 empty] | * | * | * | 1500 | H2 generator down, still on cylinder
2019 | 1 | 23 | 23 | 2580 | TN | * |  | * | 1340 | 85 |  | 2 | Y |  | 245 | 500 | 101 | [6 empty] | * | * | * | 1475 | H2 generator down, still on cylinder
2019 | 1 | 24 | 24 | 2581 | TN | * |  | * | 1340 | 85 |  | 2 | Y |  | 245 | 500 | 100 | [6 empty] | * | * | * | 1470 | H2 generator down, still on cylinder
2019 | 1 | 25 | 25 | 2582 | TN | * |  | * | 1360 | 85 |  | 2 | Y |  | 250 | 500 | 101 | [6 empty] | * | * | * | 1450 | H2 generator down, still on cylinder
2019 | 1 | 26 | 26 | 2583 | TN | * |  | * | 1350 | 85 |  | 2 | Y |  | 245 | 500 | 103 | [6 empty] | * | * | * | 1420 | H2 generator down, still on cylinder
2019 | 1 | 27 | 27 | 2584 | TN | * |  | * | [19 empty] | H2 generator down, still on cylinder
2019 | 1 | 28 | 28 | 2585 | TN | * |  | * | 1360 | 85 |  | 2 | Y |  | 250 | 500 | 103 | [6 empty] | * | * | * | 1415 | H2 generator down, still on cylinder
2019 | 1 | 29 | 29 | 2586 | TN | * |  | * | 1360 | 85 |  | 2 | Y |  | 250 | 500 | 103 | [6 empty] | * | * | * | 1375 | H2 generator down, still on cylinder
2019 | 1 | 30 | 30 | 2587 | TN | * |  | * | 1340 | 85 |  | 2 | Y |  | 250 | 500 | 103 | [6 empty] | * | * | * | 1340 | H2 generator down, still on cylinder
2019 | 1 | 31 | 31 | 2588 | TN | * |  | * | 1340 | 85 |  | 2 | Y |  | 245 | 500 | 104 | [6 empty] | * | * | * | 1320 | H2 generator down, still on cylinder
2019 | 2 | 1 | 32 | 2589 | TN | * |  | * | 1330 | 85 |  | 2 | Y |  | 245 | 500 | 103 | [6 empty] | * | * | * | 1270 | H2 generator down, still on cylinder
2019 | 2 | 2 | 33 | 2590 | TN | * |  | * | 1335 | 85 |  | 2 | Y |  | 245 | 500 | 104 | [6 empty] | * | * | * | 1250 | H2 generator down, still on cylinder
2019 | 2 | 3 | 34 | 2591 | TN | * |  | * | [19 empty] | H2 generator down, still on cylinder
2019 | 2 | 4 | 35 | 2592 | TN | * |  | * | 1330 | 85 |  | 2 | y |  | 245 | 500 | 104 | [6 empty] | * | * | * | 1220 | H2 generator down, still on cylinder
2019 | 2 | 5 | 36 | 2593 | TN | * |  | * | 1325 | 85 |  | 2 | Y |  | 245 | 500 | 104 | [6 empty] | * | * | * | 1215 | H2 generator down, still on cylinder
2019 | 2 | 6 | 37 | 2594 | TN | * |  | * | 1315 | 85 |  | 2 | Y |  | 245 | 495 | 105 | [6 empty] | * | * | * | 1200 | H2 generator down, still on cylinder
2019 | 2 | 7 | 38 | 2595 | TN | * |  | * | 1325 | 85 |  | 2 | Y |  | 240 | 500 | 105 | [6 empty] | * | * | * | 1200 | H2 generator down, still on cylinder
2019 | 2 | 8 | 39 | 2596 | TN | * |  | * | 1325 | 85 |  | 2 | Y |  | 240 | 500 | 105 | [6 empty] | * | * | * | 1145 | H2 generator down, still on cylinder
2019 | 2 | 9 | 40 | 2597 | TN/MM | * |  | * | 1325 | 85 |  | 2 | Y |  | 245 | 495 | 105 | [6 empty] | * | * | * | 1125 | H2 generator down, still on cylinder
2019 | 2 | 10 | 41 | 2598 | TN/MM | * |  | * | 1325 | 85 |  | 2 | Y |  | 245 | 500 | 105 | [6 empty] | * | * | * | 1100 | H2 generator down, still on cylinder
2019 | 2 | 11 | 42 | 2599 | [23 empty] | Replaced float chamber. H2 generator pressure checks success. Switched back to H2 generator for hydrogen supply.
2019 | 2 | 12 | 43 | 2600 | MM | * | * | * | 1310 | 40 |  | 2 | Y |  | 235 | 500 | 105 | [6 empty] | * | * | * | 1050
2019 | 2 | 13 | 44 | 2601 | TN | 61 | * | 40 | 1310 | 85 |  | 2 | Y |  | 240 | 495 | 96 | [6 empty] | * | * | * | 1060
2019 | 2 | 14 | 45 | 2602 | MM | 61 | * | 20 | 1310 | 0 |  | 2 | Y |  | 240 | 500 | 105 | [6 empty] | * | * | * | 1060
2019 | 2 | 15 | 46 | 2603
2019 | 2 | 16 | 47 | 2604 | MM | 61 | * | 40 | 1310 | 0 |  | 2 | Y |  | 240 | 500 | 105 | [6 empty] | * | * | * | 1060
2019 | 2 | 17 | 48 | 2605
2019 | 2 | 18 | 49 | 2606
2019 | 2 | 19 | 50 | 2607 | MM | 61 | * | 40 | 1290 | 0 |  | 2 | Y |  | 240 | 500 | 105 | [6 empty] | * | * | * | 1060 | Added 1L of DI water at 1955z
2019 | 2 | 20 | 51 | 2608
2019 | 2 | 21 | 52 | 2609
2019 | 2 | 22 | 53 | 2610 | MM | 61 | * | 40 | 1280 | 0 |  | 2 | Y |  | 240 | 500 | 100 | [6 empty] | * | * | * | 1060
2019 | 2 | 23 | 54 | 2611
2019 | 2 | 24 | 55 | 2612
2019 | 2 | 25 | 56 | 2613 | MM | 61 | * | 40 | 1290 | 0 |  | 2 | Y |  | 240 | 500 | 105 | [6 empty] | * | * | * | 1050
2019 | 2 | 26 | 57 | 2614
2019 | 2 | 27 | 58 | 2615
2019 | 2 | 28 | 59 | 2616 | MM | 61 | * | 40 | 1280 | 0 |  | 2 | Y |  | 240 | 500 | 100 | [6 empty] | * | * | * | 1000
2019 | 3 | 1 | 60 | 2617
2019 | 3 | 2 | 61 | 2618
2019 | 3 | 3 | 62 | 2619
2019 | 3 | 4 | 63 | 2620 | MM | 61 | * | 40 | 1280 | 0 |  | 2 | Y |  | 230 | 500 | 105 | [6 empty] | * | * | * | 1000
2019 | 3 | 5 | 64 | 2621 | MM | 61 | * | 40 | 1280 | 0 |  | 2 | Y |  | 230 | 500 | 105 | [6 empty] | * | * | * | 1000
2019 | 3 | 6 | 65 | 2622
2019 | 3 | 7 | 66 | 2623
2019 | 3 | 8 | 67 | 2624 | MM | 61 | * | 40 | 1275 | 0 |  | 2 | Y |  | 230 | 500 | 105 | [6 empty] | * | * | * | 1000 | Added 1L of DI water at 2102z
2019 | 3 | 9 | 68 | 2625 | [23 empty] | Repaired damaged wire on CUGC inlet from 1422-1443z
2019 | 3 | 10 | 69 | 2626
2019 | 3 | 11 | 70 | 2627 | MM | 61 | * | 40 | 1275 | 0 |  | 2 | Y |  | 215 | 500 | 100 | [6 empty] | * | * | * | 1000
2019 | 3 | 12 | 71 | 2628
2019 | 3 | 13 | 72 | 2629
2019 | 3 | 14 | 73 | 2630 | MM | 61 | * | 40 | 1250 | 0 |  | 2 | Y |  | 210 | 500 | 100 | [6 empty] | * | * | * | 1000
2019 | 3 | 15 | 74 | 2631
2019 | 3 | 16 | 75 | 2632
2019 | 3 | 17 | 76 | 2633 | MM | 61 | * | 40 | 1240 | 0 |  | 2 | Y |  | 210 | 500 | 90 | [6 empty] | * | * | * | 960
2019 | 3 | 18 | 77 | 2634 | MM | 61 | * | 40 | 1230 | 0 |  | 2 | Y |  | 200 | 500 | 95 | [6 empty] | * | * | * | 950 | Added 600ml of DI water at 2308z
2019 | 3 | 19 | 78 | 2635
2019 | 3 | 20 | 79 | 2636
2019 | 3 | 21 | 80 | 2637
2019 | 3 | 22 | 81 | 2638 | MM | 61 | * | 40 | 1210 | 0 |  | 2 | Y |  | 200 | 500 | 100 | [6 empty] | * | * | * | 940 | Confirmed that the hydrogen cylinder is closed since there has been a drop in psi. Colder temperatures in the TAWO vestibule this week.
2019 | 3 | 23 | 82 | 2639
2019 | 3 | 24 | 83 | 2640
2019 | 3 | 25 | 84 | 2641 | MM | 61 | * | 40 | 1200 | 0 |  | 2 | Y |  | 205 | 500 | 100 | [6 empty] | * | * | * | 950
2019 | 3 | 26 | 85 | 2642
2019 | 3 | 27 | 86 | 2643
2019 | 3 | 28 | 87 | 2644 | MM | 61 | * | 40 | 1210 | 0 |  | 2 | Y |  | 210 | 490 | 100 | [6 empty] | * | * | * | 960
2019 | 3 | 29 | 88 | 2645
2019 | 3 | 30 | 89 | 2646
2019 | 3 | 31 | 90 | 2647 | MM | 61 | * | 40 | 1210 | 0 |  | 2 | Y |  | 210 | 490 | 105 | [6 empty] | * | * | * | 960
2019 | 4 | 1 | 91 | 2648
2019 | 4 | 2 | 92 | 2649
2019 | 4 | 3 | 93 | 2650 | MM | 61 | * | 40 | 1200 | 0 |  | 2 | Y |  | 200 | 490 | 105 | [6 empty] | * | * | * | 940 | Added 1L of DI water at 1653z
2019 | 4 | 4 | 94 | 2651
2019 | 4 | 5 | 95 | 2652 | MM | 61 | * | 40 | 1190 | 0 |  | 2 | Y |  | 200 | 490 | 105 | [6 empty] | * | * | * | 930
2019 | 4 | 6 | 96 | 2653
2019 | 4 | 7 | 97 | 2654
2019 | 4 | 8 | 98 | 2655
2019 | 4 | 9 | 99 | 2656 | MM | 61 | * | 40 | 1190 | 0 |  | 2 | Y |  | 200 | 490 | 105 | [6 empty] | * | * | * | 940
2019 | 4 | 10 | 100 | 2657
2019 | 4 | 11 | 101 | 2658
2019 | 4 | 12 | 102 | 2659 | MM | 61 | * | 40 | 1190 | 0 |  | 2 | Y |  | 205 | 490 | 105 | [6 empty] | * | * | * | 940
2019 | 4 | 13 | 103 | 2660
2019 | 4 | 14 | 104 | 2661
2019 | 4 | 15 | 105 | 2662 | MM | 61 | * | 40 | 1190 | 0 |  | 2 | Y |  | 200 | 490 | 105 | [6 empty] | * | * | * | 900
2019 | 4 | 16 | 106 | 2663
2019 | 4 | 17 | 107 | 2664
2019 | 4 | 18 | 108 | 2665 | MM | 61 | * | 40 | 1150 | 0 |  | 2 | Y |  | 200 | 490 | 100 | [6 empty] | * | * | * | 900
2019 | 4 | 19 | 109 | 2666
2019 | 4 | 20 | 110 | 2667
2019 | 4 | 21 | 111 | 2668 | MM | 61 | * | 40 | 1150 | 0 |  | 2 | Y |  | 200 | 460 | 100 | [6 empty] | * | * | * | 890 | Added 400ml of DI water at 1316z
2019 | 4 | 22 | 112 | 2669
2019 | 4 | 23 | 113 | 2670
2019 | 4 | 24 | 114 | 2671 | MM | 61 | * | 40 | 1120 | 0 |  | 2 | Y |  | 200 | 460 | 100 | [6 empty] | * | * | * | 850
2019 | 4 | 25 | 115 | 2672
2019 | 4 | 26 | 116 | 2673
2019 | 4 | 27 | 117 | 2674 | MM | 61 | * | 40 | 1120 | 0 |  | 2 | Y |  | 200 | 460 | 105 | [6 empty] | * | * | * | 860 | Added 1L of DI water at 1334z
2019 | 4 | 28 | 118 | 2675
2019 | 4 | 29 | 119 | 2676 | [23 empty] | Jacques Hueber onsite
2019 | 4 | 30 | 120 | 2677 | MM/JM | [22 empty] | Jacques Hueber onsite
2019 | 5 | 1 | 121 | 2678 | MM/JM | [22 empty] | Jacques Hueber onsite
2019 | 5 | 2 | 122 | 2679 | MM/JM | [22 empty] | Jacques Hueber onsite
2019 | 5 | 3 | 123 | 2680 | MM/JM | - | - | - | 1100 | 0 |  | 1 | y |  | 200 | 400 | 110 | [6 empty] | - | * | * | 1050 | Currently running on H2 cylinder, not generator
2019 | 5 | 4 | 124 | 2681
2019 | 5 | 5 | 125 | 2682
2019 | 5 | 6 | 126 | 2683
2019 | 5 | 7 | 127 | 2684 | JM | - | - | - | 1100 | 0 |  | 1 | y |  | 200 | 380 | 110 |  | * |  |  |  | * |  | * | * | 950
2019 | 5 | 8 | 128 | 2685
2019 | 5 | 9 | 129 | 2686
2019 | 5 | 10 | 130 | 2687 | JM | - | - | - | 1100 | 0 |  | 1 | y |  | 200 | 440 | 110 | 8.5 | * | 1850 | 1925 | 1975 | * |  | * | * | 860 | Added columns S-W to daily checks spreadsheet
2019 | 5 | 11 | 131 | 2688
2019 | 5 | 12 | 132 | 2689
2019 | 5 | 13 | 133 | 2690 | JM | - | - | - | 1100 | 0 |  | 1 | y |  | 200 | 450 | 110 | 8.5 | * | 1850 | 1950 | 1975 | * |  | * | * | 800
2019 | 5 | 14 | 134 | 2691
2019 | 5 | 15 | 135 | 2692
2019 | 5 | 16 | 136 | 2693 | JM | - | - | - | 1100 | 0 |  | 1 | y |  | 200 | 440 | 110 | 8.5 | * | 1900 | 1975 | 2000 | * |  | * | * | 750
2019 | 5 | 17 | 137 | 2694
2019 | 5 | 18 | 138 | 2695
2019 | 5 | 19 | 139 | 2696 | JM | - | - | - | 1100 | 0 |  | 1 | y |  | 200 | 440 | 100 | 8.5 | * | 1900 | 1975 | 2000 | * |  | * | * | 690
2019 | 5 | 20 | 140 | 2697
2019 | 5 | 21 | 141 | 2698
2019 | 5 | 22 | 142 | 2699 | JM | - | - | - | 1100 | 0 |  | 1 | y |  | 200 | 440 | 100 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 620
2019 | 5 | 23 | 143 | 2700
2019 | 5 | 24 | 144 | 2701
2019 | 5 | 25 | 145 | 2702 | JM | - | - | - | 1100 | 0 |  | 1 | y |  | 200 | 440 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 550
2019 | 5 | 26 | 146 | 2703
2019 | 5 | 27 | 147 | 2704
2019 | 5 | 28 | 148 | 2705 | JM | - | - | - | 1100 | 0 |  | 1 | y |  | 200 | 440 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 490
2019 | 5 | 29 | 149 | 2706
2019 | 5 | 30 | 150 | 2707
2019 | 5 | 31 | 151 | 2708
2019 | 6 | 1 | 152 | 2709 | JM | - | - | - | 1100 | 0 |  | 1 | y |  | 200 | 440 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 390
2019 | 6 | 2 | 153 | 2710
2019 | 6 | 3 | 154 | 2711
2019 | 6 | 4 | 155 | 2712 | JM | - | - | - | 1050 | 0 |  | 1 | y |  | 200 | 440 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 310
2019 | 6 | 5 | 156 | 2713
2019 | 6 | 6 | 157 | 2714
2019 | 6 | 7 | 158 | 2715 | JM | - | - | - | 1050 | 0 |  | 1 | y |  | 200 | 440 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 250
2019 | 6 | 8 | 159 | 2716
2019 | 6 | 9 | 160 | 2717
2019 | 6 | 10 | 161 | 2718 | JM | - | - | - | 1050 | 0 |  | 1 | y |  | 200 | 440 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 190
2019 | 6 | 11 | 162 | 2719
2019 | 6 | 12 | 163 | 2720 | JM | - | - | - | 1050 | 0 |  | 1 | y |  | 200 | 440 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 125 | Changed out H2 tank, 2500 psi on new.
2019 | 6 | 13 | 164 | 2721
2019 | 6 | 14 | 165 | 2722
2019 | 6 | 15 | 166 | 2723 | JM | - | - | - | 1010 | 0 |  | 1 | y |  | 195 | 440 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 2410
2019 | 6 | 16 | 167 | 2724
2019 | 6 | 17 | 168 | 2725
2019 | 6 | 18 | 169 | 2726 | JM | - | - | - | 1000 | 0 |  | 1 | y |  | 190 | 420 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 2325
2019 | 6 | 19 | 170 | 2727
2019 | 6 | 20 | 171 | 2728
2019 | 6 | 21 | 172 | 2729 | JM | - | - | - | 1000 | 0 |  | 1 | y |  | 160 | 400 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 2250
2019 | 6 | 22 | 173 | 2730
2019 | 6 | 23 | 174 | 2731
2019 | 6 | 24 | 175 | 2732 | JM | - | - | - | 1000 | 0 |  | 1 | y |  | 190 | 410 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 2190
2019 | 6 | 25 | 176 | 2733
2019 | 6 | 26 | 177 | 2734
2019 | 6 | 27 | 178 | 2735 | JM | - | - | - | 1000 | 0 |  | 1 | y |  | 190 | 410 | 95 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 2100
2019 | 6 | 28 | 179 | 2736
2019 | 6 | 29 | 180 | 2737
2019 | 6 | 30 | 181 | 2738 | JM | - | - | - | 1000 | 0 |  | 1 | y |  | 190 | 410 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 2050
2019 | 7 | 1 | 182 | 2739
2019 | 7 | 2 | 183 | 2740
2019 | 7 | 3 | 184 | 2741 | JM | - | - | - | 980 | 0 |  | 1 | y |  | 160 | 400 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 1980
2019 | 7 | 4 | 185 | 2742
2019 | 7 | 5 | 186 | 2743
2019 | 7 | 6 | 187 | 2744 | JM | - | - | - | 980 | 0 |  | 1 | y |  | 150 | 390 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 1900
2019 | 7 | 7 | 188 | 2745
2019 | 7 | 8 | 189 | 2746
2019 | 7 | 9 | 190 | 2747 | JM | - | - | - | 990 | 0 |  | 1 | y |  | 150 | 400 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 1850
2019 | 7 | 10 | 191 | 2748 | JM | [21 empty] | 1800 | Picarro shut down at 1313z by Jacques. Back on at ~1645 with return of power. CUGC shut down temporarily during UPS battery swap. Instrument turned off from 1525-1738
2019 | 7 | 11 | 192 | 2749
2019 | 7 | 12 | 193 | 2750 | JM | - | - | - | 990 | 0 |  | 1 | y |  | 160 | 390 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 1790
2019 | 7 | 13 | 194 | 2751
2019 | 7 | 14 | 195 | 2752
2019 | 7 | 15 | 196 | 2753 | JM | - | - | - | 990 | 0 |  | 1 | y |  | 150 | 390 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 1700
2019 | 7 | 16 | 197 | 2754
2019 | 7 | 17 | 198 | 2755
2019 | 7 | 18 | 199 | 2756 | JM | - | - | - | 990 | 0 |  | 1 | y |  | 150 | 390 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 1640
2019 | 7 | 19 | 200 | 2757
2019 | 7 | 20 | 201 | 2758
2019 | 7 | 21 | 202 | 2759 | JM | - | - | - | 950 | 0 |  | 1 | y |  | 150 | 390 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 1580
2019 | 7 | 22 | 203 | 2760
2019 | 7 | 23 | 204 | 2761
2019 | 7 | 24 | 205 | 2762
2019 | 7 | 25 | 206 | 2763 | JM | - | - | - | 920 | 0 |  | 1 | y |  | 150 | 390 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 1480
2019 | 7 | 26 | 207 | 2764
2019 | 7 | 27 | 208 | 2765
2019 | 7 | 28 | 209 | 2766 | JM | - | - | - | 900 | 0 |  | 1 | y |  | 150 | 390 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 1400
2019 | 7 | 29 | 210 | 2767
2019 | 7 | 30 | 211 | 2768
2019 | 7 | 31 | 212 | 2769 | JM | [22 empty] | Shut down overnight due to warm temperatures
2019 | 8 | 1 | 213 | 2770 | JM | - | - | - | 900 | 0 |  | 1 | y |  | 150 | 360 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 1320
2019 | 8 | 2 | 214 | 2771
2019 | 8 | 3 | 215 | 2772
2019 | 8 | 4 | 216 | 2773 | JM | - | - | - | 900 | 0 |  | 1 | y |  | 150 | 360 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 1250
2019 | 8 | 5 | 217 | 2774
2019 | 8 | 6 | 218 | 2775
2019 | 8 | 7 | 219 | 2776 | JM | - | - | - | 900 | 0 |  | 1 | y |  | 140 | 360 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 1190
2019 | 8 | 8 | 220 | 2777
2019 | 8 | 9 | 221 | 2778
2019 | 8 | 10 | 222 | 2779 | JM | - | - | - | 890 | 0 |  | 1 | y |  | 140 | 360 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 1100
2019 | 8 | 11 | 223 | 2780
2019 | 8 | 12 | 224 | 2781
2019 | 8 | 13 | 225 | 2782 | JM | - | - | - | 890 | 0 |  | 1 | y |  | 110 | 360 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 1050
2019 | 8 | 14 | 226 | 2783
2019 | 8 | 15 | 227 | 2784
2019 | 8 | 16 | 228 | 2785 | JM | - | - | - | 890 | 0 |  | 1 | y |  | 110 | 360 | 110 | 8.5 | * | 1900 | 1950 | 1975 | * |  | * | * | 990
2019 | 8 | 17 | 229 | 2786
2019 | 8 | 18 | 230 | 2787
2019 | 8 | 19 | 231 | 2788 | NB | - | - | - | 850 | 0 |  | 1 | y |  | 130 | 360 | 111 | 8.5 | * | 1890 | 1960 | 1990 | * |  | * | * | 900 | Unplanned power outage 0839-0848z; planned power outage 1935-1937z.
2019 | 8 | 20 | 232 | 2789
2019 | 8 | 21 | 233 | 2790
2019 | 8 | 22 | 234 | 2791 | NB | - | - | - | 850 | 0 |  | 1 | y |  | 125 | 350 | 110 | 8.5 | * | 1890 | 1960 | 1990 | no |  | * | * | 825 | Picarro CO2 graph not updating.
2019 | 8 | 23 | 235 | 2792
2019 | 8 | 24 | 236 | 2793
2019 | 8 | 25 | 237 | 2794 | NB | - | - | - | 820 | 0 |  | 1 | y |  | 120 | 340 | 110 | 8.5 | * | 1870 | 1940 | 1970 | * |  | * | * | 780
2019 | 8 | 26 | 238 | 2795
2019 | 8 | 27 | 239 | 2796
2019 | 8 | 28 | 240 | 2797 | NB | - | - | - | 820 | 0 |  | 1 | y |  | 120 | 350 | 100 | 8.5 | * | 1870 | 1950 | 1970 | * |  | * | * | 710
2019 | 8 | 29 | 241 | 2798 | NB | 60 | add 2 L | 51 | 820 | 0 |  | 1 | y |  | 120 | 350 | 110 | 8.5 | * | 1870 | 1950 | 1970 | * | * | * | * | 680 | Installed and switched to new H2 generator at ~1645z.
2019 | 8 | 30 | 242 | 2799 | NB | 60 | add .5 L | 57 | 810 | off |  | 1 | y |  | 110 | 350 | 100 | 8.5 | * | 1870 | 1940 | 1970 | * | * | * | * | off | CH4 Std 2 status unchanged… just logging as "off" instead of "0" for clarity.
2019 | 8 | 31 | 243 | 2800 | NB | 60 | * | 61 | 810 | off |  | 1 | y |  | 110 | 350 | 110 | 8.5 | * | 1870 | 1940 | 1970 | * | * | * | * | off | Ran old H2 generator for ~15 min.
2019 | 9 | 1 | 244 | 2801
2019 | 9 | 2 | 245 | 2802 | NB | 60 | * | 60 | 810 | off |  | 1 | y |  | 100 | 340 | 110 | 8.5 | * | 1870 | 1940 | 1970 | * | * | * | * | off
2019 | 9 | 3 | 246 | 2803 | NB | [9 empty] | 1705 | [8 empty] | Installed new filter |  |  |  | 1600-1700z: Switched from NMHC standard 1 to new BH#6 standard. End pressure on NMHC cylinder = 100 psi, start pressure on BH#6 cylinder = 1705. Installed new desiccant filter on H2 generator.
2019 | 9 | 4 | 247 | 2804 | NB | 60 | * | 57 | 800 | off |  | 1 | y |  | 1695 | 340 | 110 | 8.5 | * | 1840 | 1940 | 1940 | * | * | * | * | off | Cooler building temperature in TAWO = slight drop in all cylinder psi.
2019 | 9 | 5 | 248 | 2805
2019 | 9 | 6 | 249 | 2806 | NB | 60 | * | 49 | 800 | off |  | 1 | y |  | 1700 | 340 | 110 | 8.5 | * | 1870 | 1940 | 1970 | * | * | * | * | off
2019 | 9 | 7 | 250 | 2807
2019 | 9 | 8 | 251 | 2808 | NB | 60 | * | 61 | 790 | off |  | 1 | y |  | 1700 | 330 | 110 | 8.5 | * | 1860 | 1930 | 1960 | * | * | * | * | off
2019 | 9 | 9 | 252 | 2809
2019 | 9 | 10 | 253 | 2810 | NB | 60 | * | 59 | 780 | off |  | 1 | y |  | 1700 | 330 | 105 | 8.5 | * | 1870 | 1930 | 1960 | * | * | * | * | off
2019 | 9 | 11 | 254 | 2811
2019 | 9 | 12 | 255 | 2812 | NB | 60 | * | 61 | 760 | off |  | 1 | y |  | 1660 | 330 | 105 | 8.5 | * | 1830 | 1890 | 1910 | * | * | * | * | off | Cooler building temperature in TAWO = slight drop in all cylinder psi.
2019 | 9 | 13 | 256 | 2813
2019 | 9 | 14 | 257 | 2814 | NB | 60 | * | 63 | 750 | off |  | 1 | y |  | 1660 | 330 | 105 | 8.5 | * | 1830 | 1900 | 1910 | * | * | * | * | off | cooler building temps
2019 | 9 | 15 | 258 | 2815
2019 | 9 | 16 | 259 | 2816 | NB | 60 | * | 62 | 750 | off |  | 1 | y |  | 1660 | 320 | 105 | 8.5 | * | 1820 | 1890 | 1910 | * | * | * | * | off
2019 | 9 | 17 | 260 | 2817
2019 | 9 | 18 | 261 | 2818 | NB | 60 | * | 62 | 740 | off |  | 1 | y |  | 1650 | 320 | 100 | 8.5 | * | 1820 | 1880 | 1910 | * | * | * | * | off
2019 | 9 | 19 | 262 | 2819
2019 | 9 | 20 | 263 | 2820 | NB | 60 | filled | 52 | 730 | off |  | 1 | y |  | 1660 | 320 | 100 | 8.5 | * | 1820 | 1890 | 1910 | * | * | * | * | off
2019 | 9 | 21 | 264 | 2821
2019 | 9 | 22 | 265 | 2822 | NB | 60 | * | 62 | 730 | off |  | 1 | y |  | 1680 | 320 | 105 | 8.5 | * | 1850 | 1910 | 1930 | * | * | * | * | off
2019 | 9 | 23 | 266 | 2823
2019 | 9 | 24 | 267 | 2824 | NB | 60 | * | 49 | 730 | off |  | 1 | y |  | 1690 | 320 | 105 | 8.5 | * | 1840 | 1920 | 1940 | * | * | * | * | off
2019 | 9 | 25 | 268 | 2825
2019 | 9 | 26 | 269 | 2826 | NB | 60 | * | 60 | 710 | off |  | 1 | y |  | 1670 | 310 | 105 | 8.5 | * | 1820 | 1890 | 1910 | * | * | * | * | off
2019 | 9 | 27 | 270 | 2827
2019 | 9 | 28 | 271 | 2828 | NB | 60 | * | 60 | 720 | off |  | 1 | y |  | 1680 | 320 | 105 | 8.5 | * | 1840 | 1900 | 1920 | * | * | * | * | off
2019 | 9 | 29 | 272 | 2829
2019 | 9 | 30 | 273 | 2830 | NB | 60 | * | 50 | 700 | off |  | 1 | y |  | 1670 | 320 | 105 | 8.5 | * | 1820 | 1900 | 1920 | * | * | * | * | off
2019 | 10 | 1 | 274 | 2831
2019 | 10 | 2 | 275 | 2832 | NB | 60 | * | 58 | 700 | off |  | 1 | y |  | 1670 | 310 | 105 | 8.5 | * | 1810 | 1900 | 1910 | * | * | * | * | off
2019 | 10 | 3 | 276 | 2833
2019 | 10 | 4 | 277 | 2834 | NB | 60 | * | 61 | 690 | off |  | 1 | y |  | 1670 | 310 | 100 | 8.5 | * | 1830 | 1900 | 1920 | * | * | * | * | off
2019 | 10 | 5 | 278 | 2835
2019 | 10 | 6 | 279 | 2836 | NB | 60 | * | 57 | 690 | off |  | 1 | y |  | 1670 | 310 | 100 | 8.5 | * | 1830 | 1900 | 1920 | * | * | * | * | off
2019 | 10 | 7 | 280 | 2837
2019 | 10 | 8 | 281 | 2838 | NB | 60 | refilled | 54 | 680 | off |  | 1 | y |  | 1670 | 310 | 100 | 8.5 | * | 1820 | 1900 | 1920 | * | * | * | * | off
2019 | 10 | 9 | 282 | 2839
2019 | 10 | 10 | 283 | 2840 | NB | 60 | * | 51 | 680 | off |  | 1 | y |  | 1660 | 300 | 95 | 8.5 | * | 1810 | 1900 | 1920 | * | * | * | * | off
2019 | 10 | 11 | 284 | 2841
2019 | 10 | 12 | 285 | 2842 | HA/SWD | 60 | * | 57 | 680 | off |  | 0.9 | Y |  | 1650 | 300 | 100 | 8.5 | * | 1810 | 1890 | 1910 | * | * | * | * | off
2019 | 10 | 13 | 286 | 2843
2019 | 10 | 14 | 287 | 2844 | HA | 60 | * | 56 | 660 | off |  | 0.9 | Y |  | 1650 | 300 | 100 | 8.5 | * | 1800 | 1900 | 1920 | * | * | * | * | off
2019 | 10 | 15 | 288 | 2845 | 0 | [22 empty] | Planned power outage 1310-1311, 1318-1319 Z; all instrument parameters within acceptable range
2019 | 10 | 16 | 289 | 2846 | HA | 60 | * | 59 | 660 | off |  | 0.9 | Y |  | 1650 | 300 | 100 | 8.5 | * | 1810 | 1900 | 1920 | * | * | * | * | off
2019 | 10 | 17 | 290 | 2847 | HA | 60 | * | 59 | 660 | off |  | 0.9 | Y |  | 1660 | 300 | 100 | 8.5 | * | 1810 | 1890 | 1920 | * | * | * | * | off | Network issues experienced at TAWO from 0800 UTC to ~1530 UTC with intermittent losses of connectivity
2019 | 10 | 18 | 291 | 2848 | 0
2019 | 10 | 19 | 292 | 2849 | HA | 60 | * | 61 | 660 | off |  | 0.9 | Y |  | 1660 | 300 | 100 | 8.5 | * | 1810 | 1890 | 1920 | * | * | * | * | off
2019 | 10 | 20 | 293 | 2850 | 0
2019 | 10 | 21 | 294 | 2851 | 0
2019 | 10 | 22 | 295 | 2852 | HA | 60 | * | 58 | 610 | off |  | 0.9 | Y |  | 1640 | 300 | 95 | 8.5 | * | 1800 | 1890 | 1900 | * | * | * | * | off
2019 | 10 | 23 | 296 | 2853 | HA | 60 | * | 58 | 610 | off |  | 0.9 | Y |  | 1610 | 290 | 95 | 8.5 | * | 1800 | 1870 | 1900 | * | * | *
2019 | 10 | 24 | 297 | 2854 | 0 | [20 empty] | * | off
2019 | 10 | 25 | 298 | 2855 | 0
2019 | 10 | 26 | 299 | 2856 | HA | 60 | refilled | 58 | 610 | off |  | 0.9 | Y |  | 1620 | 290 | 95 | 8.5 | * | 1800 | 1850 | 1910 | * | * | * | * | off | Refilled water at 1323 UTC
2019 | 10 | 27 | 300 | 2857 | 0
2019 | 10 | 28 | 301 | 2858 | HA | 60 | * | 56 | 610 | off |  | 0.9 | Y |  | 1640 | 290 | 95 | 8.5 | * | 1800 | 1890 | 1910 | * | * | * | * | off
2019 | 10 | 29 | 302 | 2859 | 0
2019 | 10 | 30 | 303 | 2860 | HA | 60 | * | 60 | 600 | off |  | 0.9 | Y |  | 1620 | 280 | 95 | 8.5 | * | 1800 | 1890 | 1910 | * | * | * | * | off
2019 | 10 | 31 | 304 | 2861 | 0
2019 | 11 | 1 | 305 | 2862 | HA | 60 | * | 58 | 590 | off |  | 0.9 | Y |  | 1630 | 280 | 95 | 8.5 | * | 1800 | 1890 | 1910 | * | * | * | * | off
2019 | 11 | 2 | 306 | 2863 | 0
2019 | 11 | 3 | 307 | 2864 | 0
2019 | 11 | 4 | 308 | 2865 | HA | 60 | * | 52 | 590 | off |  | 0.9 | Y |  | 1620 | 260 | 90 | 8.5 | * | 1800 | 1880 | 1900 | * | * | * | * | off
2019 | 11 | 5 | 309 | 2866 | 0
2019 | 11 | 6 | 310 | 2867 | 0
2019 | 11 | 7 | 311 | 2868 | HA | 60 | * | 59 | 580 | off |  | 0.9 | Y |  | 1630 | 260 | 92 | 8.5 | * | 1800 | 1890 | 1910 | * | * | * | * | off
2019 | 11 | 8 | 312 | 2869 | 0 | [22 empty] | Started up old H2 generator not being used; charged Bios Flow Meter for 24 hrs.
2019 | 11 | 9 | 313 | 2870 | 0
2019 | 11 | 10 | 314 | 2871 | HA | 60 | refilled | 57 | 550 | off |  | 0.9 | Y |  | 1620 | 260 | 94 | 8.5 | * | 1800 | 1880 | 1900 | * | * | * | * | off | Refilled H2 generator water at 1401 UTC
2019 | 11 | 11 | 315 | 2872 | 0
2019 | 11 | 12 | 316 | 2873 | 0
2019 | 11 | 13 | 317 | 2874 | [23 empty] | Could not visit TAWO due to high wind storm conditions
2019 | 11 | 14 | 318 | 2875 | [23 empty] | Could not visit TAWO due to high wind storm conditions
2019 | 11 | 15 | 319 | 2876 | HA | 60 | * | 58 | 520 | off |  | 0.9 | Y |  | 1630 | 260 | 87 | 8.5 | * | 1800 | 1890 | 1910 | * | * | * | * | off | Loss of connectivity from tower instruments to TAWO from 2019-11-14 1800 UTC to 2019-11-15 1937 UTC, as well as aerosol instruments. Interruption suspected to be due to high wind/storm conditions.
2019 | 11 | 16 | 320 | 2877 | 0
2019 | 11 | 17 | 321 | 2878 | 0
2019 | 11 | 18 | 322 | 2879 | HA | 60 | * | 60 | 520 | off |  | 0.9 | Y |  | 1620 | 260 | 90 | 8.5 | * | 1800 | 1880 | 1900 | * | * | * | * | off
2019 | 11 | 19 | 323 | 2880 | 0
2019 | 11 | 20 | 324 | 2881 | 0
2019 | 11 | 21 | 325 | 2882 | HA | 60 | * | 60 | 500 | off |  | 0.9 | Y |  | 1620 | 250 | 95 | 8.5 | * | 1800 | 1890 | 1910 | * | * | * | * | off
2019 | 11 | 22 | 326 | 2883 | 0
2019 | 11 | 23 | 327 | 2884 | 0
2019 | 11 | 24 | 328 | 2885 | 0 | [22 empty] | Did not visit TAWO
2019 | 11 | 25 | 329 | 2886 | HA | 60 | refilled | 58 | 500 | off |  | 0.9 | Y |  | 1620 | 250 | 96 | 8.5 | * | 1800 | 1890 | 1910 | * | * | * | * | off | Topped off water in H2 generator
2019 | 11 | 26 | 330 | 2887 | 0
2019 | 11 | 27 | 331 | 2888 | 0
2019 | 11 | 28 | 332 | 2889 | HA | 60 | * | 49 | 490 | off |  | 0.9 | Y |  | 1610 | 250 | 90 | 8.5 | * | 1800 | 1890 | 1900 | * | * | * | * | off
2019 | 11 | 29 | 333 | 2890 | 0
2019 | 11 | 30 | 334 | 2891 | 0
2019 | 12 | 1 | 335 | 2892 | 0 | [22 empty] | Could not safely travel to TAWO due to 40-50 kn gusting winds
2019 | 12 | 2 | 336 | 2893 | HA | 60 | * | 56 | 490 | off |  | 0.9 | Y |  | 1620 | 240 | 85 | 8.5 | * | 1800 | 1890 | 1900 | * | * | * | * | off | Summit experienced an unplanned power outages from 1935-1943Z, and 1944-1947Z. TAWO successfully backed up by UPS power, all instruments confirmed powered and functioning properly after outage.
2019 | 12 | 3 | 337 | 2894 | 0
2019 | 12 | 4 | 338 | 2895 | 0
2019 | 12 | 5 | 339 | 2896 | HA | 60 | * | 57 | 480 | off |  | 0.9 | Y |  | 1610 | 240 | 90 | 8.5 | * | 1800 | 1890 | 1900 | * | * | * | * | off
2019 | 12 | 6 | 340 | 2897 | 0
2019 | 12 | 7 | 341 | 2898 | 0
2019 | 12 | 8 | 342 | 2899 | HA | 60 | * | 61 | 450 | off |  | 0.9 | Y |  | 1610 | 240 | 90 | 8.5 | * | 1800 | 1880 | 1890 | * | * | * | * | off
2019 | 12 | 9 | 343 | 2900 | 0
2019 | 12 | 10 | 344 | 2901 | 0
2019 | 12 | 11 | 345 | 2902 | HA | 60 | * | 59 | 390 | off |  | 0.9 | Y |  | 1610 | 225 | 93 | 8.5 | * | 1790 | 1870 | 1890 | * | * | * | * | off
2019 | 12 | 12 | 346 | 2903 | 0
2019 | 12 | 13 | 347 | 2904 | 0
2019 | 12 | 14 | 348 | 2905 | HA | 60 | * | 61 | 400 | off |  | 0.8 | Y |  | 1600 | 230 | 85 | 8.5 | * | 1790 | 1870 | 1890 | * | * | * | * | off
2019 | 12 | 15 | 349 | 2906 | 0
2019 | 12 | 16 | 350 | 2907 | 0 |  | refill w/ 1 liter
2019 | 12 | 17 | 351 | 2908 | 0 | [22 empty] | Did not visit TAWO
2019 | 12 | 18 | 352 | 2909 | HA | 60 | * | 47 | 390 | off |  | 0.9 | Y |  | 1610 | 220 | 93 | 8.5 | * | 1800 | 1880 | 1890 | * | * | * | * | off
2019 | 12 | 19 | 353 | 2910 | 0
2019 | 12 | 20 | 354 | 2911 | 0
2019 | 12 | 21 | 355 | 2912 | HA | 60 | * | 54 | 390 | off |  | 0.9 | Y |  | 1610 | 220 | 85 | 8.5 | * | 1800 | 1880 | 1890 | * | * | * | * | off
2019 | 12 | 22 | 356 | 2913 | 0
2019 | 12 | 23 | 357 | 2914 | 0
2019 | 12 | 24 | 358 | 2915 | HA | 60 | * | 56 | 390 | off |  | 0.9 | Y |  | 1610 | 220 | 85 | 8.5 | * | 1790 | 1870 | 1890 | * | * | * | * | off
2019 | 12 | 25 | 359 | 2916 | 0
2019 | 12 | 26 | 360 | 2917 | 0
2019 | 12 | 27 | 361 | 2918 | HA | 60 | * | 57 | 380 | off |  | 0.9 | Y |  | 1605 | 220 | 83 | 8.5 | * | 1790 | 1860 | 1890 | * | * | * | * | off
2019 | 12 | 28 | 362 | 2919 | 0
2019 | 12 | 29 | 363 | 2920 | 0
2019 | 12 | 30 | 364 | 2921 | HA | 60 | refill w/ 1 liter | 57 | 320 | off |  | 0.9 | Y |  | 1600 | 210 | 85 | 8.5 | * | 1790 | 1870 | 1890 | * | * | * | * | off
2019 | 12 | 31 | 365 | 2922
2020 | 1 | 1 | 1 | 2923
2020 | 1 | 2 | 2 | 2924 | HA | 60 | * | 57 | 310 | off |  | 0.9 | Y |  | 1590 | 210 | 85 | 8.5 | * | 1780 | 1840 | 1880 | * | * | * | * | * | off
2020 | 1 | 3 | 3 | 2925
2020 | 1 | 4 | 4 | 2926
2020 | 1 | 5 | 5 | 2927 | HA | 60 | * | 57 | 290 | off |  | 0.85 | Y |  | 1590 | 210 | 85 | 8 | * | 1790 | 1870 | 1890 | * | * | * | * | * | off
2020 | 1 | 6 | 6 | 2928
2020 | 1 | 7 | 7 | 2929
2020 | 1 | 8 | 8 | 2930 | HA | 60 | * | 57 | 290 | off |  | 0.8 | Y |  | 1590 | 200 | 80 | 8 | * | 1780 | 1860 | 1890 | * | * | * | * | * | off
2020 | 1 | 9 | 9 | 2931
2020 | 1 | 10 | 10 | 2932
2020 | 1 | 11 | 11 | 2933 | HA | 60 | refilled with 1 liter DI water | 49 | 280 | off |  | 0.9 | Y |  | 1590 | 200 | 85 | 8 | * | 1790 | 1860 | 1880 | * | * | * | * | * | off
2020 | 1 | 12 | 12 | 2934
2020 | 1 | 13 | 13 | 2935
2020 | 1 | 14 | 14 | 2936 | HA | 60 | * | 59 | 270 | off |  | 0.8 | Y |  | 1600 | 190 | 80 | 8.5 | * | 1790 | 1870 | 1890 | * | * | * | * | * | off
2020 | 1 | 15 | 15 | 2937
2020 | 1 | 16 | 16 | 2938
2020 | 1 | 17 | 17 | 2939 | HA | 60 | * | 57 | 220 | off |  | 0.9 | Y |  | 1590 | 180 | 85 | 8.5 | * | 1780 | 1860 | 1890 | * | * | * | * | * | off
2020 | 1 | 18 | 18 | 2940
2020 | 1 | 19 | 19 | 2941
2020 | 1 | 20 | 20 | 2942 | HA | 60 | * | 50 | 210 | off |  | 0.9 | Y |  | 1600 | 175 | 85 | 8 | * | 1790 | 1870 | 1890 | * | * | * | * | * | off
2020 | 1 | 21 | 21 | 2943
2020 | 1 | 22 | 22 | 2944
2020 | 1 | 23 | 23 | 2945 | HA | 60 | * | 60 | 210 | off |  | 0.9 | Y |  | 1600 | 175 | 85 | 8 | * | 1790 | 1870 | 1890 | * | * | * | * | * | off
2020 | 1 | 24 | 24 | 2946
2020 | 1 | 25 | 25 | 2947
2020 | 1 | 26 | 26 | 2948 | HA | 60 | * | 59 | 210 | off |  | 0.9 | Y |  | 1600 | 180 | 85 | 8 | * | 1790 | 1870 | 1890 | * | * | * | * | * | off
2020 | 1 | 27 | 27 | 2949 |  |  | refilled with 1 liter DI water
2020 | 1 | 28 | 28 | 2950
2020 | 1 | 29 | 29 | 2951 | HA | 60 | * | 56 | 210 | off |  | 0.8 | Y |  | 1590 | 180 | 85 | 8 | * | 1780 | 1860 | 1890 | * | * | * | * | * | off
2020 | 1 | 30 | 30 | 2952 | [23 empty] | Increased inlet sampling line heater settings from "4" to "5"
2020 | 1 | 31 | 31 | 2953
2020 | 2 | 1 | 32 | 2954 | HA | 60 | * | 59 | 210 | off |  | 0.8 | Y |  | 1590 | 180 | 85 | 8.2 | * | 1780 | 1870 | 1890 | * | * | * | * | * | off
2020 | 2 | 2 | 33 | 2955
2020 | 2 | 3 | 34 | 2956 | HA | 60 | * | 57 | 150 | off |  | 0.9 | Y |  | 1600 | 180 | 85 | 8 | * | 1780 | 1870 | 1890 | * | * | * | * | *
2020 | 2 | 4 | 35 | 2957 | [23 empty] | Increased inlet sampling line heater settings from "5" to "6"
2020 | 2 | 5 | 36 | 2958 | HA | 60 | * | 57 | 130 | off |  | 0.9 | Y |  | 1600 | 180 | 85 | 8 | * | 1780 | 1870 | 1900 | * | * | * | * | * | off
2020 | 2 | 6 | 37 | 2959
2020 | 2 | 7 | 38 | 2960 | HA | 60 | * | 55 | 120 | off |  | 0.8 | Y |  | 1600 | 175 | 83 | 7 | * | 1780 | 1870 | 1900 | * | * | * | * | off
2020 | 2 | 8 | 39 | 2961
2020 | 2 | 9 | 40 | 2962
2020 | 2 | 10 | 41 | 2963 | HA | 60 | refilled with 1 liter DI water | 59 | 110 | off |  | 0.9 | Y |  | 1600 | 175 | 83 | 8 | * | 1790 | 1880 | 1900 | * | * | * | * | off
2020 | 2 | 11 | 42 | 2964
2020 | 2 | 12 | 43 | 2965
2020 | 2 | 13 | 44 | 2966 | HA | 60 | * | 49 | 175 | off |  | 0.9 | Y |  | 1590 | 100 | 82 | 8.2 | * | 1780 | 1870 | 1890 | * | * | * | * | off
2020 | 2 | 14 | 45 | 2967
2020 | 2 | 15 | 46 | 2968 | HA/MM | 60 | * | 56 | - | off |  | 0.9 | Y |  | 1600 | 175 | 85 | 8.5 | * | 1790 | 1890 | 1900 | * | * | * | * | off
2020 | 2 | 16 | 47 | 2969
2020 | 2 | 17 | 48 | 2970
2020 | 2 | 18 | 49 | 2971 | MM | 60 | * | 40 | zero | off |  | 1 | Y |  | 1590 | 150 | 85 | 8.5 | * | 1770 | 1900 | 1880 | * | * | * | * | off
2020 | 2 | 19 | 50 | 2972
2020 | 2 | 20 | 51 | 2973 | MM | 60 | * | 40 | zero | off |  | 1 | Y |  | 1600 | 150 | 80 | 8 | * | 1780 | 1890 | 1900 | * | * | * | * | off | CH4 Std 1 derlivery regulator psi 45
2020 | 2 | 21 | 52 | 2974
2020 | 2 | 22 | 53 | 2975 | MM | 60 | * | 40 | zero | off |  | 1 | Y |  | 1600 | 150 | 80 | 8.3 | * | 1780 | 1880 | 1890 | * | * | * | * | off | CH4 Std 1 derlivery regulator psi 37
2020 | 2 | 23 | 54 | 2976
2020 | 2 | 24 | 55 | 2977
2020 | 2 | 25 | 56 | 2978 | MM | 60 | * | 40 | zero | off |  | 1 | Y |  | 1600 | 148 | 60 | 8.5 | * | 1780 | 1890 | 1890 | * | * | * | * | off | Refilled H2 generator with 1L DI water. CH4 Std 1 derlivery regulator psi 17
2020 | 2 | 26 | 57 | 2979
2020 | 2 | 27 | 58 | 2980
2020 | 2 | 28 | 59 | 2981
2020 | 2 | 29 | 60 | 2982 | MM | 60 | * | 40 | zero | off |  | 1 | Y |  | 1590 | 150 | 77 | 8.5 | * | 1770 | 1880 | 1890 | * | * | * | * | off | CH4 Std 1 derlivery regulator psi 17
2020 | 3 | 1 | 61 | 2983
2020 | 3 | 2 | 62 | 2984
2020 | 3 | 3 | 63 | 2985 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1590 | 145 | 80 | 8.5 | * | 1780 | 1880 | 1900 | * | * | * | * | off
2020 | 3 | 4 | 64 | 2986
2020 | 3 | 5 | 65 | 2987
2020 | 3 | 6 | 66 | 2988 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1590 | 125 | 80 | 8.5 | * | 1780 | 1890 | 1900 | * | * | * | * | off
2020 | 3 | 7 | 67 | 2989
2020 | 3 | 8 | 68 | 2990
2020 | 3 | 9 | 69 | 2991 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1590 | 125 | 80 | 8.5 | * | 1780 | 1890 | 1900 | * | * | * | * | off
2020 | 3 | 10 | 70 | 2992
2020 | 3 | 11 | 71 | 2993
2020 | 3 | 12 | 72 | 2994 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1590 | 125 | 85 | 8 | * | 1780 | 1880 | 1890 | * | * | * | * | off
2020 | 3 | 13 | 73 | 2995
2020 | 3 | 14 | 74 | 2996
2020 | 3 | 15 | 75 | 2997
2020 | 3 | 16 | 76 | 2998 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1590 | 125 | 85 | 8.5 | * | 1780 | 1875 | 1890 | * | * | * | * | off
2020 | 3 | 17 | 77 | 2999
2020 | 3 | 18 | 78 | 3000 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1590 | 125 | 85 | 8.5 | * | 1780 | 1870 | 1890 | * | * | * | * | off | Added ~1L DI water to H2 generator. Ran backup H2 generator (with leak) per monthly tasking instructions. Charging the BioDefender flow meter for 24 hrs.
2020 | 3 | 19 | 79 | 3001
2020 | 3 | 20 | 80 | 3002
2020 | 3 | 21 | 81 | 3003
2020 | 3 | 22 | 82 | 3004 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1590 | 125 | 85 | 8.5 | * | 1780 | 1880 | 1890 | * | * | * | * | off
2020 | 3 | 23 | 83 | 3005
2020 | 3 | 24 | 84 | 3006
2020 | 3 | 25 | 85 | 3007
2020 | 3 | 26 | 86 | 3008 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1580 | 125 | 80 | 8.5 | * | 1780 | 1880 | 1890 | * | * | * | * | off
2020 | 3 | 27 | 87 | 3009
2020 | 3 | 28 | 88 | 3010 | MM | 60 | * | 40 | off | off |  | 0.25 | Y |  | 1570 | 125 | 85 | 5 | * | 1780 | 1880 | 1890 | * | * | * | * | off
2020 | 3 | 29 | 89 | 3011 | [23 empty] | CUGC Inlet impacted with blown snow. Attempted to remove but was unable to reach/access all of the inlet. Flows down.
2020 | 3 | 30 | 90 | 3012 | [23 empty] | Climbed TAWO inlet mast at 1330-1350z to inspect CUGC inlet
2020 | 3 | 31 | 91 | 3013
2020 | 4 | 1 | 92 | 3014 | MM | 60 | * | 40 | off | off |  | off | 3/4 |  | 1570 | 125 | 85 | 0 | * | 1750 | 1870 | 1880 | * | * | * | * | off | Inlet iced over. Inlet line flow pump off.
2020 | 4 | 2 | 93 | 3015
2020 | 4 | 3 | 94 | 3016
2020 | 4 | 4 | 95 | 3017 | MM | 60 | * | 40 | off | off |  | 0.75 | Y |  | 1560 | 125 | 85 | 7.5 | * | 1780 | 1880 | 1890 | * | * | * | * | off
2020 | 4 | 5 | 96 | 3018
2020 | 4 | 6 | 97 | 3019
2020 | 4 | 7 | 98 | 3020
2020 | 4 | 8 | 99 | 3021 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1570 | 125 | 85 | 8.5 | * | 1780 | 1890 | 1890 | * | * | * | * | off
2020 | 4 | 9 | 100 | 3022
2020 | 4 | 10 | 101 | 3023
2020 | 4 | 11 | 102 | 3024 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1570 | 125 | 85 | 8.5 | * | 1780 | 1890 | 1890 | * | * | * | * | off
2020 | 4 | 12 | 103 | 3025
2020 | 4 | 13 | 104 | 3026
2020 | 4 | 14 | 105 | 3027 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1570 | 125 | 85 | 8.5 | * | 1780 | 1890 | 1890 | * | * | * | * | off
2020 | 4 | 15 | 106 | 3028
2020 | 4 | 16 | 107 | 3029
2020 | 4 | 17 | 108 | 3030 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1570 | 125 | 85 | 8.5 | * | 1780 | 1890 | 1890 | * | * | * | * | off | Ran backup H2 generator (with leak) per monthly tasking instructions.
2020 | 4 | 18 | 109 | 3031
2020 | 4 | 19 | 110 | 3032
2020 | 4 | 20 | 111 | 3033 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1570 | 125 | 85 | 8.5 | * | 1780 | 1890 | 1890 | * | * | * | Grey-green | off | O2 indicator is appears more grey-green as compared to new in-line filter, but not yet as dark green-grey as our spent-dessicant example
2020 | 4 | 21 | 112 | 3034
2020 | 4 | 22 | 113 | 3035
2020 | 4 | 23 | 114 | 3036
2020 | 4 | 24 | 115 | 3037 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1570 | 125 | 80 | 8.5 | * | 1780 | 1890 | 1890 | * | * | * | Grey-green | off
2020 | 4 | 25 | 116 | 3038
2020 | 4 | 26 | 117 | 3039
2020 | 4 | 27 | 118 | 3040 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1570 | 125 | 85 | 8.5 | * | 1780 | 1890 | 1890 | * | * | * | Grey-green | off | Refilled H2 generator with 1L DI water 1442z
2020 | 4 | 28 | 119 | 3041
2020 | 4 | 29 | 120 | 3042
2020 | 4 | 30 | 121 | 3043
2020 | 5 | 1 | 122 | 3044 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1570 | 125 | 85 | 8.5 | * | 1780 | 1890 | 1890 | * | * | * | Grey-green | off
2020 | 5 | 2 | 123 | 3045
2020 | 5 | 3 | 124 | 3046
2020 | 5 | 4 | 125 | 3047 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1570 | 120 | 85 | 8.5 | * | 1780 | 1890 | 1890 | * | * | * | Grey-green | off
2020 | 5 | 5 | 126 | 3048
2020 | 5 | 6 | 127 | 3049
2020 | 5 | 7 | 128 | 3050
2020 | 5 | 8 | 129 | 3051 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1560 | 115 | 85 | 8.5 | * | 1780 | 1890 | 1890 | * | * | * | Grey-green | off | Refilled H2 generator with 1L DI water 1315z
2020 | 5 | 9 | 130 | 3052
2020 | 5 | 10 | 131 | 3053
2020 | 5 | 11 | 132 | 3054 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1560 | 115 | 85 | 8.5 | * | 1780 | 1890 | 1890 | * | * | * | Grey-green | off
2020 | 5 | 12 | 133 | 3055
2020 | 5 | 13 | 134 | 3056
2020 | 5 | 14 | 135 | 3057
2020 | 5 | 15 | 136 | 3058 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1560 | 115 | 85 | 8.5 | * | 1780 | 1890 | 1890 | * | * | * | Grey-green | off
2020 | 5 | 16 | 137 | 3059
2020 | 5 | 17 | 138 | 3060
2020 | 5 | 18 | 139 | 3061
2020 | 5 | 19 | 140 | 3062 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1560 | 115 | 85 | 8.5 | * | 1790 | 1900 | 1900 | * | * | * | Grey-green | off | Tightened sample line nut at the point where the line enters the back of the instrument at 1900z
2020 | 5 | 20 | 141 | 3063
2020 | 5 | 21 | 142 | 3064 | [23 empty] | Ran backup H2 generator (with leak) per monthly tasking instructions. Charging the BioDefender flow meter for 24 hrs.
2020 | 5 | 22 | 143 | 3065
2020 | 5 | 23 | 144 | 3066 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1550 | 115 | 85 | 8.5 | * | 1790 | 1890 | 1890 | * | * | * | Grey-green | off
2020 | 5 | 24 | 145 | 3067
2020 | 5 | 25 | 146 | 3068 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1550 | 115 | 85 | 8.5 | * | 1790 | 1890 | 1890 | * | * | * | Grey-green | off | Refilled H2 generator with 1.5L DI water 1337z
2020 | 5 | 26 | 147 | 3069
2020 | 5 | 27 | 148 | 3070
2020 | 5 | 28 | 149 | 3071 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1550 | 115 | 85 | 8.5 | * | 1790 | 1890 | 1890 | * | * | * | Grey-green | off
2020 | 5 | 29 | 150 | 3072
2020 | 5 | 30 | 151 | 3073
2020 | 5 | 31 | 152 | 3074
2020 | 6 | 1 | 153 | 3075 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1560 | 115 | 85 | 8.5 | * | 1790 | 1890 | 1890 | * | * | * | Grey-green | off
2020 | 6 | 2 | 154 | 3076
2020 | 6 | 3 | 155 | 3077
2020 | 6 | 4 | 156 | 3078
2020 | 6 | 5 | 157 | 3079 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1560 | 115 | 85 | 8.5 | * | 1790 | 1890 | 1890 | * | * | * | Grey-green | off
2020 | 6 | 6 | 158 | 3080
2020 | 6 | 7 | 159 | 3081
2020 | 6 | 8 | 160 | 3082 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1560 | 0? | 85 | 8.5 | * | 1790 | 1890 | 1890 | * | * | * | Grey-green | off
2020 | 6 | 9 | 161 | 3083
2020 | 6 | 10 | 162 | 3084
2020 | 6 | 11 | 163 | 3085
2020 | 6 | 12 | 164 | 3086 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1550 | 0? | 85 | 8 | * | 1790 | 1890 | 1890 | * | * | * | Grey-green | off
2020 | 6 | 13 | 165 | 3087
2020 | 6 | 14 | 166 | 3088
2020 | 6 | 15 | 167 | 3089 | MM | 60 | * | 40 | off | off |  | 1 | Y |  | 1550 | 125 | 90 | 8 | * | 1790 | 1890 | 1890 | * | * | * | Grey-green | off | Refilled H2 generator with 1.5L DI water 1330z. Ran backup H2 generator (with leak) per monthly tasking instructions. Charging the BioDefender flow meter for 24 hrs.
2020 | 6 | 16 | 168 | 3090
2020 | 6 | 17 | 169 | 3091
2020 | 6 | 18 | 170 | 3092
2020 | 6 | 19 | 171 | 3093
2020 | 6 | 20 | 172 | 3094
2020 | 6 | 21 | 173 | 3095
2020 | 6 | 22 | 174 | 3096
2020 | 6 | 23 | 175 | 3097
2020 | 6 | 24 | 176 | 3098
2020 | 6 | 25 | 177 | 3099
2020 | 6 | 26 | 178 | 3100
2020 | 6 | 27 | 179 | 3101
2020 | 6 | 28 | 180 | 3102
2020 | 6 | 29 | 181 | 3103
2020 | 6 | 30 | 182 | 3104
2020 | 7 | 1 | 183 | 3105
2020 | 7 | 2 | 184 | 3106
2020 | 7 | 3 | 185 | 3107
2020 | 7 | 4 | 186 | 3108
2020 | 7 | 5 | 187 | 3109
2020 | 7 | 6 | 188 | 3110
2020 | 7 | 7 | 189 | 3111
2020 | 7 | 8 | 190 | 3112
2020 | 7 | 9 | 191 | 3113
2020 | 7 | 10 | 192 | 3114
2020 | 7 | 11 | 193 | 3115
2020 | 7 | 12 | 194 | 3116
2020 | 7 | 13 | 195 | 3117
2020 | 7 | 14 | 196 | 3118
2020 | 7 | 15 | 197 | 3119
2020 | 7 | 16 | 198 | 3120
2020 | 7 | 17 | 199 | 3121
2020 | 7 | 18 | 200 | 3122
2020 | 7 | 19 | 201 | 3123
2020 | 7 | 20 | 202 | 3124
2020 | 7 | 21 | 203 | 3125
2020 | 7 | 22 | 204 | 3126
2020 | 7 | 23 | 205 | 3127
2020 | 7 | 24 | 206 | 3128
2020 | 7 | 25 | 207 | 3129
2020 | 7 | 26 | 208 | 3130
2020 | 7 | 27 | 209 | 3131
2020 | 7 | 28 | 210 | 3132
2020 | 7 | 29 | 211 | 3133
2020 | 7 | 30 | 212 | 3134
2020 | 7 | 31 | 213 | 3135
2020 | 8 | 1 | 214 | 3136
2020 | 8 | 2 | 215 | 3137
2020 | 8 | 3 | 216 | 3138
2020 | 8 | 4 | 217 | 3139
2020 | 8 | 5 | 218 | 3140
2020 | 8 | 6 | 219 | 3141
2020 | 8 | 7 | 220 | 3142
2020 | 8 | 8 | 221 | 3143
2020 | 8 | 9 | 222 | 3144
2020 | 8 | 10 | 223 | 3145
2020 | 8 | 11 | 224 | 3146
2020 | 8 | 12 | 225 | 3147
2020 | 8 | 13 | 226 | 3148
2020 | 8 | 14 | 227 | 3149
2020 | 8 | 15 | 228 | 3150
2020 | 8 | 16 | 229 | 3151
2020 | 8 | 17 | 230 | 3152
2020 | 8 | 18 | 231 | 3153
2020 | 8 | 19 | 232 | 3154 | MM/JT | 60 | - | 60 | off | off |  | 1 | 4 |  | 1500 | off | 95 | 6.5 | * | 1790 | 1860 | 1890 | [5 empty] | Coordinated with John to run the CUGC and H2 gen, ran overnight.
2020 | 8 | 20 | 233 | 3155 | MM/JT | off | - | off | off | off |  | 1 | 1 |  | 1525 | off | off | 6.5 | * | 1790 | 1860 | 1890 | - | - | - | - | off | Coordinated with John to shutdown the CUGC and H2 gen.
2020 | 8 | 21 | 234 | 3156
2020 | 8 | 22 | 235 | 3157
2020 | 8 | 23 | 236 | 3158
2020 | 8 | 24 | 237 | 3159 | JT | off | - | off | off | off |  | 1 | 1 |  | 1370 | off | off | 6.5 | * | 1800 | 1870 | 1890 | - | * | - | - | off | H2 cylinder regulator needs to be replaced
2020 | 8 | 25 | 238 | 3160 | JT | [9 empty] | 1490 | [12 empty] | After 16 hours NHMC Std 1 reported 1490psi
2020 | 8 | 26 | 239 | 3161 | JT | [9 empty] | 1425 | [12 empty] | Repressurized NMHC std 1 regulator to 1510
2020 | 8 | 27 | 240 | 3162 | JT | off | - | off | off | off |  | 1 | 1 |  | 1475 | off | off | 6 | * | 1790 | 1870 | 1890 | - | * | - | - | - | Unplanned power outage 1136-1137UTC. Continued troubleshooting of NMHC Std 1: opened main cylinder, pressurized regulator to 1510, closed shut off valve, closed main cylinder.
2020 | 8 | 28 | 241 | 3163 | JT | off | - | off | off | off |  | 1 | 1 | [14 empty] | Pressure dropped to 1460 after yesterdays procedure
2020 | 8 | 29 | 242 | 3164 | JT | off | - | off | off | off |  | 1 | 1 |  | 1410 | off | off | 6 | * | 1790 | 1870 | 1890 | - | * | - | - | off | Attempted to retighten main cylinder. Reprusserized regulator to 1510psi
2020 | 8 | 30 | 243 | 3165 | JT | [9 empty] | 1450 | [12 empty] | regulator continues to show a decrease in pressure
2020 | 8 | 31 | 244 | 3166 | JT | off | - | off | off | off |  | 1 | 1 |  | 1410 | off | off | 6 | * | 1790 | 1870 | 1890 | - | * | - | - | off | regulator continues to show a decrease in pressure
2020 | 9 | 1 | 245 | 3167 | JT | [9 empty] | off | [12 empty] | 1390psi , regulator continues to show a decrease in pressure. Stopping regulator leak trouble shooting.
2020 | 9 | 2 | 246 | 3168
2020 | 9 | 3 | 247 | 3169 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 6 | * | 1790 | 1850 | 1890 | - | * | - | - | off
2020 | 9 | 4 | 248 | 3170
2020 | 9 | 5 | 249 | 3171 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 6 | * | 1790 | 1850 | 1890 | - | * | - | - | off
2020 | 9 | 6 | 250 | 3172
2020 | 9 | 7 | 251 | 3173 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 6 | * | 1790 | 1840 | 1890 | - | * | - | - | off
2020 | 9 | 8 | 252 | 3174
2020 | 9 | 9 | 253 | 3175 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 6 | * | 1790 | 1830 | 1890 | - | * | - | - | off | tightened connection with NMHC Std 1 regulator and primary cylinder, leak has appeared to stop.
2020 | 9 | 10 | 254 | 3176
2020 | 9 | 11 | 255 | 3177 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 6 | * | 1790 | 1830 | 1890 | - | * | - | - | off
2020 | 9 | 12 | 256 | 3178
2020 | 9 | 13 | 257 | 3179
2020 | 9 | 14 | 258 | 3180 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 6 | * | 1790 | 1820 | 1890 | - | * | - | - | off
2020 | 9 | 15 | 259 | 3181
2020 | 9 | 16 | 260 | 3182
2020 | 9 | 17 | 261 | 3183 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 6 | * | 1790 | 1830 | 1890 | - | * | - | - | off | NMHC regulator is still leaking
2020 | 9 | 18 | 262 | 3184
2020 | 9 | 19 | 263 | 3185 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 6 | * | 1790 | 1830 | 1890 | - | * | - | - | off
2020 | 9 | 20 | 264 | 3186
2020 | 9 | 21 | 265 | 3187 | JT | 60 | * | 2/5 | off | off |  | 1 | 4 |  | 1500 | off | 105 | 6 | * | 1790 | 1820 | 1890 | - | * | - | - | off | H2 gen and GC run with John began at 1200 UTC, john aeare that NMHC Std1 reg continues to leak
2020 | 9 | 22 | 266 | 3188
2020 | 9 | 23 | 267 | 3189
2020 | 9 | 24 | 268 | 3190 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 6 | * | 1790 | 1820 | 1890 | - | * | - | - | off | H2 gen and GC run with John completed at 1200 UTC
2020 | 9 | 25 | 269 | 3191
2020 | 9 | 26 | 270 | 3192 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1790 | 1820 | 1890 | - | * | - | - | off
2020 | 9 | 27 | 271 | 3193
2020 | 9 | 28 | 272 | 3194 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1790 | 1820 | 1890 | - | * | - | - | off
2020 | 9 | 29 | 273 | 3195
2020 | 9 | 30 | 274 | 3196 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1790 | 1820 | 1890 | - | * | - | - | off | NMHC std 1 is off pressure in regulator is at 1050psi
2020 | 10 | 1 | 275 | 3197
2020 | 10 | 2 | 276 | 3198
2020 | 10 | 3 | 277 | 3199 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1790 | 1820 | 1890 | - | * | - | - | off | NMHC std 1 is off pressure in regulator is at 990psi
2020 | 10 | 4 | 278 | 3200
2020 | 10 | 5 | 279 | 3201 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1790 | 1820 | 1890 | - | * | - | - | off | NMHC std 1 is off pressure in regulator is at 900psi
2020 | 10 | 6 | 280 | 3202
2020 | 10 | 7 | 281 | 3203
2020 | 10 | 8 | 282 | 3204 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1790 | 1820 | 1890 | - | * | - | - | off | 1205-1213 UTC Inlet mast climb all intakes appeared clear and free of frost, NMHC std 1 is off pressure in regulator is at 875psi
2020 | 10 | 9 | 283 | 3205
2020 | 10 | 10 | 284 | 3206 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1790 | 1820 | 1890 | - | * | - | - | off | 1217-1227 shoveled roof
2020 | 10 | 11 | 285 | 3207
2020 | 10 | 12 | 286 | 3208 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1790 | 1820 | 1890 | - | * | - | - | off
2020 | 10 | 13 | 287 | 3209
2020 | 10 | 14 | 288 | 3210
2020 | 10 | 15 | 289 | 3211 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1790 | 1820 | 1890 | - | * | - | - | off
2020 | 10 | 16 | 290 | 3212 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1790 | 1820 | 1890 | - | * | - | - | off | 1048 UTC unplanned power outage lasting about 35s
2020 | 10 | 17 | 291 | 3213 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1790 | 1820 | 1890 | - | * | - | - | off
2020 | 10 | 18 | 292 | 3214
2020 | 10 | 19 | 293 | 3215 | JT | off | - | off | off | off |  | 1 | 4 |  | 1500 | off | 110 | 5.5 | * | 1790 | 1820 | 1890 | - | * | - | - | off | 1200 UTC initiated monthly run with John, H2 water is about 1cm below the FULL line
2020 | 10 | 20 | 294 | 3216 | JT | 60 | 1cm below full line OK | 2/5 | off | off |  | 1 | 4 |  | 1500 | off | 110 | 5.5 | * | 1790 | 1820 | 1890 | - | * | - | - | off | 1226 - 1237 UTC planned power outage
2020 | 10 | 21 | 295 | 3217
2020 | 10 | 22 | 296 | 3218 | JT | 60 | - | 3/5 | off | off |  | 1 | 4 |  | 1500 | off | 110 | 5.5 | * | 1790 | 1820 | 1890 | - | * | - | - | off
2020 | 10 | 23 | 297 | 3219
2020 | 10 | 24 | 298 | 3220 | JT | 60 | - | 2/5 | off | off |  | 1 | 4 |  | 1490 | off | 110 | 5.5 | * | 1790 | 1810 | 1890 | - | * | - | - | off
2020 | 10 | 25 | 299 | 3221 | [23 empty] | Summit Station timezone changed from West Greenland Summer Time (WGST UTC-2) to West Greenland Time (WGT UTC -3)
2020 | 10 | 26 | 300 | 3222 | JT | 60 | - | 2/5 | off | off |  | 1 | 4 |  | 1490 | off | 110 | 5.5 | * | 1790 | 1810 | 1890 | - | * | - | - | off
2020 | 10 | 27 | 301 | 3223
2020 | 10 | 28 | 302 | 3224
2020 | 10 | 29 | 303 | 3225 | JT | 60 | OK | 2/5 | off | off |  | 1 | 4 |  | 1490 | off | 105 | 5.5 | * | 1790 | 1810 | 1890 | - | * | - | - | off | H2 gen water level about 1cm lower than last check 10 days ago
2020 | 10 | 30 | 304 | 3226
2020 | 10 | 31 | 305 | 3227 | JT | 60 | OK | 2/5 | off | off |  | 1 | 4 |  | 1490 | off | 100 | 5.5 | * | 1790 | 1810 | 1890 | - | * | - | - | off
2020 | 11 | 1 | 306 | 3228
2020 | 11 | 2 | 307 | 3229 | JT | 60 | OK | 2/5 | off | off |  | 1 | 4 |  | 1490 | off | 105 | 5.5 | * | 1790 | 1810 | 1890 | - | * | - | - | off
2020 | 11 | 3 | 308 | 3230
2020 | 11 | 4 | 309 | 3231 | JT | 60 | OK | 2/5 | off | off |  | 1 | 4 |  | 1490 | off | 105 | 5.5 | * | 1790 | 1810 | 1890 | - | * | - | - | off | Wind overnight southerly 41kt
2020 | 11 | 5 | 310 | 3232 | JT | 60 | OK | 3/5 | off | off |  | 1 | 4 |  | 1480 | off | 105 | 5.5 | * | 1790 | 1810 | 1890 | - | * | - | - | off | Refilled H2 Gen
2020 | 11 | 6 | 311 | 3233 | [23 empty] | 12:50/12:58Z TAWO inlet mast climb and inspection. All inlets secure and free of frost.
2020 | 11 | 7 | 312 | 3234 | [23 empty] | Did not visit TAWO due to condition 2 blowing snow, strong winds, poor visibility, and worsening weather conditions
2020 | 11 | 8 | 313 | 3235 | JT | 60 | OK | 2/5 | off | off |  | 1 | 4 |  | 1460 | off | 105 | 5.5 | * | 1780 | 1810 | 1890 | - | * | - | - | off
2020 | 11 | 9 | 314 | 3236 | JT | 60 | OK | 2/5 | off | off |  | 1 | 4 |  | 1460 | off | 95 | 5.5 | * | 1780 | 1810 | 1890 | - | * | - | - | off | Wind overnight up to 21m/s southerly
2020 | 11 | 10 | 315 | 3237
2020 | 11 | 11 | 316 | 3238 | JT | 60 | OK | 2/5 | off | off |  | 1 | 4 |  | 1460 | off | 105 | 5.5 | * | 1780 | 1810 | 1890 | - | * | - | -
2020 | 11 | 12 | 317 | 3239 | [23 empty] | did not visit TAWO today due to off station tasking
2020 | 11 | 13 | 318 | 3240 | JT | 60 | OK | 2/5 | off | off |  | 1 | 4 |  | 1450 | off | 105 | 5.5 | * | 1780 | 1805 | 1860 | - | * | - | - |  | did not visit TAWO today due to off station tasking and limited light with Condition 2 windchill
2020 | 11 | 14 | 319 | 3241
2020 | 11 | 15 | 320 | 3242
2020 | 11 | 16 | 321 | 3243 | JT | 60 | OK | 2/5 | off | off |  | 1 | 4 |  | 1450 | off | 100 | 5.5 | * | 1760 | 1805 | 1850 | - | * | - | -
2020 | 11 | 17 | 322 | 3244 |  |  | refill?
2020 | 11 | 18 | 323 | 3245
2020 | 11 | 19 | 324 | 3246 | JT | 60 | OK | 3/5 | off | off |  | 1 | 4 |  | 1450 | off | 100 | 5.5 | * | 1750 | 1805 | 1850 | - | * | - | - | off
2020 | 11 | 20 | 325 | 3247
2020 | 11 | 21 | 326 | 3248 | JT | 60 | added water | 2/5 | off | off |  | 1 | 4 |  | 1450 | off | 100 | 5.5 | * | 1750 | 1805 | 1850 | - | * | - | - | off | water was at top of refill range, now 1cm below full, ran out of DI at TAWO, will resupply
2020 | 11 | 22 | 327 | 3249
2020 | 11 | 23 | 328 | 3250 | JT | 60 | OK | 2/5 | off | off |  | 1 | 4 |  | 1450 | off | 105 | 5.5 | * | 1750 | 1805 | 1850 | - | * | - | - | off | 12:48/12:56Z inlet mast climb to remove Neph and CLAP Inlet, aditional activiety at mast 12:35Z and 13:26Z
2020 | 11 | 24 | 329 | 3251 | [23 empty] | 14:19/1428Z inlet mast climb to reinstall Neph/CLAP inlet, monthly inlet mast inspection also performed, all inlets secure and frost free
2020 | 11 | 25 | 330 | 3252 | JT | 60 | OK | 2/5 | off | off |  | 1 | 4 |  | 1450 | off | 105 | 5.5 | * | 1750 | 1805 | 1850 | - | * | - | - | off
2020 | 11 | 26 | 331 | 3253 | JT | 60 | OK | 2/5 | off | off |  | 1 | 4 |  | 1460 | off | 100 | 5.5 | * | 1750 | 1805 | 1850 | - | * | - | - | off
2020 | 11 | 27 | 332 | 3254 | JT | 60 | OK | 2/5 | off | off |  | off | off |  | 1430 | off | off | off | * | 1730 | 1805 | 1840 | - | * | - | - | off | received CUGC update from John
2020 | 11 | 28 | 333 | 3255 | JT | 60 | OK | 2/5 | off | off |  | 1 | 1 |  | 1430 | off | off | 5.5 | * | 1730 | 1805 | 1840 | - | * | - | - | off | 14:14/14:24Z snow shoveling to clear TAWO roof
2020 | 11 | 29 | 334 | 3256 | JT | [22 empty] | 13:11Z light snow shoveling
2020 | 11 | 30 | 335 | 3257 | JT | 60 | OK | 2/5 | off | off |  | 1 | 1 |  | 1430 | off | off | 5.5 | * | 1710 | 1800 | 1820 | - | * | - | - | off
2020 | 12 | 1 | 336 | 3258
2020 | 12 | 2 | 337 | 3259
2020 | 12 | 3 | 338 | 3260 | JT | 60 | OK | 2/5 | off | off |  | 1 | 1 |  | 1430 | off | off | 5.5 | * | 1710 | 1805 | 1810 | - | * | - | - | off | Plan to go through shut down process
2020 | 12 | 4 | 339 | 3261 | JT | [22 empty] | Did not visit TAWO due to condition 1 visibility due to blowing snow and compunded by winter darkness
2020 | 12 | 5 | 340 | 3262 | JT | [22 empty] | Did not visit TAWO due to condition 1 and 2 visibility due to blowing snow and compunded by winter darkness
2020 | 12 | 6 | 341 | 3263 | JT | 60 | OK | 2/5 | off | off |  | 1 | 1 |  | 1450 | off | off | 5.5 | * | 1760 | 1810 | 1850 | - | * | - | - | off | Turned off NMHC std 1 and unpluged catylyst heater
2020 | 12 | 7 | 342 | 3264 | JT | off | OK/off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1760 | 1810 | 1850 | - | * | - | - | off | 13:15/14:00Z Diagnostic and shut down session with John
2020 | 12 | 8 | 343 | 3265
2020 | 12 | 9 | 344 | 3266 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 |  | 1730 | 1805 | 1830 | - | * | - | - | off
2020 | 12 | 10 | 345 | 3267 | JT | [22 empty] | Did not visit TAWO today due to off station tasking
2020 | 12 | 11 | 346 | 3268
2020 | 12 | 12 | 347 | 3269 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1780 | 1810 | 1850 | - | * | - | - | off | reviewing re-wire of pump
2020 | 12 | 13 | 348 | 3270
2020 | 12 | 14 | 349 | 3271 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1810 | 1850 | - | * | - | - | off
2020 | 12 | 15 | 350 | 3272
2020 | 12 | 16 | 351 | 3273
2020 | 12 | 17 | 352 | 3274 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1810 | 1850 | - | * | - | - | off | began locating parts for new pump power cord
2020 | 12 | 18 | 353 | 3275
2020 | 12 | 19 | 354 | 3276 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1720 | 1805 | 1830 | - | * | - | - | off
2020 | 12 | 20 | 355 | 3277 | [23 empty] | attempted to build power cord for new line purge pump but was missing a tool
2020 | 12 | 21 | 356 | 3278 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1805 | 1850 | - | * | - | - | off | built power cable for new line purge pump
2020 | 12 | 22 | 357 | 3279 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1805 | 1850 | - | * | - | - | off | built power cable for new line purge pump
2020 | 12 | 23 | 358 | 3280 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1805 | 1850 | - | * | - | - | off | 1545/1630Z worked with JO and put new power supply to original line purge pump, restarted computer which had a dead battery reason for unplug unknown, identified serial line for JO diagnostics instrument identification
2020 | 12 | 24 | 359 | 3281 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1805 | 1850 | - | * | - | - | off
2020 | 12 | 25 | 360 | 3282 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1805 | 1840 | - | * | - | - | off
2020 | 12 | 26 | 361 | 3283
2020 | 12 | 27 | 362 | 3284 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1720 | 1805 | 1820 | - | * | - | - | off | Conditon 1, 100m visibility, did not visit TAWO
2020 | 12 | 28 | 363 | 3285 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1760 | 1840 | 1820 | - | * | - | - | off
2020 | 12 | 29 | 364 | 3286 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1760 | 1840 | 1820 | - | * | - | - | off
2020 | 12 | 30 | 365 | 3287
2020 | 12 | 31 | 366 | 3288 | JT | off | - | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1840 | 1810 | - | * | - | - | off
2021 | 1 | 1 | 1 | 3289 | JT | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1810 | 1850 | - | * | * | - | off
2021 | 1 | 2 | 2 | 3290 | JT | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1810 | 1840 | - | * | * | - | off
2021 | 1 | 3 | 3 | 3291
2021 | 1 | 4 | 4 | 3292 | JT | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1805 | 1840 | - | * | * | - | off
2021 | 1 | 5 | 5 | 3293 | JT | [22 empty] | 13:05/13:35Z TAWO inlet mast climb and inspection, all inets secure and frost free as possible, Surface O3 inlet particle filter replaced
2021 | 1 | 6 | 6 | 3294 | JT | [22 empty] | 15:05/15:08Z Snow removal on TAWO roof
2021 | 1 | 7 | 7 | 3295 | JT | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1805 | 1840 | - | * | * | - | off | 13:33/13:38Z Snow removal from TAWO roof
2021 | 1 | 8 | 8 | 3296
2021 | 1 | 9 | 9 | 3297 | JT | off | off | off | 0 | 0 |  | 1 | 1 |  | 1440 | 0 | off | 5.5 | * | 1720 | 1800 | 1820 | - | * | * | - | off | large NMHC Std1 is 1440psi, small NMHC Std1 (on rack) is empty
2021 | 1 | 10 | 10 | 3298
2021 | 1 | 11 | 11 | 3299 | JT | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1740 | 1810 | 1840 | - | * | * | * | off
2021 | 1 | 12 | 12 | 3300
2021 | 1 | 13 | 13 | 3301
2021 | 1 | 14 | 14 | 3302 | JT | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1760 | 1810 | 1860 | - | * | * | * | off
2021 | 1 | 15 | 15 | 3303
2021 | 1 | 16 | 16 | 3304 | JT | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1780 | 1810 | 1880 | - | * | * | * | off
2021 | 1 | 17 | 17 | 3305
2021 | 1 | 18 | 18 | 3306 | JT | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1740 | 1810 | 1840 | - | * | * | * | off
2021 | 1 | 19 | 19 | 3307
2021 | 1 | 20 | 20 | 3308
2021 | 1 | 21 | 21 | 3309 | JT | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1810 | 1850 | - | * | * | * | off
2021 | 1 | 22 | 22 | 3310
2021 | 1 | 23 | 23 | 3311 | JT | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1810 | 1850 | - | * | * | * | off | 14:50/15:20Z snow removal on TAWO roof
2021 | 1 | 24 | 24 | 3312 | JT | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1810 | 1850 | - | * | * | * | off | 12:26/17:47Z Working on TAWO roof to lower NOAA aerosol and CUGC inlets
2021 | 1 | 25 | 25 | 3313 | JT | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1810 | 1850 | - | * | * | * | off
2021 | 1 | 26 | 26 | 3314 | [23 empty] | did not visit TAWO due to Con1 temperature and visibility
2021 | 1 | 27 | 27 | 3315 | JT | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1740 | 1805 | 1840 | - | * | * | * | off | Abbreviated TAWO visit due to challenging and deteriorating conditions conditions
2021 | 1 | 28 | 28 | 3316 | JT | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1740 | 1805 | 1840 | - | * | * | * | off
2021 | 1 | 29 | 29 | 3317 | JT | 60 | full | 0/4 | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1740 | 1805 | 1840 | - | * | * | * | off | 13:00Z Started H2 gen for a 24 hour run, 14:20Z measuring inlet position
2021 | 1 | 30 | 30 | 3318 | JT | 60 | full | 0/4 | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1740 | 1805 | 1840 | - | * | * | * | off | 12:45Z turned off H2 gen, H2 pressure at expected 60psi, flow 4/4 while building pressure not flow once pressure reached
2021 | 1 | 31 | 31 | 3319
2021 | 2 | 1 | 32 | 3320 | JT | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1740 | 1805 | 1840 | - | * | * | * | off
2021 | 2 | 2 | 33 | 3321
2021 | 2 | 3 | 34 | 3322 | JT | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1740 | 1805 | 1840 | - | * | * | * | off
2021 | 2 | 4 | 35 | 3323 | JT | [22 empty] | No TAWO visit today due to blowing snow and 20meter visibility
2021 | 2 | 5 | 36 | 3324 | JT | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1810 | 1850 | - | * | * | * | off
2021 | 2 | 6 | 37 | 3325 | JT | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1810 | 1850 | - | * | * | * | off
2021 | 2 | 7 | 38 | 3326 | JT | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1810 | 1850 | - | * | * | * | off
2021 | 2 | 8 | 39 | 3327
2021 | 2 | 9 | 40 | 3328
2021 | 2 | 10 | 41 | 3329
2021 | 2 | 11 | 42 | 3330 | JT | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1810 | 1850 | - | * | * | * | off
2021 | 2 | 12 | 43 | 3331 | JT | [22 empty] | Did not visit TAWO due to Con1 visibility
2021 | 2 | 13 | 44 | 3332 | JT | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1810 | 1850 | - | * | * | * | off
2021 | 2 | 14 | 45 | 3333 | JT | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1810 | 1850 | - | * | * | * | off | PM visit due to poor AM visibility
2021 | 2 | 15 | 46 | 3334 | JT | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1810 | 1850 | - | * | * | * | off | 13:45/14:43Z Lowered inlet anchors
2021 | 2 | 16 | 47 | 3335
2021 | 2 | 17 | 48 | 3336 | JT | [22 empty] | Flight day no AM TAWO visit, twin otter flight times: land 14:32Z, take off 14:55Z
2021 | 2 | 18 | 49 | 3337 | [23 empty] | Con1 visibility no TAWO visit
2021 | 2 | 19 | 50 | 3338 | JT | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1810 | 1850 | - | * | * | * | off | 13:30Z/13:31Z and 13:44Z/13:45Z two brief planned power outages
2021 | 2 | 20 | 51 | 3339
2021 | 2 | 21 | 52 | 3340 | NB | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1820 | 1840 | * | * | * | * | off | Twin otter ondeck 1403-1438z;
2021 | 2 | 22 | 53 | 3341
2021 | 2 | 23 | 54 | 3342 | NB | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1820 | 1840 | * | * | * | * | off
2021 | 2 | 24 | 55 | 3343
2021 | 2 | 25 | 56 | 3344 | NB | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1820 | 1850 | * | * | * | * | off
2021 | 2 | 26 | 57 | 3345
2021 | 2 | 27 | 58 | 3346 | NB | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1820 | 1850 | * | * | * | * | off | Noted Picarro computer time ~12 min 15 sec behind actual
2021 | 2 | 28 | 59 | 3347
2021 | 3 | 1 | 60 | 3348 | NB | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1740 | 1820 | 1850 | * | * | * | * | off | Picarro computer time remains incorrect.
2021 | 3 | 2 | 61 | 3349
2021 | 3 | 3 | 62 | 3350 | NB | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1820 | 1850 | * | * | * | * | off | Picarro computer time remains incorrect. Tracking time offset here: https://conus.summitcamp.org/mirror/summit/ftp/science/CUGC/
2021 | 3 | 4 | 63 | 3351
2021 | 3 | 5 | 64 | 3352 | NB | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1820 | 1850 | * | * | * | * | off | Picarro computer time remains incorrect.
2021 | 3 | 6 | 65 | 3353
2021 | 3 | 7 | 66 | 3354 | NB | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1820 | 1850 | * | * | * | * | off | Picarro computer time remains incorrect.
2021 | 3 | 8 | 67 | 3355
2021 | 3 | 9 | 68 | 3356 | NB | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1750 | 1820 | 1850 | * | * | * | * | off | Picarro computer time remains incorrect.
2021 | 3 | 10 | 69 | 3357
2021 | 3 | 11 | 70 | 3358 | NB | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1740 | 1820 | 1850 | * | * | * | * | off | Picarro computer time remains incorrect.
2021 | 3 | 12 | 71 | 3359
2021 | 3 | 13 | 72 | 3360 | NB | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1740 | 1820 | 1850 | * | * | * | * | off | Picarro computer time remains incorrect.
2021 | 3 | 14 | 73 | 3361
2021 | 3 | 15 | 74 | 3362 | NB | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1740 | 1820 | 1850 | * | * | * | * | off | Picarro computer time remains incorrect.
2021 | 3 | 16 | 75 | 3363
2021 | 3 | 17 | 76 | 3364 | NB | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1740 | 1820 | 1850 | * | * | * | * | off | Picarro computer time remains incorrect.
2021 | 3 | 18 | 77 | 3365 | [23 empty] | Unplanned power outage 1719-1802UTC. Instrument supported by UPS during this time and was unaffected.
2021 | 3 | 19 | 78 | 3366 | NB | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1740 | 1820 | 1850 | * | * | * | * | off | Planned power outage 1638-1639Z. Instrument supported by UPS during this time and was unaffected
2021 | 3 | 20 | 79 | 3367
2021 | 3 | 21 | 80 | 3368 | NB | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1740 | 1820 | 1850 | * | * | * | * | off | Picarro computer time remains incorrect.
2021 | 3 | 22 | 81 | 3369
2021 | 3 | 23 | 82 | 3370 | NB | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1740 | 1820 | 1850 | * | * | * | * | off | Picarro computer time remains incorrect.
2021 | 3 | 24 | 83 | 3371 | DH | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1720 | 1810 | 1830 | * | * | * | * | off | Picarro computer time remains incorrect.
2021 | 3 | 25 | 84 | 3372
2021 | 3 | 26 | 85 | 3373 | NB | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1720 | 1810 | 1830 | * | * | * | * | off | Picarro computer time remains incorrect.
2021 | 3 | 27 | 86 | 3374
2021 | 3 | 28 | 87 | 3375 | NB | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1720 | 1810 | 1830 | * | * | * | * | off | Picarro computer time remains incorrect.
2021 | 3 | 29 | 88 | 3376
2021 | 3 | 30 | 89 | 3377 | NB | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1720 | 1810 | 1830 | * | * | * | * | off | Picarro computer time remains incorrect.
2021 | 3 | 31 | 90 | 3378
2021 | 4 | 1 | 91 | 3379 | NB | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1720 | 1810 | 1830 | * | * | * | * | off | Picarro computer time remains incorrect.
2021 | 4 | 2 | 92 | 3380
2021 | 4 | 3 | 93 | 3380 | NB | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1720 | 1810 | 1830 | * | * | * | * | off | Picarro computer time remains incorrect.
2021 | 4 | 4 | 94 | 3380
2021 | 4 | 5 | 95 | 3380 | NB | off | off | off | off | off |  | 1 | 1 |  | off | off | off | 5.5 | * | 1720 | 1810 | 1830 | * | * | * | * | off | Picarro computer time remains incorrect.
2021 | 4 | 6 | 96 | 3380
2021 | 4 | 7 | 97 | 3380
2021 | 4 | 8 | 98 | 3380
2021 | 4 | 9 | 99 | 3380
2021 | 4 | 10 | 100 | 3380
2021 | 4 | 11 | 101 | 3380
2021 | 4 | 12 | 102 | 3380
2021 | 4 | 13 | 103
2021 | 4 | 14 | 104
2021 | 4 | 15 | 105
2021 | 4 | 16 | 106
2021 | 4 | 17 | 107
2021 | 4 | 18 | 108
2021 | 4 | 19 | 109
2021 | 4 | 20 | 110
2021 | 4 | 21 | 111
2021 | 4 | 22 | 112
2021 | 4 | 23 | 113
2021 | 4 | 24 | 114
2021 | 4 | 25 | 115
2021 | 4 | 26 | 116
2021 | 4 | 27 | 117
2021 | 4 | 28 | 118
2021 | 4 | 29 | 119
2021 | 4 | 30 | 120
2021 | 5 | 1 | 121
2021 | 5 | 2 | 122
2021 | 5 | 3 | 123
2021 | 5 | 4 | 124
2021 | 5 | 5 | 125
2021 | 5 | 6 | 126
2021 | 5 | 7 | 127
2021 | 5 | 8 | 128
2021 | 5 | 9 | 129
2021 | 5 | 10 | 130
2021 | 5 | 11 | 131
2021 | 5 | 12 | 132
2021 | 5 | 13 | 133
2021 | 5 | 14 | 134
2021 | 5 | 15 | 135
2021 | 5 | 16 | 136
2021 | 5 | 17 | 137
2021 | 5 | 18 | 138
2021 | 5 | 19 | 139
2021 | 5 | 20 | 140
2021 | 5 | 21 | 141
2021 | 5 | 22 | 142
2021 | 5 | 23 | 143
2021 | 5 | 24 | 144
2021 | 5 | 25 | 145
2021 | 5 | 26 | 146
2021 | 5 | 27 | 147
2021 | 5 | 28 | 148
2021 | 5 | 29 | 149
2021 | 5 | 30 | 150
2021 | 5 | 31 | 151
2021 | 6 | 1 | 152
2021 | 6 | 2 | 153
2021 | 6 | 3 | 154
2021 | 6 | 4 | 155
2021 | 6 | 5 | 156
2021 | 6 | 6 | 157
2021 | 6 | 7 | 158
2021 | 6 | 8 | 159
2021 | 6 | 9 | 160
2021 | 6 | 10 | 161
2021 | 6 | 11 | 162
2021 | 6 | 12 | 163
2021 | 6 | 13 | 164
2021 | 6 | 14 | 165
2021 | 6 | 15 | 166
2021 | 6 | 16 | 167
2021 | 6 | 17 | 168
2021 | 6 | 18 | 169
2021 | 6 | 19 | 170
2021 | 6 | 20 | 171
2021 | 6 | 21 | 172
2021 | 6 | 22 | 173
2021 | 6 | 23 | 174
2021 | 6 | 24 | 175
2021 | 6 | 25 | 176
2021 | 6 | 26 | 177
2021 | 6 | 27 | 178
2021 | 6 | 28 | 179
2021 | 6 | 29 | 180
2021 | 6 | 30 | 181
2021 | 7 | 1 | 182
2021 | 7 | 2 | 183
2021 | 7 | 3 | 184
2021 | 7 | 4 | 185
2021 | 7 | 5 | 186
2021 | 7 | 6 | 187
2021 | 7 | 7 | 188
2021 | 7 | 8 | 189
2021 | 7 | 9 | 190
2021 | 7 | 10 | 191
2021 | 7 | 11 | 192
2021 | 7 | 12 | 193
2021 | 7 | 13 | 194
2021 | 7 | 14 | 195
2021 | 7 | 15 | 196
2021 | 7 | 16 | 197
2021 | 7 | 17 | 198
2021 | 7 | 18 | 199
2021 | 7 | 19 | 200
2021 | 7 | 20 | 201
2021 | 7 | 21 | 202
2021 | 7 | 22 | 203
2021 | 7 | 23 | 204
2021 | 7 | 24 | 205
2021 | 7 | 25 | 206
2021 | 7 | 26 | 207
2021 | 7 | 27 | 208
2021 | 7 | 28 | 209
2021 | 7 | 29 | 210
2021 | 7 | 30 | 211
2021 | 7 | 31 | 212
2021 | 8 | 1 | 213
2021 | 8 | 2 | 214
2021 | 8 | 3 | 215
2021 | 8 | 4 | 216
2021 | 8 | 5 | 217
2021 | 8 | 6 | 218
2021 | 8 | 7 | 219
2021 | 8 | 8 | 220
2021 | 8 | 9 | 221
2021 | 8 | 10 | 222
2021 | 8 | 11 | 223
2021 | 8 | 12 | 224
2021 | 8 | 13 | 225
2021 | 8 | 14 | 226
2021 | 8 | 15 | 227
2021 | 8 | 16 | 228
2021 | 8 | 17 | 229
2021 | 8 | 18 | 230
2021 | 8 | 19 | 231
2021 | 8 | 20 | 232
2021 | 8 | 21 | 233
2021 | 8 | 22 | 234
2021 | 8 | 23 | 235
2021 | 8 | 24 | 236
2021 | 8 | 25 | 237
2021 | 8 | 26 | 238
2021 | 8 | 27 | 239
2021 | 8 | 28 | 240
2021 | 8 | 29 | 241
2021 | 8 | 30 | 242
2021 | 8 | 31 | 243
2021 | 9 | 1 | 244
2021 | 9 | 2 | 245
2021 | 9 | 3 | 246
2021 | 9 | 4 | 247
2021 | 9 | 5 | 248
2021 | 9 | 6 | 249
2021 | 9 | 7 | 250
2021 | 9 | 8 | 251
2021 | 9 | 9 | 252
2021 | 9 | 10 | 253
2021 | 9 | 11 | 254
2021 | 9 | 12 | 255
2021 | 9 | 13 | 256
2021 | 9 | 14 | 257
2021 | 9 | 15 | 258
2021 | 9 | 16 | 259
2021 | 9 | 17 | 260
2021 | 9 | 18 | 261
2021 | 9 | 19 | 262
2021 | 9 | 20 | 263
2021 | 9 | 21 | 264
2021 | 9 | 22 | 265
2021 | 9 | 23 | 266
2021 | 9 | 24 | 267
2021 | 9 | 25 | 268
2021 | 9 | 26 | 269
2021 | 9 | 27 | 270
2021 | 9 | 28 | 271
2021 | 9 | 29 | 272
2021 | 9 | 30 | 273
2021 | 10 | 1 | 274
2021 | 10 | 2 | 275
2021 | 10 | 3 | 276
2021 | 10 | 4 | 277
2021 | 10 | 5 | 278
2021 | 10 | 6 | 279
2021 | 10 | 7 | 280
2021 | 10 | 8 | 281
2021 | 10 | 9 | 282
2021 | 10 | 10 | 283
2021 | 10 | 11 | 284
2021 | 10 | 12 | 285
2021 | 10 | 13 | 286
2021 | 10 | 14 | 287
2021 | 10 | 15 | 288
2021 | 10 | 16 | 289
2021 | 10 | 17 | 290
2021 | 10 | 18 | 291
2021 | 10 | 19 | 292
2021 | 10 | 20 | 293
2021 | 10 | 21 | 294
2021 | 10 | 22 | 295
2021 | 10 | 23 | 296
2021 | 10 | 24 | 297
2021 | 10 | 25 | 298
2021 | 10 | 26 | 299
2021 | 10 | 27 | 300
2021 | 10 | 28 | 301
2021 | 10 | 29 | 302
2021 | 10 | 30 | 303
2021 | 10 | 31 | 304
2021 | 11 | 1 | 305
2021 | 11 | 2 | 306
2021 | 11 | 3 | 307
2021 | 11 | 4 | 308
2021 | 11 | 5 | 309
2021 | 11 | 6 | 310
2021 | 11 | 7 | 311
2021 | 11 | 8 | 312
2021 | 11 | 9 | 313
2021 | 11 | 10 | 314
2021 | 11 | 11 | 315
2021 | 11 | 12 | 316
2021 | 11 | 13 | 317
2021 | 11 | 14 | 318
2021 | 11 | 15 | 319
2021 | 11 | 16 | 320
2021 | 11 | 17 | 321
2021 | 11 | 18 | 322
2021 | 11 | 19 | 323
2021 | 11 | 20 | 324
2021 | 11 | 21 | 325
2021 | 11 | 22 | 326
2021 | 11 | 23 | 327
2021 | 11 | 24 | 328
2021 | 11 | 25 | 329
2021 | 11 | 26 | 330
2021 | 11 | 27 | 331
2021 | 11 | 28 | 332
2021 | 11 | 29 | 333
2021 | 11 | 30 | 334
2021 | 12 | 1 | 335
2021 | 12 | 2 | 336
2021 | 12 | 3 | 337
2021 | 12 | 4 | 338
2021 | 12 | 5 | 339
2021 | 12 | 6 | 340
2021 | 12 | 7 | 341
2021 | 12 | 8 | 342
2021 | 12 | 9 | 343
2021 | 12 | 10 | 344
2021 | 12 | 11 | 345
2021 | 12 | 12 | 346
2021 | 12 | 13 | 347
2021 | 12 | 14 | 348
2021 | 12 | 15 | 349
2021 | 12 | 16 | 350
2021 | 12 | 17 | 351
2021 | 12 | 18 | 352
2021 | 12 | 19 | 353
2021 | 12 | 20 | 354
2021 | 12 | 21 | 355
2021 | 12 | 22 | 356
2021 | 12 | 23 | 357
2021 | 12 | 24 | 358
2021 | 12 | 25 | 359
2021 | 12 | 26 | 360
2021 | 12 | 27 | 361
2021 | 12 | 28 | 362
2021 | 12 | 29 | 363
2021 | 12 | 30 | 364
2021 | 12 | 31 | 365

## NIST_2_std
NIST-2 | from Jerry | uncert
ppt
ethane | 497.9 | 2.1
propane | 522.3 | 3.6
propene | 495 | .7/10
i-butane | 521.5 | 0.8
n-butane | 508.4 | 0.8
i-butene | 519.5 | .8/-30
i-pentane | 479.1 | 1
n-pentane | 487.4 | 1
1-pentene | 480.1 | 1.1/-40
hexane | 485.9 | 1.1
benzene | 498.7 | 1.1
n-heptane | 498.1 | 1.1
i-octane | 484.9 | 0.9
octane | 479.9 | 0.9
toluene | 484.1 | 1.3
nonane | 473.4 | 1
o-xylene | 491.5 | 1.1
decane | 489.9 | 0.9

## site_visit_logs
Maintenance Visits Logs
DH arrived on April 24, 2013
trap housing purge 35 scale units
inlet line purge 2.2 scale units
new 12 VDC batteries for trap heat are 26 AH, have 13.06 V when unpacked
old batteries are 2x 12.0 AH, read 13.56 while on charger
old batteries without charger after 15 min: 13.5 V
filled up hydrogen generator with DI water, 1 quart
replaced mol sieve cartridge on hydrogen generator
noticed that trap heat would abort mid into bakeout and then mid into injection
2013-04-25
-47 ded C night in tent
replaced old batteries with new 26AH battery
excess blank flow during sampling at 1.8 V 63 ml per min
switched to ads trap B, water trap A
backflush purge 28 ml per min
dry purge 228 ml per min
Ward changed sampling line inlet filter
std 1 excess flow during sampling 44 ml per min
2013-04-26
at 10.00 turned off TAWO heat venting system to check if GC baseline drift improves
ran BA std over night with increasing sample volumes for linearity test
hydrogen backup tank pressure: 2200 psi
methane flow tests:
std1: 69 ml/min
std2: 68 ml/min
zero: 77 ml/min
was usin ads trp A wtr trp B,
11.30 switched to ads trap B and wtr trap A, started BA std runs
14.30 connected NIST-2 Std to std channel 2
NMHC Std 1 flow 81.5 ml/min
NMHC Std 1 flow 82 ml/min
14.45 switched back to ads trap A, wtrp B
NMHC Std 1 excess flow during sample: 30 ml/min |  | -> spling flow is 51 ml/min
NMHC zero flow: |  | 2000 s volume is 2000/60*51=1700 ml volumetric
15:20 replaced little inlet filter on zero air inlet on roof of TAWO
17.30 reset the fid signal cables. Fid collector electrode of front FID was loose. Tightened the nut snug.
charged both spare 12 AH batteries
NMHC blank flow rate 108 ml/min
GC column is Agilent HP-Al/KCl, 50 m, 0.53 mm, 15 micron
2013-04-27
system ran fine over night doing methane and NMHC runs
did more tests on the FID noise, put dampers cardboard on North side air inlet and plastic foil on East Side inlet.
ANG did not fly into Summit as scheduled due to winds, spending another day
lowered ads trap overheat shutoff temp from 400 to 360 deg C
2013-04-28
trap heat cycle aborted again soon after inject
replaced solid state resettable 30 amp fuse with regular 30 amp fuse
1st heat cycle after that looked very good
room is pretty warm, at 69F, after work on cooling system yesterday. G. worked on fan and damper controls,
reduced speed on fan to less than half of previously
gc peaks are shifted to earlier times
2017-08-18
Jacques and Reiner on site to perform a VOC audit.
Jacques will shut down the CU GC for the winter or leave instructions to do so.
2017-08-25
CU GC turned off for winter ops
discuss with Jacques
no counter in setup step | 300s
nafion dryer, how to change mol sieve
consistent labels
ozone scrubber
need dentistry mirrow
more heat for gc transfer line
labels on plumbing box
spare fuse for trap heat
bring some aligator clip wires for grounding GC chassis
extend gc nmhc run time by one minute