## Neph
Date | Julian Day | Time (UTC) | Observer's Initials | Sample inlet? | Check Daily Aerosol Email | NEPH Lamp On/Current (<7A) | NEPH Flag Status | NEPH Flow (~18 lpm) | CLAP flow ~0.60 lpm | CLAP tr>0.7 | CLAP Spot # | Biweekly Span Check | Comments
2018-01-01 | 1 | 1719 | HA | rm frost | * | 5.7 | 100 | 18 | 0.59 | 0.811 | 3
2018-01-02 | 2 | [11 empty] | Could not visit TAWO due to weather
2018-01-03 | 3 | [11 empty] | Could not visit TAWO due to weather
2018-01-04 | 4 | [11 empty] | Could not visit TAWO due to weather
2018-01-05 | 5 | 1454 | HA | rm hvy frost | * | 5.6 | 100 | 18 | 0.61 | 0.798 | 3
2018-01-06 | 6
2018-01-07 | 7
2018-01-08 | 8 | 1423 | HA | rm frost | * | 5.6 | 100 | 18 | 0.6 | 0.785 | 3 | x | span check completed from 1714-1751 UTC, passed at 2% error. CO2 tank pressure- 600 psi. CPD3 laptop shutdown from 1640-1645.
2018-01-09 | 9
2018-01-10 | 10 | [11 empty] | Flow was adjusted from 18 to 24 on 1/10/2018 at Betsy's request due to AE33 being removed from system.
2018-01-11 | 11
2018-01-12 | 12 | 1440 | HA | rm frost | * | 5.6 | 100 | 24 | 0.54 | 0.801 | 3 |  | Flow was adjusted from 18 to 24 on 1/10/2018 at Betsy's request due to AE33 being removed from system.
2018-01-13 | 13
2018-01-14 | 14 | [11 empty] | Could not visit TAWO due to 30-40 knot winds
2018-01-15 | 15 | 1413 | HA | rm frost | * | 5.6 | 100 | 22 | 0.62 | 0.745 | 3 |  | Adjusted CLAP flow from .62 to .60
2018-01-16 | 16
2018-01-17 | 17
2018-01-18 | 18 | 1407 | HA | rm frost |  | 5.7 | 100 | 24 | 0.65 | 0.741 | 3 |  | Adjusted CLAP flow from .65 to .6 at 1411 UTC
2018-01-19 | 19
2018-01-20 | 20 | 1323 | HA | rm frost | * | 5.6 | 100 | 23 | 0.6 | 0.703 | 3 |  | Unexpected power outage from 18:21-18:22Z and 19:00-19:04Z
2018-01-21 | 21
2018-01-22 | 22 | 1758 | HA | rm frost | * | 5.6 | 100 | 24 | 0.58 | 0.686 | 3 |  | Adjusted CLAP flow from .58 to .60, Neph Span Check
2018-01-23 | 23
2018-01-24 | 24 | 1409 | HA | rm frost | * | 5.6 | 100 | 24 | 0.61 | 0.681 | 3 | x | Neph span check completed from 1507-1542 UTC; error: 1.8%, co2 tank pressure is 600 psi
2018-01-25 | 25
2018-01-26 | 26
2018-01-27 | 27
2018-01-28 | 28
2018-01-29 | 29 | 1432 | HA | rm frost | * | 5.6 | 100 | 22 | 0.58 | 0.673 | 3
2018-01-30 | 30
2018-01-31 | 31 | 1434 | HA | rm frost | * | 5.6 | 100 | 22 | 0.59 | 0.998 | 4 |  | CLAP spot advanced due to low transmittance
2018-02-01 | 32
2018-02-02 | 33
2018-02-03 | 34 | 1411 | HA | rm frost | * | 5.6 | 100 | 22 | 0.56 | 0.991 | 4
2018-02-04 | 35
2018-02-05 | 36 | 1424 | HA | rm frost | - | 5.6 | 100 | 22 | 0.59 | 0.966 | 4
2018-02-06 | 37 | 1428 | HA |  | * | 5.6 | 100 | 22 | 0.6 | 0.899 | 4 | x | Neph span check completed from 1436-1502 UTC; error: 1.8%, CO2 tank pressure 600 psi.
2018-02-07 | 38 | 1612 | HA | rm frost | * | 5.6 | 100 | 22 | 0.59 | 0.892 | 4
2018-02-08 | 39
2018-02-09 | 40 | 1314 | HA | rm frost | * | 5.6 | 100 | 22 | 0.59 | 0.889 | 4
2018-02-10 | 41
2018-02-11 | 42
2018-02-12 | 43 | 1458 | HA, SWD | rm hvy frost | * | 5.6 | 100 | 22 | 0.61 | 0.883 | 4
2018-02-13 | 44
2018-02-14 | 45 | 1418 | SWD | rm frost | * | 5.5 | 100 | 22 | 0.61 | 0.881 | 4
2018-02-15 | 46 |  |  |  | *
2018-02-16 | 47 |  |  |  | *
2018-02-17 | 48 | 1353 | SWD | rm frost | * | 5.5 | 100 | 22 | 0.64 | 0.871 | 4 |  | Adjusted CLAP flow to 0.61 1355z.
2018-02-18 | 49 |  |  |  | *
2018-02-19 | 50 |  |  |  | *
2018-02-20 | 51 |  |  |  | *
2018-02-21 | 52 | 1428 | SWD | rm frost | * | 5.5 | 100 | 22 | 0.61 | 0.87 | 4 | x | Neph span check completed from 1429-1503Z; error: 2.1%, CO2 tank pressure 630 psi.
2018-02-22 | 53 |  |  |  | *
2018-02-23 | 54 |  |  |  | *
2018-02-24 | 55 |  |  |  | *
2018-02-25 | 56 | 1503 | SWD | rm frost | * | 5.5 | 100 | 22 | 0.61 | 0.858 | 4
2018-02-26 | 57 | 1539 | SWD | clr | * | 5.5 | 100 | 22 | 0.61 | 0.856 | 4
2018-02-27 | 58 | 1851 | SWD | rm frost | * | 5.5 | 100 | 22 | 0.62 | 0.854 | 4 |  | Adjusted CLAP flow to 0.61 1855z.
2018-02-28 | 59 |  |  |  | *
2018-03-01 | 60 |  |  |  | *
2018-03-02 | 61 | 1657 | SWD | rm frost | * | 5.5 | 0 | 22 | 0.61 | 0.837 | 4
2018-03-03 | 62 |  |  |  | *
2018-03-04 | 63 |  |  |  | *
2018-03-05 | 64 | 1439 | SWD | rm frost | * | 5.5 | 100 | 22 | 0.61 | 0.756 | 4 | x | Neph span check completed from 14:43-15:19Z; error: 2.0%, CO2 tank pressure 680 psi. Adjusted CLAP flow to 0.61 lpm 1442z.
2018-03-06 | 65 |  |  |  | *
2018-03-07 | 66 | 1517 | SWD | rm frost | * | 5.5 | 100 | 22 | 0.61 | 0.982 | 5 |  | Adjusted CLAP flow to 0.61 1518z.
2018-03-08 | 67 |  |  |  | *
2018-03-09 | 68 |  |  |  | *
2018-03-10 | 69 | 1243 | SWD | rm frost | * | 5.5 | 100 | 22 | 0.61 | 0.971 | 5
2018-03-11 | 70 |  |  |  | *
2018-03-12 | 71 | 1454 | SWD | rm frost | * | 5.5 | 100 | 22 | 0.61 | 0.963 | 5 |  | Adjusted CLAP flow to 0.61 1457z.
2018-03-13 | 72 |  |  |  | *
2018-03-14 | 73 | 1529 | SWD | rm hf | * | 5.5 | 100 | 22 | 0.61 | 0.944 | 5 |  | Active pump 'Dignan' discovered to have failed (14:42z). Swapped to pump 'Colgan', repowered, and adjusted flow rates at 15:01z.
2018-03-15 | 74 |  |  |  | *
2018-03-16 | 75 | 1244 | SWD | rm lf | * | 5.5 | 100 | 22 | 0.61 | 0.906 | 5
2018-03-17 | 76 |  |  |  | *
2018-03-18 | 77 |  |  |  | *
2018-03-19 | 78 | 1459 | SWD | rm hf | * | 5.4 | 0 | 22 | 0.61 | 0.857 | 5 | x | Neph span check completed from 15:10Z to 15:41Z; error:1.8%, CO2 tank pressure 710 psi. Adjusted CLAP flow to 0.61 lpm 1500Z
2018-03-20 | 79 |  |  |  | *
2018-03-21 | 80 | 1320 | SWD | rm lf | * | 5.4 | 100 | 22 | 0.61 | 0.852 | 5 |  | Adjusted CLAP flow to 0.61 1324z.
2018-03-22 | 81 |  |  |  | *
2018-03-23 | 82 | 1300 | SWD/RN | rm lf | * | 5.4 | 100 | 22 | 0.61 | 0.847 | 5 |  | Adjusted CLAP flow to 0.61 1302z.
2018-03-24 | 83 |  |  |  | *
2018-03-25 | 84 | 1227 | SWD | rm lf | * | 5.4 | 100 | 22 | 0.62 | 0.841 | 5 |  | Adjusted CLAP flow to 0.61 1230z.
2018-03-26 | 85 |  |  |  | *
2018-03-27 | 86 |  |  |  | *
2018-03-28 | 87 | 1503 | NB | rm hf | * | 5.4 | 100 | 22 | 0.61 | 0.804 | 5
2018-03-29 | 88 |  |  |  | *
2018-03-30 | 89 | 1722 | NB | rm lf | * | 5.4 | 100 | 22 | 0.61 | 0.781 | 5
2018-03-31 | 90 |  |  |  | *
2018-04-01 | 91 | 1157 | NB | rm hf | * | 5.7 | 100 | 22 | 0.6 | 0.76 | 5
2018-04-02 | 92 | 1306 | NB | rm lf | * | 5.7 | 100 | 22 | 0.6 | 0.752 | 5 | x | Neph span check completed from 1310-1342 z; error 1.5 %, CO2 tank pressure 690 psi.
2018-04-03 | 93 |  |  |  | *
2018-04-04 | 94 |  |  |  | *
2018-04-05 | 95 | 1203 | NB | rm lf | * | 5.6 | 100 | 22 | 0.58 | 0.717 | 5 |  | Adjusted CLAP flow to 0.60 1208z.
2018-04-06 | 96 |  |  |  | *
2018-04-07 | 97 | 1122 | NB | rm lf | * | 5.7 | 100 | 22 | 0.59 | 0.691 | 5 |  | Adjusted CLAP flow to 0.60 1123z.
2018-04-08 | 98 |  |  |  | *
2018-04-09 | 99 | 1135 | NB | rm lf | * | 5.6 | 100 | 22 | 0.63 | 0.924 | 6 |  | Adjusted CLAP flow to 0.60 1226z.
2018-04-10 | 100 |  |  |  | *
2018-04-11 | 101 |  |  |  | *
2018-04-12 | 102 | 1130 | NB | rm lf | * | 5.6 | 100 | 22 | 0.61 | 0.901 | 6
2018-04-13 | 103 |  |  |  | *
2018-04-14 | 104 | 1314 | NB | clr | * | 5.6 | 0 | 22 | 0.62 | 0.894 | 6 |  | Adjusted CLAP flow to 0.60 1316z.
2018-04-15 | 105 |  |  |  | *
2018-04-16 | 106 | 1127 | NB | rm lf | * | 5.6 | 100 | 22 | 0.59 | 0.877 | 6 | x | Neph span check completed from 1215-1245z. Error 1.1%. CO2 tank pressure 700 psi. Adjusted CLAP flow to 0.60 at 1248z.
2018-04-17 | 107 |  |  |  | *
2018-04-18 | 108 | 1145 | NB | rm hf | * | 5.6 | 100 | 22 | 0.6 | 0.865 | 6
2018-04-19 | 109 |  |  |  | *
2018-04-20 | 110 | 1141 | NB | rm lf | * | 5.6 | 0 | 22 | 0.61 | 0.854 | 6
2018-04-21 | 111
2018-04-22 | 112 | 1207 | NB | rm lf |  | 5.8 | 100 | 22 | 0.62 | 0.882 | 7 |  | Adjusted CLAP flow to .60 at 1210z.
2018-04-23 | 113 | 1144 | NB | rm hf |  | 5.8 | 100 | 22 | 0.6 | 0.857 | 7
2018-04-24 | 114 |  |  |  | *
2018-04-25 | 115 |  |  |  | *
2018-04-26 | 116 | 1130 | NB | rm lf | * | 5.9 | 100 | 22 | 0.6 | 0.781 | 7
2018-04-27 | 117 | 1327 | NB/MB | rm lf | * | 5.8 | 100 | 23 | 0.6 | 0.769 | 7
2018-04-28 | 118 |  |  |  | *
2018-04-29 | 119 |  |  |  | *
2018-04-30 | 120 | 1228 | NB/MB | clr | * | 5.8 | 100 | 23 | 0.59 | 0.689 | 7 | x | CLAP flag 0004. Advanced CLAP to spot 8 at 1245z. Adjusted CLAP flow to 0.60 lpm at 1305z. Neph span check completed from 1310-1343z; error 2.6%; Tank pressure 620 psi.
2018-05-01 | 121 | 1224 | MB | rm lf | * | 5.8 | 100 | 23 | 0.6 | 0.975 | 8
2018-05-02 | 122 |  |  |  | * | [7 empty] | Aerosol email: "Neph lamp power not within 10 percent of the setpoint for 1% (8 minute(s)) of data"
2018-05-03 | 123 |  |  |  | * | [7 empty] | Aerosol email: "Neph lamp power not within 10 percent of the setpoint for 1% (9 minute(s)) of data"
2018-05-04 | 124 | 1227 | MB | clr | * | 5.9 | 100 | 23 | 0.58 | 0.92 | 8 |  | Connection Lost error message on Aerosol laptop. Reboot computer ~1200z. Data appears normal after reboot. Adjusted CLAP flow to 0.60 at 1233z. Aerosol email: "Neph lamp power not within 10 percent of the setpoint for 1% (8 minute(s)) of data"
2018-05-05 | 125 | 1217 | MB | rm lf | * | 5.8 | 100 | 23 | 0.6 | 0.92 | 8
2018-05-06 | 126 |  |  |  | *
2018-05-07 | 127 | 1224 | MB | clr | * | 5.6 | 100 | 23 | 0.61 | 0.865 | 8
2018-05-08 | 128 | 1247 | MB | rm lf | * | 5.7 | 100 | 23 | 0.61 | 0.858 | 8
2018-05-09 | 129 |  |  |  | *
2018-05-10 | 130 | 1247 | MB | rm lf | * | 5.7 | 100 | 23 | 0.62 | 0.839 | 8 |  | Adjusted CLAP flow to 0.60 lpm at 1252z.
2018-05-11 | 131 |  |  |  | *
2018-05-12 | 132 | 1721 | MB | clr | * | 5.8 | 0 | 23 | 0.61 | 0.814 | 8
2018-05-13 | 133 |  |  |  | *
2018-05-14 | 134 | 1153 | MB | clr | * | 5.9 | 100 | 23 | 0.61 | 0.792 | 8 | x | Neph span check completed from 1305-1336z. Error 2.2%. CO2 tank pressure 700 psi.
2018-05-15 | 135 | 1515 | MB | rm hf | * | 5.8 | 100 | 23 | 0.61 | 0.749 | 8 |  | CLAP filter change 1442-1454z.
2018-05-16 | 136 |  |  |  | *
2018-05-17 | 137 | 1205 | MB | rm lf | * | 5.8 | 0 | 23 | 0.61 | 0.83 | 1
2018-05-18 | 138 |  |  |  | *
2018-05-19 | 139 | 1219 | MB | clr | * | 5.7 | 100 | 23 | 0.6 | 0.819 | 1
2018-05-20 | 140 |  |  |  | *
2018-05-21 | 141 | 1658 | MB | rm lf | * | 5.8 | 100 | 23 | 0.6 | 0.803 | 1
2018-05-22 | 142 | 1859 | MB | rm lf | * | 6.1 | 100 | 23 | 0.61 | 0.793 | 1
2018-05-23 | 143 | 1243 | MB | rm lf | * | 6.1 | 100 | 23 | 0.61 | 0.788 | 1
2018-05-24 | 144 | 1226 | MB | clr | * | 6.1 | 100 | 23 | 0.61 | 0.777 | 1
2018-05-25 | 145 |  |  |  | *
2018-05-26 | 146 |  |  |  | *
2018-05-27 | 147 |  |  |  | *
2018-05-28 | 148 | 1220 | MB | clr | * | 6.2 | 100 | 23 | 0.61 | 0.763 | 1 | x | Neph span check completed from 1309-1340z. Error 2.2%. CO2 tank pressure 650 psi.
2018-05-29 | 149 | 1347 | MB | clr | * | 6.2 | 100 | 23 | 0.62 | 0.754 | 1 |  | Adjusted CLAP flow to 0.60 lpm at 1349z.
2018-05-30 | 150 | 1201 | MB | clr | * | 6.6 | 0 | 23 | 0.6 | 0.747 | 1
2018-05-31 | 151 |  |  |  | *
2018-06-01 | 152 | 1129 | MB | clr | * | 6.5 | 0 | 23 | 0.6 | 0.735 | 1
2018-06-02 | 153 |  |  |  | *
2018-06-03 | 154 | 1203 | MB | clr | * | 6.6 | 0 | 23 | 0.6 | 0.718 | 1 |  | Neph lamp changed 12:05-12:31z. Lamp current 5.5 A after change.
2018-06-04 | 155 |  |  |  | *
2018-06-05 | 156 | 1302 | MB/MM | clr | * | 5.7 | 0 | 23 | 0.608 | 0.696 | 1 |  | Adjusted CLAP flow to 0.60 lpm for turnover training purposes at 1332z
2018-06-06 | 157 | 1341 | MB/MM | clr | * | 5.7 | 0 | 23 | 0.6 | 0.682 | 1
2018-06-07 | 158 | 1209 | MB/MM | clr | * | 5.7 | 0 | 22 | 0.6 | 0.671 | 1
2018-06-08 | 159 | 1220 | MM | clr | * | 5.7 | 0 | 23 | 0.59 | 0.6668 | 1
2018-06-09 | 160 | 1647 | MM | clr | * | 5.7 | 0 | 23 | 0.602 | 0.95 | 2 |  | Vestibule of TAWO felt warm in the afternoon due to warm outdoor ambient temperatures, Aerosol pump felt warm to touch, set Solo controller from 16C to 15C at 1820z
2018-06-10 | 161 |  |  |  | *
2018-06-11 | 162 | 1659 | MB/MM | clr | * | 5.7 | 0 | 23 | 0.61 | 0.942 | 2 | x | Neph span check completed from 1706-1740z. Error 1.9%. CO2 tank pressure 700 psi.
2018-06-12 | 163 | 1802 | MM | rm hf | * | 5.7 | 0 | 23 | 0.604 | 0.9363 | 2
2018-06-13 | 164 |  |  |  | *
2018-06-14 | 165 |  |  |  | *
2018-06-15 | 166 | 1828 | MM | rm lf | * | 5.7 | 0 | 23 | 0.599 | 0.8636 | 2 |  | The aerosol pump was off during a planned power outage from 1655z-1859z, NEPH and CLAP stayed running on the UPS
2018-06-16 | 167 | 1746 | MM | clr | * | 5.7 | 0 | 23 | 0.593 | 0.803 | 2
2018-06-17 | 168 |  |  |  | *
2018-06-18 | 169 |  |  |  | * | [7 empty] | The aerosol pump was off during during two brief planned power outages from 0115z-0130z and 0645z-0650z, NEPH and CLAP stayed running on the UPS
2018-06-19 | 170 | 1326 | MM | clr | * | 5.7 | 0 | 23 | 0.6 | 0.768 | 2
2018-06-20 | 171 | 2023 | MM | clr | * | 5.7 | 0 | 23 | 0.597 | 0.753 | 2
2018-06-21 | 172 |  |  |  | *
2018-06-22 | 173 | 1358 | MM | clr | * | 5.7 | 0 | 23 | 0.6 | 0.732 | 2 |  | Adjusted CLAP flow to 0.60 lpm at 1402z
2018-06-23 | 174 |  |  |  | *
2018-06-24 | 175 |  |  |  | *
2018-06-25 | 176 | 1802 | MM | rm lf | * | 5.7 | 0 | 23 | 0.599 | 0.726 | 2
2018-06-26 | 177 | 1849 | MM | clr | * | 5.7 | 0 | 23 | 0.597 | 0.73237 | 2 | x | Neph span check completed from 1836-1858z. Error 2.3%. CO2 tank pressure 710 psi.
2018-06-27 | 178 | 1649 | MM | clr | * | 5.7 | 0 | 23 | 0.599 | 0.699 | 2 |  | Adjusted CLAP flow to 0.60 lpm at 1652z
2018-06-28 | 179 | 1558 | MM | clr | * | 5.7 | 0 | 23 | 0.603 | 0.6933 | 2
2018-06-29 | 180 |  |  |  | *
2018-06-30 | 181 | 2047 | MM | clr | * | 5.7 | 0 | 23 | 0.604 | 0.6853 | 2
2018-07-01 | 182 |  |  |  | *
2018-07-02 | 183 |  |  |  | *
2018-07-03 | 184 | 1603 | MM | clr | * | 5.7 | 0 | 23 | 0.601 | 0.6785 | 2
2018-07-04 | 185 |  |  |  | *
2018-07-05 | 186 | 1424 | MM | clr | * | 5.7 | 0 | 23 | 0.596 | 0.673 | 2
2018-07-06 | 187 |  |  |  | *
2018-07-07 | 188 |  |  |  | *
2018-07-08 | 189 |  |  |  | *
2018-07-09 | 190 |  |  |  | *
2018-07-10 | 191 | 1223 | MM | clr | * | 5.7 | 0 | 23 | 0.603 | 0.989 | 3 | x | Neph span check completed from 1244-1320z. Error 2.8%. CO2 tank pressure 700 psi.
2018-07-11 | 192 | 1811 | MM | clr | * | 5.7 | 0 | 23 | 0.603 | 0.9834 | 3
2018-07-12 | 193 |  |  |  | *
2018-07-13 | 194 |  |  |  | *
2018-07-14 | 195 |  |  |  | *
2018-07-15 | 196 |  |  |  | *
2018-07-16 | 197 | 2209 | MM | clr | * | 5.7 | 0 | 23 | 0.596 | 0.938 | 3
2018-07-17 | 198 |  |  |  | *
2018-07-18 | 199 | 1430 | MM | clr | * | 5.7 | 0 | 23 | 0.604 | 0.846 | 3
2018-07-19 | 200 |  |  |  | *
2018-07-20 | 201 |  |  |  | *
2018-07-21 | 202 | 1317 | MM | rm lf | * | 5.7 | 0000/0100 | 23 | 0.6 | 0.767 | 3 |  | NEPH flag blinking between 0000 and 0100: Checked TSI manual for neph which idicated the 0100 flag indicates a chopper fault. This error did not show up in the daily aerosol email.
2018-07-22 | 203 |  |  |  | *
2018-07-23 | 204 |  |  |  | * | [7 empty] | Planned power outage from 2153-2300z and again from 2307-2315z. The NEPH and CLAP remained on while the aerosol pumps and the TEI were off. Internt was also down during the outage.
2018-07-24 | 205 |  |  |  | *
2018-07-25 | 206 | 2335 | MM | rm lf | * | 5.7 | 0 | 23 | 0 | - | 0 | x | Neph span check completed from 2250-2325z. Error 3.8%. CO2 tank pressure 750 psi. CLAP discovered to have no flow and read SPOT 0 and FILTER CHANGE REQUIRED. Planned power outage from 2114-2139z, NEPH and CLAP remained on while the aerosol pump and TEI O3 instrument were off
2018-07-26 | 207 |  |  |  | *
2018-07-27 | 208 | 204 | MM | clr | * | 5.7 | 0 | 23 | 0.6 | 1 | 1 |  | Changed the CLAP filter to address the FILTER CHANGE REQUIRED directive after the power outage
2018-07-28 | 209 |  |  |  | * | [7 empty] | TAWO planned power outage from 1218-1228z; Used simple green and rubbing alcohol in the TAWO vestibule
2018-07-29 | 210 | 258 | MM | rm hf | * | 5.7 | 0 | 23 | 0.06 | 0.978 | 1 |  | Unplanned power outage from 0231-0246z. Adjusted CLAP flow from .61 to .60
2018-07-30 | 211 | 1234 | MM | clr | * | 5.7 | 0 | 23 | 0.598 | 0.9366 | 1
2018-07-31 | 212 |  |  |  | *
2018-08-01 | 213 |  |  |  | *
2018-08-02 | 214 | 2147 | MM | clr | * | 5.7 | 0 | 23 | 0.6 | 0.876 | 1 |  | Computer was off during planned power outage from 1458-1646z Adjusted CLAP flow from .58 to .60
2018-08-03 | 215 | 1810 | MM | clr | * | 5.7 | 0 | 23 | 0.6 | 0.874 | 1 |  | Planned power outage from 1458-1740z. TEI, NEPH and CLAP and CPD computer powered off during the outage. Adjusted CLAP flow from .61 to .60
2018-08-04 | 216 | 1430 | MM | rm lf | * | 5.7 | 0 | 23 | 0.6 | 1.0007 | 2 |  | Planned power outage from 1030-1402z. TEI, NEPH and CLAP and CPD computer powered off during the outage.
2018-08-05 | 217 |  |  |  | *
2018-08-06 | 218 |  |  |  | *
2018-08-07 | 219 | 2037 | MM | rm lf | * | 5.7 | 0 | 23 | 0.604 | 0.9979 | 3 | x | Planned power outage from 1028-1312z. TEI, NEPH and CLAP and CPD computer powered off during the outage. Neph span check completed from 2006-2041z. Error 1.6%. CO2 tank pressure 700 psi.
2018-08-08 | 220 |  |  |  | *
2018-08-09 | 221 | 1524 | MM | rm hf | * | 5.7 | 0 | 23 | 0.598 | 0.959 | 3 |  | Brief planned power outage during which the TEI was turned off from 1035-1038z and the aerosol pump was off
2018-08-10 | 222 |  |  |  | * | [7 empty] | Planned power outage from 1700-1702z, TEI and aerosol pumps off
2018-08-11 | 223 |  |  |  | *
2018-08-12 | 224 |  |  |  | *
2018-08-13 | 225 | 2200 | MM | rm hf | * | 5.7 | 0 | 23 | 0.6 | 0.904 | 3 |  | Adjusted CLAP flow from .61-.60
2018-08-14 | 226 |  |  |  | *
2018-08-15 | 227 |  |  |  | * | [7 empty] | Planned power outage from 1700-1702z, TEI and aerosol pumps off
2018-08-16 | 228 |  |  |  | *
2018-08-17 | 229 | 1614 | MM | rm hf | * | 5.7 | 0 | 23 | 0.6 | 0.724 | 3 |  | Planned power outage from 1220-1222z, TEI and aerosol pumps off
2018-08-18 | 230 |  |  |  | *
2018-08-19 | 231 |  |  |  | *
2018-08-20 | 232 |  |  |  | *
2018-08-21 | 233 | 1746 | MM | rm lf | * | 5.7 | 0 | 23 | 0.6 | 0.6804 | 3 | x | Neph span check completed from 1710-1741z. Error 2.7%. CO2 tank pressure 710 psi. Adjusted CLAP flow from .58 to .60 at 1750z
2018-08-22 | 234 |  |  |  | *
2018-08-23 | 235 |  |  |  | *
2018-08-24 | 236 |  |  |  | *
2018-08-25 | 237 | 1930 | MM | rm hf | * | 5.7 | 0 | 23 | 0.6 | 0.949 | 4 |  | Adjusted CLAP flow from .63 to .60
2018-08-26 | 238 | 1747 | MM | clr | * | 5.7 | 0 | 23 | 0.6 | 0.935 | 4 |  | Climbed TAWO inlet tower at 1740 to change CUGC inlet filter
2018-08-27 | 239 | 1251 | MM | clr | * | 5.7 | 100 | 23 | 0.6 | 0.897 | 4 |  | Two brief planned power outages during which the aerosol pump was turned off-- 1050-1054 and 1257-1302. Neph flag error a result of the pumps being off briefly while instrument was running. Adjusted CLAP flow from .58 to .60
2018-08-28 | 240 |  |  |  | *
2018-08-29 | 241 |  |  |  | *
2018-08-30 | 242 |  |  |  | *
2018-08-31 | 243 | 1433 | MM | clr | * | 5.7 | 0 | 23 | 0.6 | 0.828 | 4 |  | Adjusted CLAP flow from .59 to .60
2018-09-01 | 244 |  |  |  | *
2018-09-02 | 245 |  |  |  | *
2018-09-03 | 246 |  |  |  | *
2018-09-04 | 247 | 1613 | MM | rm fr | * | 5.7 | 100 | 23 | 0.6 | 0.7249 | 4 | x | Neph span check completed from 1520-1550z. Error 2.1%. CO2 tank pressure 710 psi.
2018-09-05 | 248 | 1700 | MM |  | * | 5.7 | 0 | 23 | 0.6 | 0.701 | 4
2018-09-06 | 249 |  |  |  | *
2018-09-07 | 250 |  |  |  | *
2018-09-08 | 251 |  |  |  | *
2018-09-09 | 252 | 1632 | MM | clr | * | 5.7 | 0 | 23 | 0.6 | 0.671 | 4 |  | Adjusted CLAP flow from .59 to .60
2018-09-10 | 253 |  |  |  | *
2018-09-11 | 254 |  |  |  | *
2018-09-12 | 255 | 1414 | MM | clr | * | 5.7 | 0 | 23 | 0.6 | 0.9676 | 5 |  | Adjusted CLAP flow from .62 to .60
2018-09-13 | 256 |  |  |  | *
2018-09-14 | 257 |  |  |  | *
2018-09-15 | 258 | 1950 | MM | rm hf | * | 5.7 | 100 | 23 | 0.6 | 0.952 | 5
2018-09-16 | 259 |  |  |  | *
2018-09-17 | 260 |  |  |  | *
2018-09-18 | 261 | 1445 | MM | rm hf | * | 5.7 | 100 | 23 | 0.6 | 0.9223 | 5 | x | Neph span check completed from 1509-1550z. Error 1.8%. CO2 tank pressure 690 psi.
2018-09-19 | 262 |  |  |  | *
2018-09-20 | 263 |  |  |  | *
2018-09-21 | 264 | 1900 | MM | rm hf | * | 5.7 | 100 | 23 | 0.6 | 0.9033 | 5
2018-09-22 | 265 |  |  |  | *
2018-09-23 | 266 |  |  |  | *
2018-09-24 | 267 | 1925 | MM | rm lf | * | 5.6 | 100 | 23 | 0.6 | 0.884 | 5
2018-09-25 | 268 |  |  |  | *
2018-09-26 | 269 |  |  |  | *
2018-09-27 | 270 |  |  |  | *
2018-09-28 | 271 | 2000 | MM | clr | * | 5.6 | 100 | 23 | 0.6 | 0.858 | 5
2018-09-29 | 272 |  |  |  | *
2018-09-30 | 273 |  |  |  | *
2018-10-01 | 274 |  |  |  | *
2018-10-02 | 275 | 1937 | MM | rm lf | * | 5.6 | 100 | 23 | 0.6 | 0.6995 | 5 | x | Neph span check completed from 1906-1940z. Error 1.5%. CO2 tank pressure 710 psi.
2018-10-03 | 276 |  |  |  | *
2018-10-04 | 277 |  |  |  | *
2018-10-05 | 278 | 1747 | MM | rm lf | * | 5.6 | 100 | 23 | 0.6 | 0.8457 | 6
2018-10-06 | 279 |  |  |  | *
2018-10-07 | 280 |  |  |  | *
2018-10-08 | 281 | 1843 | MM | rm hf | * | 5.6 | 100 | 23 | 0.6 | 0.787 | 6 |  | Adjusted CLAP flow from .59 to .60
2018-10-09 | 282 |  |  |  | *
2018-10-10 | 283 |  |  |  | * | [7 empty] | Replaced O3 Analyzer inlet filter on TAWO mast from 1631-1645z.
2018-10-11 | 284 |  |  |  | *
2018-10-12 | 285 | 1124 | MM | rm lf | * | 5.6 | 100 | 23 | 0.6 | 0.751 | 6 |  | Adjusted SOLO controller from 15 to 16 C at 1046z
2018-10-13 | 286 |  |  |  | *
2018-10-14 | 287 |  |  |  | *
2018-10-15 | 288 | 1800 | MM | rm lf | * | 5.6 | 100 | 23 | 0.6 | 0.695 | 6 |  | Adjusted CLAP flow from .58 to .60
2018-10-16 | 289 | 1750 | MM/TN | rm hf | * | 5.6 | 100 | 22.5 | 0.6 | 0.693 | 6 | x | Neph span check completed from 1616-16570z. Error 1.9%. CO2 tank pressure 700 psi. Adjusted CLAP flow from .61 to .60
2018-10-17 | 290 |  |  |  | *
2018-10-18 | 291 | 1755 | MM/TN | rm lf | * | 5.6 | 100 | 22.25 | 0.6 | 0.683 | 6 |  | Adjusted CLAP flow from .58 to .60
2018-10-19 | 292 | 1848 | MM/TN | rm lf | * | 5.6 | 100 | 22.3 | 0.6 | 0.678 | 6 |  | Planned power outage from 1538-1541, and again from 1602-1606, TEI surface ozone analyzer and aerosol pumps off. Adjusted CLAP flow from .59 to .60
2018-10-20 | 293
2018-10-21 | 294 | 1630 | TN | rm hf | * | 5.6 | 100 | 22.3 | 0.6 | 0.675 | 6 |  | Adjusted CLAP flow from .61 to .60
2018-10-22 | 295 | 1616 | TN | rm lf | * | 5.6 | 100 | 23 | 0.6 | 0.953 | 7
2018-10-23 | 296 | 1656 | TN | rm lf | * | 5.6 | 100 | 22.6 | 0.6 | 0.941 | 7 |  | Adjusted CLAP flow from .61 to .60
2018-10-24 | 297 | 1710 | TN | rm lf | * | 5.6 | 100 | 22.5 | 0.6 | 0.939 | 7 |  | Adjusted CLAP flow from .61 to .60
2018-10-25 | 298 | 1711 | TN | rm hf | * | 5.6 | 100 | 22.8 | 0.6 | 0.933 | 7
2018-10-26 | 299 | 1640 | TN | rm hf | * | 5.6 | 100 | 22.5 | 0.6 | 0.922 | 7 |  | Adjusted CLAP flow from .61 to .60
2018-10-27 | 300 | 1536 | TN | clr | * | 5.6 | 100 | 22.5 | 0.6 | 0.914 | 7 |  | Adjusted CLAP flow from .59 to .60
2018-10-28 | 301 | 1526 | TN | rm lf | * | 5.6 | 100 | 22.3 | 0.6 | 0.911 | 7 |  | Adjusted CLAP flow from .58 to .60
2018-10-29 | 302 | 1658 | TN | rm hf | * | 5.6 | 100 | 22.3 | 0.6 | 0.8 | 7 |  | Adjusted CLAP flow from .61 to .60
2018-10-30 | 303 | 1506 | TN | clr | * | 5.6 | 100 | 22.3 | 0.6 | 0.798 | 7
2018-10-31 | 304 | 1445 | TN | rm lf | * | 5.6 | 100 | 22.5 | 0.6 | 0.796 | 7 | * | Neph span check conducted 1410-1440Z. Error = 2.4%. CO2 psi = 560. Adjusted CLAP flow from .61 to .6
2018-11-01 | 305 | 1742 | TN | rm lf | * | 5.6 | 100 | 22.5 | 0.6 | 0.795 | 7 |  | pump was off upon arrival at 1530Z. Working with PI's to fix. -*Pump back on at 2202Z. Carbon vane failure. Values listed are from 2215Z, after the pump was returned to service.
2018-11-02 | 306 | 1645 | TN | rm lf | * | 5.6 | 100 | 22.4 | 0.6 | 0.793 | 7 |  | Adjusted CLAP flow from .59 to .60
2018-11-03 | 307 | 1813 | TN | rm hf | * | 5.6 | 100 | 22.3 | 0.6 | 0.791 |  |  | No spot # indicated. Internet and phone down at TAWO - could it have s/t to do with that? Adjusted CLAP flow from .58 to .60
2018-11-04 | 308 | 1508 | TN | clr | * | 5.6 | 100 | 22.3 | 0.6 | 0.788 |  |  | No spot # indicated. Adjusted CLAP flow from .61 to .60. CLAP flow slow to adjust.
2018-11-05 | 309 | 1247 | TN | rm lf | * | 5.6 | 100 | 22.3 | 0.6 | 0.786 |  |  | No spot # indicated.
2018-11-06 | 310 | 1240 | TN | rm lf | * | 5.6 | 100 | 22 | 0.6 | 0.783 |  |  | No spot # indicated. Adjusted CLAP flow from .62 to .60
2018-11-07 | 311 | 1255 | TN | rm lf | * | 5.6 | 100 | 22.4 | 0.6 | 0.781 |  |  | No spot # indicated.
2018-11-08 | 312 | 1547 | TN | clr | * | 5.6 | 100 | 22.3 | 0.6 | 0.779 | 7
2018-11-09 | 313 | 1325 | TN | clr | * | 5.5 | 100 | 22.3 | 0.6 | 0.778 | 7
2018-11-10 | 314 | 1244 | TN | clr | * | 5.6 | 100 | 22.3 | 0.6 | 0.778 | 7
2018-11-11 | 315 | 1616 | TN | rm hf | * | 5.5 | 100 | 22.2 | 0.6 | 0.776 | 7
2018-11-12 | 316 | 1233 | TN | rm hf | * | 5.6 | 100 | 22.2 | 0.6 | 0.775 | 7
2018-11-13 | 317 | 1320 | TN | rm hf | * | 5.5 | 0 | 22.2 | 0.6 | 0.772 | 7
2018-11-14 | 318 | 1336 | TN | rm hf | * | 5.5 | 100 | 22 | 0.6 | 0.771 | 7 |  | Adjusted CLAP flow rate from .59 to .60
2018-11-15 | 319 | 1355 | TN | rm hf | * | 5.5 | 100 | 22 | 0.6 | 0.77 | 7 | * | Neph span check conducted from 1409 - 1441 . CO2 tank psi was 700. error was 2.7
2018-11-16 | 320 | 1914 | TN | rm hf | * | 5.6 | 100 | 22 | 0.6 | 0.99986 | 1 |  | Replaced CLAP filter. Adjusted CLAP flow rate from .62 to .60.
2018-11-17 | 321 | 1228 | TN | rm hf | * | 5.5 | 100 | 22 | 0.6 | 1 | 1 |  | Adjusted CLAP flow rate from .59 to .60
2018-11-18 | 322 | 1630 | TN | rm lf | * | 5.6 | 0 | 22 | 0.6 | 0.997 | 1
2018-11-19 | 323 | 1234 | TN | rm lf | * | 5.5 | 100 | 22.2 | 0.6 | 0.996 | 1 |  | Adjusted CLAP flow rate from .62 to .60 lpm.
2018-11-20 | 324 | 1258 | TN | clr | * | 5.5 | 0 | 22.2 | 0.6 | 0.993 | 1
2018-11-21 | 325 | 1324 | TN | clr | * | 5.5 | 0 | 22.2 | 0.6 | 0.979 | 1 |  | Adjusted CLAP flow rate from .59 to .60
2018-11-22 | 326 | 1520 | TN | clr | * | 5.5 | 100 | 22.3 | 0.6 | 0.971 | 1
2018-11-23 | 327 | 1310 | TN | clr | * | 5.5 | 100 | 22.3 | 0.6 | 0.966 | 1 |  | Adjusted CLAP flow rate from .61 to .60. CLAP screen was in "Boxcar" mode; tech hid screen, then brought back up, returning CLAP to normal screen on CPD3.
2018-11-24 | 328 | 1250 | TN | clr | * | 5.5 | 100 | 22.2 | 0.6 | 0.958 | 1
2018-11-25 | 329 | 1803 | TN | rm lf | * | 5.5 | 100 | 22.3 | 0.6 | 0.949 | 1
2018-11-26 | 330 | 1313 | TN | clr | * | 5.5 | 100 | 22.2 | 0.6 | 0.943 | 1 |  | Adjusted CLAP flow rate from .59 to .60
2018-11-27 | 331 | 1530 | TN | rm lf | * | 5.5 | 100 | 22.3 | 0.6 | 0.934 | 1 |  | Adjusted CLAP flow rate from .59 to .60
2018-11-28 | 332 | 1532 | TN | rm lf | * | 5.5 | 100 | 22.3 | 0.6 | 0.928 | 1 | * | Adjusted CLAP flow rate from .59 to .60. Neph span check conducted from 1807-1839. CO2 tank pressure was 550 psi. Delivery pressure was 25 psi at 4.3 lpm. Error was 1.1.
2018-11-29 | 333 | 1555 | TN | rm lf | * | 5.5 | 100 | 22.2 | 0.6 | 0.918 | 1 |  | Adjusted CLAP flow rate from .59 to .60
2018-11-30 | 334 | 1251 | TN | rm lf | * | 5.5 | 100 | 21 | 0.6 | 0.91 | 1
2018-12-01 | 335 | 1330 | TN | rm hf | * | 5.5 | 100 | 22 | 0.6 | 0.907 | 1
2018-12-02 | 336 | 1601 | TN | rm lf | * | 5.5 | 100 | 21.5 | 0.6 | 0.9 | 1
2018-12-03 | 337 | 1338 | TN | rm lf | * | 5.5 | 100 | 21 | 0.6 | 0.891 | 1
2018-12-04 | 338 | 1519 | TN | rm hf | * | 5.5 | 100 | 21.5 | 0.6 | 0.88 | 1
2018-12-05 | 339 | 1322 | TN | rm lf | * | 5.5 | 100 | 22 | 0.6 | 0.874 | 1
2018-12-06 | 340 | 1631 | TN | rm hf | * | 5.5 | 100 | 22 | 0.6 | 0.87 | 1 |  | CLAP flow rate adjusted from .61 to .60
2018-12-07 | 341 | 1308 | TN | rm hf | * | 5.5 | 100 | 22 | 0.6 | 0.868 | 1 |  | CLAP flow rate adjusted from .61 to .60
2018-12-08 | 342 | 1258 | TN | rm hf | * | 5.5 | 100 | 21.5 | 0.6 | 0.867 | 1
2018-12-09 | 343 | 1420 | TN | rm hf | * | 5.5 | 100 | 21.5 | 0.6 | 0.892 | 1
2018-12-10 | 344 | 1314 | TN | rm lf | * | 5.5 | 100 | 22 | 0.6 | 0.889 | 1
2018-12-11 | 345 | 1329 | TN | rm hf | * | 5.5 | 100 | 22 | 0.6 | 0.858 | 1 | * | CLAP flowr rate adjusted from .61 to .60. Neph span check conducted from 1406 - 1437; CO2 tank pressure 580 psi; flow rate of 4.7 lpm at 25 psi delivery; error of 7.9. Span check to be repeated. Second span check: 1445 - 1516 Z ; CO2 tank pressure 590 psi; Flow rate of 4.5 at 25psi delivery; error = 2.7.
2018-12-12 | 346 | 1324 | TN | rm lf | * | 5.5 | 100 | 22 | 0.6 | 0.857 | 1 |  | Adjusted CLAP flow rate from .57 to .60
2018-12-13 | 347 | 1328 | TN | rm hf | * | 5.5 | 100 | 22 | 0.6 | 0.856 | 1
2018-12-14 | 348 | 1315 | TN | rm lf | * | 5.45 | 100 | 22 | 0.6 | 0.855 | 1
2018-12-15 | 349 | 1714 | TN | rm lf | * | 5.5 | 100 | 22 | 0.6 | 0.853 | 1 |  | CLAP flow rate adjusted from .61 to .60
2018-12-16 | 350 | 1443 | TN | rm lf | * | 5.4 | 100 | 22 | 0.6 | 0.85 | 1
2018-12-17 | 351 | 1252 | TN | rm lf | * | 5.4 | 100 | 21.5 | 0.6 | 0.848 | 1
2018-12-18 | 352 | 1314 | TN | rm lf | * | 5.4 | 100 | 22 | 0.6 | 0.846 | 1
2018-12-19 | 353 | 1302 | TN | rm lf | * | 5.6 | 0 | 22 | 0.6 | 0.843 | 1
2018-12-20 | 354 | 1258 | TN | rm lf | * | 5.7 | 0 | 22 | 0.6 | 0.842 | 1
2018-12-21 | 355 | 1316 | TN | rm lf | * | 5.7 | 100 | 21.5 | 0.6 | 0.838 | 1
2018-12-22 | 356
2018-12-23 | 357
2018-12-24 | 358 | 1301 | TN | rm lf | * | 5.6 | 100 | 21.5 | 0.6 | 0.829 | 1 |  | CLAP flow rate adjusted from .59 to .60
2018-12-25 | 359 | 1354 | TN | rm hf | * | 5.6 | 100 | 21.5 | 0.6 | 0.829 | 1 | * | CLAP flow rate adjusted from .59 to .60. Neph span check conducted from 1426 -1458 . Tank psi: 710; flow rate: 4.6 lpm @ 25 psi. error: 1.7%
2018-12-26 | 360 | 1322 | TN | rm hf | * | 5.6 | 100 | 21.5 | 0.6 | 0.829 | 1
2018-12-27 | 361 | 1305 | TN | rm hf | * | 5.6 | 100 | 21.5 | 0.6 | 0.827 | 1 |  | CLAP flow rate adjusted from .59 to .60
2018-12-28 | 362
2018-12-29 | 363 | 1327 | TN | rm lf | * | 5.6 | 100 | 21.5 | 0.6 | 0.825 | 1 |  | CLAP flow rate adjusted from .61 to .60
2018-12-30 | 364
2018-12-31 | 365 | 1254 | TN | rm lf | * | 5.6 | 100 | 22 | 0.6 | 0.82 | 1 |  | CLAP flow rate adjusted from .61 to .60

## CU GC
Date | Julian Day | Time (UTC) | Observer's Initials | H2 Gen. Act Pressure (PSI) | H2 Generator Water Level | H2 Gen. Flow (ml/min) | Methane Std 1 Pressure | Methane Std 2 Pressure | Inlet Line Flow Rate (L/min) | 4 Pumps On? | NMHC Std 1 Pressure | NMHC Std 2 Pressure | Trap Housing Purge Flow (ml/min) | Graphs Look Good? | Daily H2 Gen Desiccant Check | Daily H2 Line Filter Desiccant Check | Daily Oxygen indicator (green) | H2 Cylinder Pressure | Comments (Inlet Cleaning, Inlet Filter Changes, Water Refill, etc.)
2018-01-01 | 1 | 1719 | HA | 61 | * | 42 | 1500 | 950 | 1.75 | Y | 425 | 700 | 100 |  | * | * |  | 1450
2018-01-02 | 2 | 0 | 0 | [15 empty] | Could not visit TAWO due to weather
2018-01-03 | 3 | 0 | 0 | [15 empty] | Could not visit TAWO due to weather
2018-01-04 | 4 | 0 | 0 | [15 empty] | Could not visit TAWO due to weather
2018-01-05 | 5 | 1454 | HA | 61 | * | 43 | 1500 | 900 | 1.75 | Y | 425 | 700 | 100 |  | * | * |  | 1450
2018-01-06 | 6 | 0 | 0
2018-01-07 | 7 | 0 | 0
2018-01-08 | 8 | 1423 | HA | 61 | * | 42 | 1500 | 900 | 1.75 | Y | 425 | 700 | 100 |  | * | * |  | 1450
2018-01-09 | 9 | 0 | 0
2018-01-10 | 10 | 1523 | HA | 61 | * | 40 | 1500 | 900 | 1.75 | Y | 425 | 700 | 100 |  | * | *
2018-01-11 | 11 | 0 | 0
2018-01-12 | 12 | 1440 | HA | 61 | added 1 liter water | 42 | 1500 | 900 | 1.75 | y | 425 | 700 | 100 |  | * | * |  | 1425
2018-01-13 | 13 | 0 | 0
2018-01-14 | 14 | 0 | 0
2018-01-15 | 15 | 1413 | HA | 61 | * | 42 | 1500 | 900 | 1.75 | Y | 425 | 700 | 100 |  | * | * |  | 1375
2018-01-16 | 16 | 0 | 0
2018-01-17 | 17 | 0 | 0
2018-01-18 | 18 | 1407 | HA | 61 | * | 43 | 1500 | 900 | 1.75 | Y | 425 | 700 | 100 |  | * | * |  | 1400
2018-01-19 | 19 | 0 | 0
2018-01-20 | 20 | 1323 | HA | 61 | * | 38 | 1500 | 900 | 1.75 | Y | 425 | 700 | 100 |  | * | * |  | 1400 | Unexpected power outage from 18:21-18:22Z and 19:00-19:04Z
2018-01-21 | 21 | 0 | 0
2018-01-22 | 22 | 1758 | HA | 61 | * | 37 | 1500 | 850 | 1.75 | Y | 425 | 700 | 90 |  | * | * |  | 1400
2018-01-23 | 23 | 0 | 0
2018-01-24 | 24 | 1409 | HA | 61 | * | 42 | 1500 | 850 | 1.75 | Y | 425 | 700 | 100 |  | * | * |  | 1375
2018-01-25 | 25 | 0 | 0
2018-01-26 | 26 | 0 | 0
2018-01-27 | 27 | 0 | 0
2018-01-28 | 28 | 0 | 0
2018-01-29 | 29 | 1432 | HA | 61 | added 1/2 liter DI water | 42 | 1500 | 850 | 1.75 | Y | 425 | 700 | 100 |  | * | * |  | 1400
2018-01-30 | 30 | 0 | 0
2018-01-31 | 31 | 1434 | HA | 61 | * | 42 | 1500 | 850 | 1.75 | Y | 425 | 700 | 100 |  | * | * |  | 1375
2018-02-01 | 32 | 0 | 0
2018-02-02 | 33 | 0 | 0
2018-02-03 | 34 | 1411 | HA | 61 | * | 42 | 1500 | 850 | 1.75 | Y | 425 | 700 | 100 |  | * | * |  | 1350
2018-02-04 | 35 | 0 | 0
2018-02-05 | 36 | 1424 | HA | 61 | * | 42 | 1500 | 850 | 1.75 | Y | 425 | 700 | 100 |  | * | *
2018-02-06 | 37 | 1428 | HA
2018-02-07 | 38 | 1612 | HA | 61 | * | 42 | 1500 | 825 | 1.75 | Y | 400 | 675 | 100 |  | * | * |  | 1350
2018-02-08 | 39 | 0 | 0
2018-02-09 | 40 | 1314 | HA | 61 | * | 42 | 1500 | 800 | 1.75 | Y | 400 | 675 | 100 |  | * | * |  | 1350
2018-02-10 | 41 | 0 | 0
2018-02-11 | 42 | 0 | 0
2018-02-12 | 43 | 1458 | HA, SWD | 61 | added 1 liter water | 42 | 1500 | 800 | 1.75 | Y | 400 | 680 | 95 |  | * | * |  | 1320
2018-02-13 | 44 | 0 | 0
2018-02-14 | 45 | 1418 | SWD | 61 | * | 41 | 1500 | 800 | 1.75 | Y | 380 | 670 | 95 |  | * | * | * | 1320
2018-02-15 | 46 | 0 | 0
2018-02-16 | 47 | 0 | 0
2018-02-17 | 48 | 1353 | SWD | 61 | * | 42 | 1500 | 800 | 1.75 | Y | 380 | 670 | 100 |  | * | * | * | 1320 | Added 1.6 L water at time of check. Central front panel button on H2 Generator is failing.
2018-02-18 | 49 | 0 | 0
2018-02-19 | 50 | 0 | 0
2018-02-20 | 51 | 0 | 0
2018-02-21 | 52 | 1428 | SWD | 61 | * | 41 | 1500 | 800 | 1.75 | Y | 390 | 670 | 95 |  | * | * | * | 1320
2018-02-22 | 53 | 0 | 0
2018-02-23 | 54 | 0 | 0
2018-02-24 | 55 | 0 | 0
2018-02-25 | 56 | 1503 | SWD | 61 | * | 59 | 1500 | 800 | 1.75 | Y | 390 | 670 | 95 |  | * | * | * | 1320
2018-02-26 | 57 | 1539 | SWD | 61 | * | 42 | 1500 | 800 | 1.75 | Y | 390 | 670 | 95 |  | * | * | * | 1320
2018-02-27 | 58 | 1851 | SWD | 61 | * | 42 | 1500 | 780 | 1.75 | Y | 390 | 670 | 95 |  | * | * | * | 1320
2018-02-28 | 59 | 0 | 0
2018-03-01 | 60 | 0 | 0
2018-03-02 | 61 | 1657 | SWD | 61 | * | 42 | 1500 | 780 | 1.75 | Y | 390 | 670 | 100 |  | * | * | * | 1320
2018-03-03 | 62 | 0 | 0
2018-03-04 | 63 | 0 | 0
2018-03-05 | 64 | 1439 | SWD | 61 | * | 42 | 1520 | 750 | 1.75 | Y | 380 | 670 | 100 |  | * | * | * | 1320 | Added DI water to H2 generator at time of check.
2018-03-06 | 65 | 0 | 0
2018-03-07 | 66 | 1517 | SWD | 61 | * | 40 | 1520 | 750 | 1.75 | Y | 380 | 670 | 95 |  | * | * | * | 1320 | H2 Generator display unable to show full-precision 1 mL/min flow rate due to issue with button; instead the filled-square indicator is used (1 square = 20 mL/min).
2018-03-08 | 67 | 0 | 0
2018-03-09 | 68 | 0 | 0
2018-03-10 | 69 | 1243 | SWD | 61 | * | 40 | 1520 | 750 | 1.75 | Y | 380 | 660 | 95 |  | * | * | * | 1320
2018-03-11 | 70 | 0 | 0
2018-03-12 | 71 | 1454 | SWD | 61 | * | 40 | 1520 | 740 | 1.75 | Y | 380 | 660 | 95 |  | * | * | * | 1320
2018-03-13 | 72 | 0 | 0
2018-03-14 | 73 | 1529 | SWD | 61 | * | 40 | 1500 | 740 | 1.75 | Y | 380 | 660 | 95 |  | * | * | * | 1320
2018-03-15 | 74 | 0 | 0
2018-03-16 | 75 | 1244 | SWD | 61 | * | 40 | 1510 | 720 | 1.75 | Y | 380 | 660 | 95 |  | * | * | * | 1320
2018-03-17 | 76 | 0 | 0
2018-03-18 | 77 | 0 | 0
2018-03-19 | 78 | 1459 | SWD | 61 | * | 40 | 1510 | 720 | 1.75 | Y | 380 | 660 | 100 |  | * | * | * | 1320
2018-03-20 | 79 | 0 | 0
2018-03-21 | 80 | 1320 | SWD | 61 | * | 40 | 1520 | 720 | 1.75 | Y | 380 | 660 | 95 |  | * | * | * | 1320
2018-03-22 | 81 | 0 | 0
2018-03-23 | 82 | 1300 | SWD/RN | 61 | * | 40 | 1530 | 690 | 1.75 | Y | 380 | 650 | 95 |  | * | * | * | 1350 | Added DI water to H2 generator at 1322z.
2018-03-24 | 83 | 0 | 0
2018-03-25 | 84 | 1227 | SWD | 61 | * | 40 | 1530 | 690 | 1.75 | Y | 380 | 650 | 95 |  | * | * | * | 1350
2018-03-26 | 85 | 0 | 0
2018-03-27 | 86 | 0 | 0
2018-03-28 | 87 | 1503 | NB | 61 | * | 40 | 1520 | 690 | 1.75 | Y | 380 | 650 | 100 |  | * | * | * | 1340
2018-03-29 | 88 | 0 | 0
2018-03-30 | 89 | 1722 | NB | 61 | * | 40 | 1510 | 690 | 1.75 | Y | 380 | 660 | 100 |  | * | * | * | 1300
2018-03-31 | 90 | 0 | 0
2018-04-01 | 91 | 1157 | NB | 61 | * | 40 | 1510 | 670 | 1.75 | Y | 380 | 660 | 100 |  | * | * | * | 1300
2018-04-02 | 92 | 1306 | NB | 61 | * | 40 | 1510 | 680 | 1.75 | Y | 380 | 660 | 100 |  | * | * | * | 1310
2018-04-03 | 93 | 0 | 0
2018-04-04 | 94 | 0 | 0
2018-04-05 | 95 | 1203 | NB | 61 | * | 40 | 1500 | 660 | 1.75 | Y | 370 | 650 | 100 |  | * | * | * | 1260
2018-04-06 | 96 | 0 | 0
2018-04-07 | 97 | 1122 | NB | 61 | * | 40 | 1500 | 650 | 1.75 | Y | 360 | 650 | 100 |  | * | * | * | 1260
2018-04-08 | 98 | 0 | 0
2018-04-09 | 99 | 1135 | NB | 61 | * | 40 | 1510 | 650 | 1.75 | Y | 370 | 650 | 100 |  | * | * | * | 1310
2018-04-10 | 100 | 0 | 0
2018-04-11 | 101 | 0 | 0
2018-04-12 | 102 | 1130 | NB | 61 | * | 40 | 1510 | 650 | 1.75 | Y | 370 | 650 | 100 |  | * | * | * | 1300
2018-04-13 | 103 | 0 | 0
2018-04-14 | 104 | 1314 | NB | 61 | * | 40 | 1520 | 640 | 1.75 | Y | 370 | 650 | 100 |  | * | * | * | 1350 | Added 1L DI water to H2 generator at 1322z.
2018-04-15 | 105 | 0 | 0
2018-04-16 | 106 | 1127 | NB | 61 | * | 40 | 1510 | 640 | 1.75 | Y | 370 | 650 | 100 |  | * | * | * | 1350
2018-04-17 | 107 | 0 | 0
2018-04-18 | 108 | 1145 | NB | 61 | * | 40 | 1510 | 640 | 1.75 | Y | 360 | 650 | 100 |  | * | * | * | 1350
2018-04-19 | 109 | 0 | 0
2018-04-20 | 110 | 1141 | NB | 61 | * | 40 | 1510 | 630 | 1.75 | Y | 370 | 650 | 100 |  | * | * | * | 1320
2018-04-21 | 111 | 0 | 0
2018-04-22 | 112 | 1207 | NB | 61 | * | 40 | 1510 | 620 | 1.75 | Y | 360 | 640 | 100 |  | * | * | * | 1310
2018-04-23 | 113 | 1144 | NB | 61 | * | 40 | 1510 | 620 | 1.75 | Y | 360 | 650 | 100 |  | * | * | * | 1310
2018-04-24 | 114 | 0 | 0
2018-04-25 | 115 | 0 | 0
2018-04-26 | 116 | 1130 | NB | 61 | * | 40 | 1510 | 610 | 1.75 | Y | 360 | 640 | 98 |  | * | * | * | 1280
2018-04-27 | 117 | 1327 | NB/MB | 61 | * | 40 | 1510 | 600 | 1.75 | Y | 370 | 630 | 100 |  | * | * | * | 1290
2018-04-28 | 118 | 0 | 0
2018-04-29 | 119 | 0 | MB
2018-04-30 | 120 | 1228 | NB/MB | 61 | * | 40 | 1510 | 590 | 1.75 | Y | 360 | 630 | 85 |  | * | * | * | 1260 | Trap housing purge flow increased to 98 ml/min when checked again ~1 hour later.
2018-05-01 | 121 | 1224 | MB | 61 | * | 40 | 1510 | 590 | 1.75 | Y | 360 | 630 | 98 |  | * | * | * | 1280
2018-05-02 | 122 | 0 | 0
2018-05-03 | 123 | 0 | 0
2018-05-04 | 124 | 1227 | MB | 61 | * | 40 | 1510 | 580 | 1.75 | Y | 360 | 630 | 98 |  | * | * | * | 1270 | Added 1L of DI water to H2 generator at 1248z.
2018-05-05 | 125 | 1217 | MB | 61 | * | 40 | 1510 | 570 | 1.75 | Y | 360 | 630 | 98 |  | * | * | * | 1280
2018-05-06 | 126 | 0 | 0
2018-05-07 | 127 | 1224 | MB | 61 | * | 40 | 1520 | 570 | 1.75 | Y | 360 | 630 | 98 |  | * | * | * | 1250
2018-05-08 | 128 | 1247 | MB | 61 | * | 40 | 1520 | 570 | 1.7 | Y | 360 | 630 | 98 |  | * | * | * | 1260
2018-05-09 | 129 | 0 | 0
2018-05-10 | 130 | 1247 | MB | 61 | * | 40 | 1520 | 570 | 1.7 | Y | 360 | 630 | 98 |  | * | * | * | 1260
2018-05-11 | 131 | 0 | 0
2018-05-12 | 132 | 1721 | MB | 61 | * | 40 | 1520 | 570 | 1.75 | Y | 360 | 630 | 98 |  | * | * | * | 1240
2018-05-13 | 133 | 0 | 0
2018-05-14 | 134 | 1153 | MB | 61 | * | 40 | 1520 | 570 | 1.75 | Y | 360 | 630 | 98 |  | * | * | * | 1260
2018-05-15 | 135 | 1515 | MB | 61 | * | 40 | 1520 | 570 | 1.75 | Y | 360 | 630 | 98 |  | * | * | * | 1700 | Replaced H2 generator desiccant cartridge 1330-1407z. H2 generator acutal pressure 61 psi, flow 20 mL/min (1/5 squares). Flow increased to 40 mL/min (2/5 squares) in the following hour.
2018-05-16 | 136 | 0 | 0
2018-05-17 | 137 | 1205 | MB | 61 | * | 40 | 1520 | 560 | 1.75 | Y | 360 | 630 | 98 |  | * | * | * | 1690
2018-05-18 | 138 | 0 | 0
2018-05-19 | 139 | 1219 | MB | 61 | * | 40 | 1520 | 560 | 1.75 | Y | 360 | 630 | 98 |  | * | * | * | 1680 | Added 1L of DI water to H2 generator at 1226z.
2018-05-20 | 140 | 0 | 0
2018-05-21 | 141 | 1658 | MB | 61 | * | 40 | 1520 | 560 | 1.75 | Y | 360 | 630 | 98 |  | * | * | * | 1670
2018-05-22 | 142 | 1859 | MB | 61 | * | 40 | 1520 | 560 | 1.75 | Y | 360 | 630 | 98 |  | * | * | * | 1650
2018-05-23 | 143 | 1243 | MB | 61 | * | 40 | 1520 | 560 | 1.75 | Y | 360 | 630 | 98 |  | * | * | * | 1650
2018-05-24 | 144 | 1226 | MB | 61 | * | 40 | 1520 | 540 | 1.75 | Y | 360 | 630 | 98 |  | * | * | * | 1650
2018-05-25 | 145 | 0 | 0
2018-05-26 | 146 | 0 | 0
2018-05-27 | 147 | 0 | 0
2018-05-28 | 148 | 1220 | MB | 61 | * | 40 | 1520 | 540 | 1.7 | Y | 360 | 630 | 98 |  | * | * | * | 1630
2018-05-29 | 149 | 1347 | MB | 61 | * | 40 | 1520 | 540 | 1.75 | Y | 360 | 630 | 90 |  | * | * | * | 1610
2018-05-30 | 150 | 1201 | MB | 61 | * | 40 | 1520 | 540 | 1.75 | Y | 360 | 630 | 98 |  | * | * | * | 1610
2018-05-31 | 151 | 0 | 0
2018-06-01 | 152 | 1129 | MB | 61 | * | 40 | 1520 | 530 | 1.75 | Y | 360 | 620 | 98 |  | * | * | * | 1600
2018-06-02 | 153 | 0 | 0
2018-06-03 | 154 | 1203 | MB | 61 | * | 40 | 1520 | 530 | 1.75 | Y | 360 | 620 | 98 |  | * | * | * | 1600
2018-06-04 | 155 | 0 | 0
2018-06-05 | 156 | 1302 | MB/MM | 61 | * | 40 | 1520 | 530 | 1.75 | Y | 360 | 620 | 100 |  | * | * | * | 1600
2018-06-06 | 157 | 1341 | MB/MM | 61 | * | 60 | 1520 | 500 | 1.75 | Y | 350 | 610 | 100 |  | * | * | * | 1600
2018-06-07 | 158 | 1209 | MB/MM | 61 | * | 40 | 1520 | 500 | 1.75 | Y | 350 | 610 | 100 |  | * | * | * | 1600 | Added 1L of DI water to H2 generator at 1213z.
2018-06-08 | 159 | 1211 | MM | 61 | * | 40 | 1520 | 500 | 1.75 | Y | 350 | 610 | 100 |  | * | * | * | 1600
2018-06-09 | 160 | 1647 | MM | 61 | * | 40 | 1520 | 500 | 1.75 | Y | 350 | 610 | 100 |  | * | * | * | 1600
2018-06-10 | 161 | 0 | 0
2018-06-11 | 162 | 1659 | MB/MM | 61 | * | 40 | 1520 | 500 | 1.75 | Y | 350 | 610 | 100 |  | * | * | * | 1600 | Removed H2 Generator button panel at approximately 1810z
2018-06-12 | 163 | 1806 | MM | 62 | * | 40 | 1520 | 500 | 1.75 | Y | 350 | 610 | 100 |  | * | * | * | 1590
2018-06-13 | 164 | 0 | 0
2018-06-14 | 165 | 0 | 0
2018-06-15 | 166 | 1828 | MM | 63 | * | 40 | 1510 | 500 | 1.75 | Y | 350 | 610 | 105 |  | * | * | * | 1550 | Replaced the H2 Generator button panel at approximately 1730z. The GC was remotely powered down during a planned power outage and resumed operation at approximately 1930z
2018-06-16 | 167 | 1741 | MM | 63 | * | 40 | 1510 | 500 | 1.75 | Y | 350 | 600 | 105 |  | * | * | * | 1550
2018-06-17 | 168 | 0 | 0
2018-06-18 | 169 | 0 | 0 | [15 empty] | The CU GC ran on the UPS during two brief planned power outages from 0115z-0130z and 0645z-0650z
2018-06-19 | 170 | 1326 | MM | 62 | * | 40 | 1510 | 500 | 1.75 | Y | 350 | 600 | 100 |  | * | * | * | 1550
2018-06-20 | 171 | 2023 | MM | 62 | * | 40 | 1510 | 470 | 1.75 | Y | 350 | 600 | 100 |  | * | * | * | 1525
2018-06-21 | 172 | 0 | 0
2018-06-22 | 173 | 1358 | MM | 62 | * | 40 | 1510 | 450 | 1.75 | Y | 350 | 600 | 100 |  | * | * | * | 1510
2018-06-23 | 174 | 0 | 0
2018-06-24 | 175 | 0 | 0
2018-06-25 | 176 | 1802 | MM | 62 | * | 40 | 1510 | 450 | 1.75 | Y | 350 | 600 | 100 |  | * | * | * | 1510
2018-06-26 | 177 | 1855 | MM | 62 | * | 40 | 1510 | 450 | 1.75 | Y | 350 | 600 | 100 |  | * | * | * | 1500 | Instrument offline during planned internet outages associated with Big House lift: 6/26 from 1515 -2136 and from 2345 to 0230 UTC on 6/27
2018-06-27 | 178 | 1649 | MM | 62 | * | 40 | 1510 | 450 | 1.75 | Y | 350 | 600 | 100 |  | * | * | * | 1500
2018-06-28 | 179 | 1558 | MM | 62 | * | 40 | 1510 | 450 | 1.75 | Y | 350 | 600 | 100 |  | * | * | * | 1500 | Added 1.5L of DI water to H2 generator at 1540
2018-06-29 | 180 | 0 | 0
2018-06-30 | 181 | 2047 | MM | 62 | * | 40 | 1510 | 450 | 1.75 | Y | 350 | 600 | 100 |  | * | * | * | 1500
2018-07-01 | 182 | 0 | 0
2018-07-02 | 183 | 0 | 0
2018-07-03 | 184 | 1603 | MM | 62 | * | 40 | 1510 | 450 | 1.75 | Y | 350 | 600 | 100 |  | * | * | * | 1500
2018-07-04 | 185 | 0 | 0
2018-07-05 | 186 | 1424 | MM | 62 | * | 40 | 1510 | 450 | 1.75 | Y | 350 | 600 | 100 |  | * | * | * | 1450
2018-07-06 | 187 | 0 | 0
2018-07-07 | 188 | 0 | 0
2018-07-08 | 189 | 0 | 0
2018-07-09 | 190 | 0 | 0
2018-07-10 | 191 | 1223 | MM | 62 | * | 40 | 1510 | 450 | 1.75 | Y | 350 | 600 | 100 |  | * | * | * | 1450
2018-07-11 | 192 | 1811 | MM | 62 | * | 40 | 1510 | 450 | 1.75 | Y | 350 | 600 | 100 |  | * | * | * | 1450
2018-07-12 | 193 | 0 | 0
2018-07-13 | 194 | 0 | 0
2018-07-14 | 195 | 0 | 0
2018-07-15 | 196 | 0 | 0
2018-07-16 | 197 | 2209 | MM | 62 | * | 20 | 1510 | 400 | 1.75 | Y | 350 | 600 | 100 |  | * | * | * | 1450 | Removed rime from inlet at 2230, flow returned to 2/4 bars on hydrogen generator (40 lpm)
2018-07-17 | 198 | 0 | 0
2018-07-18 | 199 | 1430 | MM | 62 | * | 40 | 1510 | 400 | 1.75 | Y | 350 | 600 | 100 |  | * | * | * | 1400
2018-07-19 | 200 | 0 | 0
2018-07-20 | 201 | 0 | 0
2018-07-21 | 202 | 1317 | MM | 62 | * | 40 | 1510 | 400 | 1.75 | Y | 310 | 600 | 100 |  | * | * | * | 1390
2018-07-22 | 203 | 0 | 0
2018-07-23 | 204 | 0 | 0 | [15 empty] | CUGC ran on UPS system during a planned power outage from 2153-2300z and again from 2307-2315z.
2018-07-24 | 205 | 0 | MM | 62 | * | 40 | 1510 | 400 | 1.75 | Y | 310 | 600 | 90 |  | * | * | * | 1390 | Added 50ml of water at 21:35z to the H2 generator
2018-07-25 | 206 | 2335 | MM | 62 | * | 40 | 1510 | 350 | 1.75 | Y | 310 | 600 | 105 |  | * | * | * | 1390
2018-07-26 | 207 | 0 | 0
2018-07-27 | 208 | 204 | MM | 62 | * | 40 | 1510 | 350 | 1.75 | Y | 310 | 600 | 105 |  | * | * | * | 1350
2018-07-28 | 209 | 0 | 0 | [15 empty] | TAWO planned power outage from 1218-1228z; Used simple green and rubbing alcohol in the TAWO vestibule
2018-07-29 | 210 | 258 | MM | 62 | * | 40 | 1510 | 350 | 1.75 | Y | 310 | 580 | 105 |  | * | * | * | 1350 | Added 1litre of DI water to H2 generator
2018-07-30 | 211 | 1234 | MM | 62 | * | 40 | 1510 | 350 | 1.75 | Y | 310 | 580 | 105 |  | * | * | * | 1350
2018-07-31 | 212 | 0 | 0
2018-08-01 | 213 | 0 | 0
2018-08-02 | 214 | 2147 | MM | 62 | * | 40 | 1510 | 350 | 1.75 | Y | 310 | 580 | 100 |  | * | * | * | 1350 | CUGC in standby mode during power outages
2018-08-03 | 215 | 1810 | MM | 62 | * | 40 | 1510 | 350 | 1.75 | Y | 310 | 575 | 105 |  | * | * | * | 1325
2018-08-04 | 216 | 1430 | MM | [15 empty] | CUGC remotely shutdown during power outages
2018-08-05 | 217 | 0 | 0
2018-08-06 | 218 | 0 | 0
2018-08-07 | 219 | 2037 | MM | 62 | * | 40 | 1510 | 350 | 0 | N | 310 | 575 | 0 |  | * | * | * | 1310 | CUGC remotely shutdown during power outages
2018-08-08 | 220 | 0 | 0
2018-08-09 | 221 | 1524 | MM | 62 | * | 40 | 1510 | 350 | 1.75 | Y | 310 | 575 | 100 |  | * | * | * | 1310
2018-08-10 | 222 | 0 | 0
2018-08-11 | 223 | 0 | 0
2018-08-12 | 224 | 0 | 0
2018-08-13 | 225 | 2200 | MM | 62 | * | 40 | 1510 | 340 | 1.75 | Y | 310 | 575 | 100 |  | * | * | * | 1300
2018-08-14 | 226 | 0 | 0
2018-08-15 | 227 | 0 | 0
2018-08-16 | 228 | 0 | 0
2018-08-17 | 229 | 1614 | MM | 62 | * | 40 | 1510 | 340 | 1.75 | Y | 315 | 575 | 100 |  | * | * | * | 1300
2018-08-18 | 230 | 0 | 0
2018-08-19 | 231 | 2221 | MM | 62 | * | 40 | 1510 | 340 | 1.75 | Y | 315 | 575 | 100 |  | * | * | * | 1300 | Added 1L of DI water to H2 generator at 2220z.
2018-08-20 | 232 | 0 | 0
2018-08-21 | 233 | 1746 | MM | 62 | * | 40 | 1510 | 320 | 2 | Y | 310 | 590 | 90 |  | * | * | * | 1300
2018-08-22 | 234 | 0 | 0
2018-08-23 | 235 | 0 | 0
2018-08-24 | 236 | 0 | 0
2018-08-25 | 237 | 1930 | MM | 62 | * | 40 | 1510 | 250 | 1.75 | Y | 310 | 590 | 100 |  | * | * | * | 1300 | Jacques onsite.
2018-08-26 | 238 | 1747 | MM | [15 empty] | Jacques onsite. Climbed TAWO inlet tower at 1740 to change CUGC inlet filter
2018-08-27 | 239 | 1251 | MM | [15 empty] | Jacques onsite.
2018-08-28 | 240 | 0 | 0 | [15 empty] | Jacques onsite.
2018-08-29 | 241 | 0 | 0
2018-08-30 | 242 | 0 | 0
2018-08-31 | 243 | 1433 | MM | 62 | * | 40 | 1510 | 290 | 1.9 | Y | 310 | 590 | 115 |  | * | * | * | 1700
2018-09-01 | 244 | 0 | 0
2018-09-02 | 245 | 0 | 0
2018-09-03 | 246 | 0 | 0
2018-09-04 | 247 | 1613 | MM | 62 | * | 40 | 1510 | 290 | 2 | Y | 310 | 590 | 115 |  | * | * | * | 1670
2018-09-05 | 248 | 1659 | MM | 62 | * | 40 | 1510 | 290 | 2 | Y | 310 | 560 | 120 |  | * | * | * | 1690 | Added 1L of DI water to H2 generator at 1700z.
2018-09-06 | 249 | 0 | 0
2018-09-07 | 250 | 0 | 0
2018-09-08 | 251 | 0 | 0
2018-09-09 | 252 | 1632 | MM | 62 | * | 40 | 1510 | 290 | 2 | Y | 310 | 560 | 120 |  | * | * | * | 1640
2018-09-10 | 253 | 0 | 0
2018-09-11 | 254 | 0 | 0
2018-09-12 | 255 | 1414 | MM | 62 | * | 40 | 1510 | 290 | 2 | Y | 310 | 560 | 120 |  | * | * | * | 1640
2018-09-13 | 256 | 0 | 0
2018-09-14 | 257 | 0 | 0
2018-09-15 | 258 | 1950 | MM | 62 | * | 40 | 1510 | 290 | 2 | Y | 310 | 560 | 105 |  | * | * | * | 1640
2018-09-16 | 259 | 0 | 0
2018-09-17 | 260 | 0 | 0
2018-09-18 | 261 | 1445 | MM | 62 | * | 40 | 1510 | 275 | 2 | Y | 310 | 560 | 115 |  | * | * | * | 1590
2018-09-19 | 262 | 0 | 0
2018-09-20 | 263 | 0 | 0
2018-09-21 | 264 | 1900 | MM | 62 | * | 40 | 1510 | 275 | 2 | Y | 310 | 560 | 118 |  | * | * | * | 1570
2018-09-22 | 265 | 0 | 0
2018-09-23 | 266 | 0 | 0
2018-09-24 | 267 | 1925 | MM | 62 | * | 40 | 1510 | 250 | 2 | Y | 310 | 580 | 118 |  | * | * | * | 1550
2018-09-25 | 268 | 0 | 0
2018-09-26 | 269 | 0 | 0
2018-09-27 | 270 | 2004 | MM | 62 | * | 40 | 1510 | 245 | 2 | Y | 310 | 575 | 113 |  | * | * | * | 1540 | Added 1L of DI water to H2 generator at 2005z.
2018-09-28 | 271 | 1954 | MM | 62 | * | 40 | 1510 | 245 | 2 | Y | 310 | 560 | 118 |  | * | * | * | 1530
2018-09-29 | 272 | 0 | 0
2018-09-30 | 273 | 0 | 0
2018-10-01 | 274 | 0 | 0
2018-10-02 | 275 | 1937 | MM | 62 | * | 40 | 1510 | 200 | 2 | Y | 300 | 560 | 118 |  | * | * | * | 1510
2018-10-03 | 276 | 0 | 0
2018-10-04 | 277 | 0 | 0
2018-10-05 | 278 | 1747 | MM | 62 | * | 40 | 1510 | 200 | 2 | Y | 300 | 560 | 115 |  | * | * | * | 1500
2018-10-06 | 279 | 0 | 0
2018-10-07 | 280 | 0 | 0
2018-10-08 | 281 | 1843 | MM | 62 | * | 40 | 1510 | 200 | 2 | Y | 295 | 510 | 113 |  | * | * | * | 1500
2018-10-09 | 282 | 0 | 0
2018-10-10 | 283 | 0 | 0 | [15 empty] | Replaced O3 Analyzer inlet filter on TAWO mast from 1631-1645z.
2018-10-11 | 284 | 0 | 0
2018-10-12 | 285 | 1124 | MM | 62 | * | 40 | 1500 | 175 | 2 | Y | 300 | 550 | 118 |  | * | * | * | 1475 | Added 1L of DI water to H2 generator at 1128z.
2018-10-13 | 286 | 0 | 0
2018-10-14 | 287 | 0 | 0
2018-10-15 | 288 | 1800 | MM | 62 | * | 40 | 1510 | 160 | 2 | Y | 300 | 550 | 118 |  | * | * | * | 1490
2018-10-16 | 289 | 1750 | MM/TN
2018-10-17 | 290 | 0 | 0
2018-10-18 | 291 | 1735 | MM/TN | 61 | * | 40 | 1515 | 150 | 2 | Y | 295 | 550 | 110 |  | * | * | * | 1660 | Changed the H2 gen desiccant cartridge 1617-1714z. (1632 powered H2 generator back on, 1643 began self-test)
2018-10-19 | 292 | 1848 | MM/TN
2018-10-20 | 293 | 0 | 0
2018-10-21 | 294 | 1630 | TN | 61 | * | 40 | 1510 | 150 | 2 | Y | 295 | 545 | 100 |  | * | * | * | 1620 | Added 1L of DI water to H2 generator at 1655z
2018-10-22 | 295 | 1616 | TN | 61 | * | 40 | 1505 | 150 | 2 | Y | 295 | 545 | 110 |  | * | * | * | 1620
2018-10-23 | 296 | 1656 | TN | 61 | * | 40 | 1500 | 150 | 2 | y | 300 | 545 | 115 |  | * | * | * | 1650
2018-10-24 | 297 | 1710 | TN | 61 | * | 40 | 1500 | 150 | 2 | Y | 295 | 550 | 110 |  | * | * | * | 1645
2018-10-25 | 298 | 1711 | TN | 61 | * | 40 | 1505 | 145 | 2 | y | 295 | 545 | 115 |  | * | * | * | 1625
2018-10-26 | 299 | 1640 | TN | 61 | * | 40 | 1500 | 150 | 2 | Y | 295 | 545 | 115 |  | * | * | * | 1620
2018-10-27 | 300 | 1536 | TN | 61 | * | 40 | 1500 | 145 | 2 | Y | 295 | 545 | 110 |  | * | * | * | 1600
2018-10-28 | 301 | 1526 | TN | 61 | * | 40 | 1505 | 145 | 2 | y | 295 | 545 | 110 |  | * | * | * | 1620
2018-10-29 | 302 | 1658 | TN | 61 | * | 40 | 1505 | 145 | 2 | Y | 295 | 545 | 110 |  | * | * | * | 1620
2018-10-30 | 303 | 1506 | TN | 61 | * | 40 | 1510 | 140 | 2 | Y | 295 | 540 | 110 |  | * | * | * | 1610
2018-10-31 | 304 | 1445 | TN | 61 | * | 40 | 1510 | 140 | 2 | Y | 295 | 545 | 110 |  | * | * | * | 1600
2018-11-01 | 305 | 1742 | TN | 61 | * | 40 | 1500 | 140 | 2 | Y | 295 | 540 | 110 |  | * | * | * | 1590
2018-11-02 | 306 | 1645 | TN | 61 | * | 40 | 1505 | 125 | 2 | Y | 295 | 540 | 110 |  | * | * | * | 1600
2018-11-03 | 307 | 1813 | TN | 61 | * | 40 | 1505 | 120 | 2 | Y | 295 | 545 | 110 |  | * | * | * | 1595 | H2 gen briefly dropped to 20 ml/min flow.
2018-11-04 | 308 | 1508 | TN | 61 | * | 40 | 1505 | 120 | 2 | Y | 295 | 540 | 110 |  | * | * | * | 1580
2018-11-05 | 309 | 1247 | TN | 61 | * | 40 | 1505 | 115 | 2 | Y | 295 | 540 | 110 |  | * | * | * | 1580
2018-11-06 | 310 | 1240 | TN | 61 | * | 40 | 1500 | 115 | 2 | Y | 295 | 540 | 110 |  | * | * | * | 1580
2018-11-07 | 311 | 1255 | TN | 61 | * | 40 | 1500 | 110 | 2 | Y | 295 | 540 | 110 |  | * | * | * | 1575
2018-11-08 | 312 | 1547 | TN | 61 | * | 40 | 1500 | 110 | 2 | y | 295 | 540 | 110 |  | * | * | * | 1580
2018-11-09 | 313 | 1325 | TN | 61 | * | 40 | 1500 | 110 | 2 | y | 290 | 530 | 110 |  | * | * | * | 1590
2018-11-10 | 314 | 1244 | TN | 61 | * | 40 | 1500 | 110 | 2 | Y | 295 | 540 | 110 |  | * | * | * | 1590
2018-11-11 | 315 | 1616 | TN | 61 | * | 40 | 1500 | 110 | 2 | Y | 290 | 545 | 110 |  | * | * | * | 1575
2018-11-12 | 316 | 1233 | TN | 61 | * | 40 | 1510 | 110 | 2 | Y | 295 | 545 | 110 |  | * | * | * | 1575
2018-11-13 | 317 | 1320 | TN | 61 | * | 40 | 1505 | 105 | 2 | Y | 295 | 540 | 110 |  | * | * | * | 1575 | Added 1L of DI water to H2 generator at 1351z
2018-11-14 | 318 | 1336 | TN | 61 | * | 40 | 1510 | 90 | 2 | Y | 290 | 540 | 110 |  | * | * | * | 1575 | H2 gen flow briefly rose to 60%
2018-11-15 | 319 | 1355 | TN | 61 | * | 40 | 1515 | 90 | 2 | Y | 295 | 540 | 110 |  | * | * | * | 1570 | Switched CH4 from Std 2 to Std1 tank at 1340 Z
2018-11-16 | 320 | 1914 | TN | 61 | * | 40 | 1515 | 90 | 2 | Y | 290 | 535 | 110 |  | * | * | * | 1560 | After PI consultation, re-opened valve to Std2. Installed Teamviewer 14 on GC computer. Changed Std tank view GC computer with guidance from PI.
2018-11-17 | 321 | 1228 | TN | 61 | * | 40 | 1510 | 90 | 2 | Y | 290 | 530 | 110 |  | * | * | * | 1550
2018-11-18 | 322 | 1630 | TN | 61 | * | 40 | 1510 | 90 | 2 | Y | 290 | 540 | 107 |  | * | * | * | 1550
2018-11-19 | 323 | 1234 | TN | 61 | * | 40 | 1510 | 90 | 2 | Y | 295 | 540 | 107 |  | * | * | * | 1550
2018-11-20 | 324 | 1258 | TN | 61 | * | 40 | 1500 | 90 | 2 | Y | 290 | 525 | 107 |  | * | * | * | 1550
2018-11-21 | 325 | 1324 | TN | 61 | * | 40 | 1490 | 90 | 2 | Y | 290 | 525 | 107 |  | * | * | * | 1550
2018-11-22 | 326 | 1520 | TN | 61 | * | 40 | 1500 | 90 | 2 | Y | 290 | 530 | 107 |  | * | * | * | 1510
2018-11-23 | 327 | 1310 | TN | 61 | * | 40 | 1500 | 90 | 2 | Y | 295 | 525 | 107 |  | * | * | * | 1510
2018-11-24 | 328 | 1250 | TN | 61 | * | 40 | 1500 | 90 | 2 | Y | 295 | 525 | 110 |  | * | * | * | 1520
2018-11-25 | 329 | 1803 | TN | 61 | * | 40 | 1505 | 90 | 2 | Y | 295 | 540 | 110 |  | * | * | * | 1515
2018-11-26 | 330 | 1313 | TN | 1 | * | 0 | 1505 | 90 | 2 | Y | 280 | 530 | 108 |  | * | * | * | 1510 | alarm sounding on H2 generator. Float champer lock. Awaiting instructions from PI. Per PI instructions, the backup cylinder was activated to temporarily replace the H2 generator. Cylinder valve opened at 2057Z at 1660psi; tri valve on CUGC unit opened 180 degrees at 2058Z; H2 generator turned off at 2059Z.
2018-11-27 | 331 | 1530 | TN | x | * | x | 1505 | 90 | 2 | Y | 275 | 525 | 108 |  | * | * | * | 1645 | backup cylinder in use; H2 generator is off.
2018-11-28 | 332 | 1532 | TN | x | * | x | 1500 | 90 | 2 | y | 270 | 520 | 107 |  | * | * | * | 1600 | backup cylinder in use; H2 generator is off.
2018-11-29 | 333 | 1648 | TN | 61 | * | 40 | 1495 | 90 | 2 | Y | 260 | 505 | 106 |  | * | * | * | 1600 | H2 generator back online as of 1640Z
2018-11-30 | 334 | 1251 | TN | 61 | * | 40 | 1500 | 90 | 2 | Y | 260 | 520 | 106 |  | * | * | * | 1600
2018-12-01 | 335 | 1330 | TN | 61 | * | 40 | 1495 | 90 | 2 | Y | 270 | 520 | 105 |  | * | * | * | 1580 | added 1 L of DI water to H2 generator reservoir.
2018-12-02 | 336 | 1601 | TN | 61 | * | 40 | 1500 | 90 | 2 | Y | 280 | 530 | 108 |  | * | * | * | 1590
2018-12-03 | 337 | 1338 | TN | 61 | * | 40 | 1505 | 90 | 2 | Y | 280 | 530 | 108 |  | * | * | * | 1595
2018-12-04 | 338 | 1519 | TN | 61 | * | 40 | 1520 | 90 | 2 | Y | 280 | 530 | 108 |  | * | * | * | 1610
2018-12-05 | 339 | 1322 | TN | 61 | * | 40 | 1490 | 90 | 2 | Y | 280 | 525 | 107 |  | * | * | * | 1560
2018-12-06 | 340 | 1631 | TN | 61 | * | 40 | 1510 | 90 | 2 | Y | 275 | 525 | 110 |  | * | * | * | 1580
2018-12-07 | 341 | 1308 | TN | 61 | * | 40 | 1510 | 90 | 2 | Y | 280 | 525 | 110 |  | * | * | * | 1590
2018-12-08 | 342 | 1258 | TN | 61 | * | 40 | 1510 | 90 | 2 | y | 270 | 525 | 110 |  | * | * | * | 1600
2018-12-09 | 343 | 1420 | TN | 61 | * | 20 | 1510 | 90 | 2 | Y | 270 | 520 | 108 |  | * | * | * | 1560 | Hydrogen gen flow down to 20% on bar graph (35 lpm actual flow).
2018-12-10 | 344 | 1314 | TN | 61 | * | 40 | 1505 | 90 | 2 | Y | 270 | 520 | 108 |  | * | * | * | 1550
2018-12-11 | 345 | 1329 | TN | 61 | * | 40 | 1510 | 90 | 2 | Y | 270 | 525 | 111 |  | * | * | * | 1560 | added .5 L DI water to H2 generator
2018-12-12 | 346 | 1324 | TN | 61 | * | 40 | 1480 | 90 | 2 | Y | 270 | 520 | 105 |  | * | * | * | 1550
2018-12-13 | 347 | 1328 | TN | 61 | * | 40 | 1500 | 85 | 2 | Y | 270 | 525 | 106 |  | * | * | * | 1555
2018-12-14 | 348 | 1315 | TN | 61 | * | 40 | 1500 | 85 | 2 | Y | 280 | 525 | 106 |  | * | * | * | 1560
2018-12-15 | 349 | 1714 | TN | 61 | * | 40 | 1500 | 85 | 2 | Y | 275 | 525 | 107 |  | * | * | * | 1550
2018-12-16 | 350 | 1443 | TN | 61 | * | 40 | 1495 | 85 | 2 | Y | 275 | 525 | 107 |  | * | * | * | 1550
2018-12-17 | 351 | 1252 | TN | 61 | * | 40 | 1495 | 85 | 2 | Y | 275 | 520 | 105 |  | * | * | * | 1545
2018-12-18 | 352 | 1314 | TN | 61 | * | 40 | 1490 | 85 | 2 | Y | 275 | 525 | 106 |  | * | * | * | 1550
2018-12-19 | 353 | 1302 | TN | 61 | * | 40 | 1490 | 85 | 2 | Y | 270 | 530 | 106 |  | * | * | * | 1550
2018-12-20 | 354 | 1258 | TN | 61 | * | 40 | 1495 | 85 | 2 | Y | 275 | 525 | 107 |  | * | * | * | 1550
2018-12-21 | 355 | 1316 | TN | 61 | * | 40 | 1495 | 85 | 2 | Y | 275 | 520 | 107 |  | * | * | * | 1540
2018-12-22 | 356
2018-12-23 | 357 | 0 | 0
2018-12-24 | 358 | 1301 | TN | 61 | * | 40 | 1490 | 85 | 2 | Y | 275 | 525 | 106 |  | * | * | * | 1500
2018-12-25 | 359 | 1354 | TN | 61 | * | 40 | 1490 | 85 | 2 | Y | 270 | 525 | 105 |  | * | * | * | 1505
2018-12-26 | 360 | 1322 | TN | 61 | * | 40 | 1490 | 85 | 2 | Y | 270 | 520 | 106 |  | * | * | * | 1520
2018-12-27 | 361 | 1305 | TN | 61 | * | 40 | 1490 | 85 | 2 | Y | 270 | 520 | 107 |  | * | * | * | 1510
2018-12-28 | 362 | 0 | 0
2018-12-29 | 363 | 1327 | TN | 61 | * | 40 | 1490 | 85 | 2 | y | 260 | 515 | 107 |  | * | * | * | 1495
2018-12-30 | 364 | 0 | 0
2018-12-31 | 365 | 1254 | TN | 61 | * | 40 | 1490 | 85 | 2 | Y | 270 | 520 | 107 |  | * | * | * | 1470 | added 0.5 l DI water to H2 generator

## Aethalometer AE16-AE33
Date | Julian Day | Time (UTC) | Observer's Initials | AE33 Sample inlet? | Panel Readout BC (ng/m3) | Flow rate, (lpm) | Weekly Tape Check | Comments
2018-01-01 | 1 | 1719 | HA | rm frost | 130 | 4 | *
2018-01-02 | 2 | 0 | 0 |  |  |  |  | Could not visit TAWO due to weather
2018-01-03 | 3 | 0 | 0 |  |  |  |  | Could not visit TAWO due to weather
2018-01-04 | 4 | 0 | 0 |  |  |  |  | Could not visit TAWO due to weather
2018-01-05 | 5 | 1454 | HA | rm hvy frost | 142 | 4 | * | Completed biannual maintenance procedures: inspecting optical chamber, flow check, calibration, and verification, lubricate optical chamber sliders. Instrument did not restart after these procedures, contacted Betsy and in the process of troubleshooting. Instrument currently not operating.
2018-01-06 | 6 | 0 | 0 |  |  |  |  | AE33 not operating
2018-01-07 | 7 | 0 | 0 |  |  |  |  | AE33 not operating
2018-01-08 | 8 | 1423 | HA | rm frost | - | - | * | CPD3 laptop shutdown from 1640-1645.
2018-01-09 | 9 | 0 | 0
2018-01-10 | 10 | 1432 | HA |  |  |  |  | AE33 being shutdown and removed from system to return to Boulder
2018-01-11 | 11 | 0 | 0
2018-01-12 | 12 | 1440 | HA
2018-01-13 | 13 | 0 | 0
2018-01-14 | 14 | 0 | 0
2018-01-15 | 15 | 1413 | HA
2018-01-16 | 16 | 0 | 0
2018-01-17 | 17 | 0 | 0
2018-01-18 | 18 | 1407 | HA
2018-01-19 | 19 | 0 | 0
2018-01-20 | 20 | 1323 | HA
2018-01-21 | 21 | 0 | 0
2018-01-22 | 22 | 1758 | HA
2018-01-23 | 23 | 0 | 0
2018-01-24 | 24 | 1409 | HA
2018-01-25 | 25 | 0 | 0
2018-01-26 | 26 | 0 | 0
2018-01-27 | 27 | 0 | 0
2018-01-28 | 28 | 0 | 0
2018-01-29 | 29 | 1432 | HA
2018-01-30 | 30 | 0 | 0
2018-01-31 | 31 | 1434 | HA
2018-02-01 | 32 | 0 | 0
2018-02-02 | 33 | 0 | 0
2018-02-03 | 34 | 1411 | HA
2018-02-04 | 35 | 0 | 0
2018-02-05 | 36 | 1424 | HA
2018-02-06 | 37 | 1428 | HA
2018-02-07 | 38 | 1455 | HA
2018-02-08 | 39 | 0 | 0
2018-02-09 | 40 | 1314 | HA
2018-02-10 | 41 | 0 | 0
2018-02-11 | 42 | 0 | 0
2018-02-12 | 43 | 1458 | HA, SWD
2018-02-13 | 44 | 0 | 0
2018-02-14 | 45 | 1418 | SWD
2018-02-15 | 46 | 0 | 0
2018-02-16 | 47 | 0 | 0
2018-02-17 | 48 | 1353 | SWD
2018-02-18 | 49 | 0 | 0
2018-02-19 | 50 | 0 | 0
2018-02-20 | 51 | 0 | 0
2018-02-21 | 52 | 1428 | SWD
2018-02-22 | 53 | 0 | 0
2018-02-23 | 54 | 0 | 0
2018-02-24 | 55 | 0 | 0
2018-02-25 | 56 | 1503 | SWD
2018-02-26 | 57 | 1539 | SWD
2018-02-27 | 58 | 1851 | SWD
2018-02-28 | 59 | 0 | 0
2018-03-01 | 60 | 0 | 0
2018-03-02 | 61 | 1657 | SWD
2018-03-03 | 62 | 0 | 0
2018-03-04 | 63 | 0 | 0
2018-03-05 | 64 | 1439 | SWD
2018-03-06 | 65 | 0 | 0
2018-03-07 | 66 | 1517 | SWD
2018-03-08 | 67 | 0 | 0
2018-03-09 | 68 | 0 | 0
2018-03-10 | 69 | 1243 | SWD
2018-03-11 | 70 | 0 | 0
2018-03-12 | 71 | 1454 | SWD
2018-03-13 | 72 | 0 | 0
2018-03-14 | 73 | 1529 | SWD
2018-03-15 | 74 | 0 | 0
2018-03-16 | 75 | 1244 | SWD
2018-03-17 | 76 | 0 | 0
2018-03-18 | 77 | 0 | 0
2018-03-19 | 78 | 1459 | SWD
2018-03-20 | 79 | 0 | 0
2018-03-21 | 80 | 1320 | SWD
2018-03-22 | 81 | 0 | 0
2018-03-23 | 82 | 1300 | SWD/RN
2018-03-24 | 83 | 0 | 0
2018-03-25 | 84 | 1227 | SWD
2018-03-26 | 85 | 0 | 0
2018-03-27 | 86 | 0 | 0
2018-03-28 | 87 | 1503 | NB
2018-03-29 | 88 | 0 | 0
2018-03-30 | 89 | 1722 | NB
2018-03-31 | 90 | 0 | 0
2018-04-01 | 91 | 1157 | NB
2018-04-02 | 92 | 1306 | NB
2018-04-03 | 93 | 0 | 0
2018-04-04 | 94 | 0 | 0
2018-04-05 | 95 | 1203 | NB
2018-04-06 | 96 | 0 | 0
2018-04-07 | 97 | 1122 | NB
2018-04-08 | 98 | 0 | 0
2018-04-09 | 99 | 1135 | NB
2018-04-10 | 100 | 0 | 0
2018-04-11 | 101 | 0 | 0
2018-04-12 | 102 | 1130 | NB
2018-04-13 | 103 | 0 | 0
2018-04-14 | 104 | 1314 | NB
2018-04-15 | 105 | 0 | 0
2018-04-16 | 106 | 1127 | NB
2018-04-17 | 107 | 0 | 0
2018-04-18 | 108 | 1145 | NB
2018-04-19 | 109 | 0 | 0
2018-04-20 | 110 | 1141 | NB
2018-04-21 | 111 | 0 | 0
2018-04-22 | 112 | 1207 | NB
2018-04-23 | 113 | 1144 | NB
2018-04-24 | 114 | 0 | 0
2018-04-25 | 115 | 0 | 0
2018-04-26 | 116 | 1130 | NB
2018-04-27 | 117 | 1327 | NB/MB
2018-04-28 | 118 | 0 | 0
2018-04-29 | 119 | 0 | 0
2018-04-30 | 120 | 1228 | NB/MB
2018-05-01 | 121 | 1224 | MB
2018-05-02 | 122 | 0 | 0
2018-05-03 | 123 | 0 | 0
2018-05-04 | 124 | 1227 | MB
2018-05-05 | 125 | 1217 | MB
2018-05-06 | 126 | 0 | 0
2018-05-07 | 127 | 1224 | MB
2018-05-08 | 128 | 1247 | MB
2018-05-09 | 129 | 0 | 0
2018-05-10 | 130 | 1247 | MB
2018-05-11 | 131 | 0 | 0
2018-05-12 | 132 | 1721 | MB
2018-05-13 | 133 | 0 | 0
2018-05-14 | 134 | 1153 | MB
2018-05-15 | 135 | 1515 | MB
2018-05-16 | 136 | 0 | 0
2018-05-17 | 137 | 1205 | MB
2018-05-18 | 138 | 0 | 0
2018-05-19 | 139 | 1219 | MB
2018-05-20 | 140 | 0 | 0
2018-05-21 | 141 | 1658 | MB
2018-05-22 | 142 | 1859 | MB
2018-05-23 | 143 | 1243 | MB
2018-05-24 | 144 | 1226 | MB
2018-05-25 | 145 | 0 | 0
2018-05-26 | 146 | 0 | 0
2018-05-27 | 147 | 0 | 0
2018-05-28 | 148 | 1220 | MB
2018-05-29 | 149 | 1347 | MB
2018-05-30 | 150 | 1201 | MB
2018-05-31 | 151 | 0 | 0
2018-06-01 | 152 | 1129 | MB
2018-06-02 | 153 | 0 | 0
2018-06-03 | 154 | 1203 | MB
2018-06-04 | 155 | 0 | 0
2018-06-05 | 156 | 1302 | MB/MM
2018-06-06 | 157 | 1341 | MB/MM
2018-06-07 | 158 | 1209 | MB/MM
2018-06-08 | 159 | 1220 | MM
2018-06-09 | 160 | 1647 | MM
2018-06-10 | 161 | 0 | 0
2018-06-11 | 162 | 1659 | MB/MM
2018-06-12 | 163 | 1802 | MM
2018-06-13 | 164 | 0 | 0
2018-06-14 | 165 | 0 | 0
2018-06-15 | 166 | 1828 | MM
2018-06-16 | 167 | 1746 | MM
2018-06-17 | 168 | 0 | 0
2018-06-18 | 169 | 0 | 0
2018-06-19 | 170 | 1326 | MM
2018-06-20 | 171 | 2023 | MM
2018-06-21 | 172 | 0 | 0
2018-06-22 | 173 | 1358 | MM
2018-06-23 | 174 | 0 | 0
2018-06-24 | 175 | 0 | 0
2018-06-25 | 176 | 1802 | MM
2018-06-26 | 177 | 1849 | MM
2018-06-27 | 178 | 1649 | MM
2018-06-28 | 179 | 1558 | MM
2018-06-29 | 180 | 0 | 0
2018-06-30 | 181 | 2047 | MM
2018-07-01 | 182 | 0 | 0
2018-07-02 | 183 | 0 | 0
2018-07-03 | 184 | 1603 | MM
2018-07-04 | 185 | 0 | 0
2018-07-05 | 186 | 1424 | MM
2018-07-06 | 187 | 0 | 0
2018-07-07 | 188 | 0 | 0
2018-07-08 | 189 | 0 | 0
2018-07-09 | 190 | 0 | 0
2018-07-10 | 191 | 1223 | MM
2018-07-11 | 192 | 1811 | MM
2018-07-12 | 193 | 0 | 0
2018-07-13 | 194 | 0 | 0
2018-07-14 | 195 | 0 | 0
2018-07-15 | 196 | 0 | 0
2018-07-16 | 197 | 2209 | MM
2018-07-17 | 198 | 0 | 0
2018-07-18 | 199 | 1430 | MM
2018-07-19 | 200 | 0 | 0
2018-07-20 | 201 | 0 | 0
2018-07-21 | 202 | 1317 | MM
2018-07-22 | 203 | 0 | 0
2018-07-23 | 204 | 0 | 0
2018-07-24 | 205 | 0 | 0
2018-07-25 | 206 | 2335 | MM
2018-07-26 | 207 | 0 | 0
2018-07-27 | 208 | 204 | MM
2018-07-28 | 209 | 0 | 0
2018-07-29 | 210 | 258 | MM
2018-07-30 | 211 | 1234 | MM
2018-07-31 | 212 | 0 | 0
2018-08-01 | 213 | 0 | 0
2018-08-02 | 214 | 2147 | MM
2018-08-03 | 215 | 1810 | MM
2018-08-04 | 216 | 1430 | MM
2018-08-05 | 217 | 0 | 0
2018-08-06 | 218 | 0 | 0
2018-08-07 | 219 | 2037 | MM
2018-08-08 | 220 | 0 | 0
2018-08-09 | 221 | 1524 | MM
2018-08-10 | 222 | 0 | 0
2018-08-11 | 223 | 0 | 0
2018-08-12 | 224 | 0 | 0
2018-08-13 | 225 | 2200 | MM
2018-08-14 | 226 | 0 | 0
2018-08-15 | 227 | 0 | 0
2018-08-16 | 228 | 0 | 0
2018-08-17 | 229 | 1614 | MM
2018-08-18 | 230 | 0 | 0
2018-08-19 | 231 | 0 | 0
2018-08-20 | 232 | 0 | 0
2018-08-21 | 233 | 1746 | MM
2018-08-22 | 234 | 0 | 0
2018-08-23 | 235 | 0 | 0
2018-08-24 | 236 | 0 | 0
2018-08-25 | 237 | 1930 | MM
2018-08-26 | 238 | 1747 | MM
2018-08-27 | 239 | 1251 | MM
2018-08-28 | 240 | 0 | 0
2018-08-29 | 241 | 0 | 0
2018-08-30 | 242 | 0 | 0
2018-08-31 | 243 | 1433 | MM
2018-09-01 | 244 | 0 | 0
2018-09-02 | 245 | 0 | 0
2018-09-03 | 246 | 0 | 0
2018-09-04 | 247 | 1613 | MM
2018-09-05 | 248 | 1700 | MM
2018-09-06 | 249 | 0 | 0
2018-09-07 | 250 | 0 | 0
2018-09-08 | 251 | 0 | 0
2018-09-09 | 252 | 1632 | MM
2018-09-10 | 253 | 0 | 0
2018-09-11 | 254 | 0 | 0
2018-09-12 | 255 | 1414 | MM
2018-09-13 | 256 | 0 | 0
2018-09-14 | 257 | 0 | 0
2018-09-15 | 258 | 1950 | MM
2018-09-16 | 259 | 0 | 0
2018-09-17 | 260 | 0 | 0
2018-09-18 | 261 | 1445 | MM
2018-09-19 | 262 | 0 | 0
2018-09-20 | 263 | 0 | 0
2018-09-21 | 264 | 1900 | MM
2018-09-22 | 265 | 0 | 0
2018-09-23 | 266 | 0 | 0
2018-09-24 | 267 | 1925 | MM
2018-09-25 | 268 | 0 | 0
2018-09-26 | 269 | 0 | 0
2018-09-27 | 270 | 0 | 0
2018-09-28 | 271 | 2000 | MM
2018-09-29 | 272 | 0 | 0
2018-09-30 | 273 | 0 | 0
2018-10-01 | 274 | 0 | 0
2018-10-02 | 275 | 1937 | MM
2018-10-03 | 276 | 0 | 0
2018-10-04 | 277 | 0 | 0
2018-10-05 | 278 | 1747 | MM
2018-10-06 | 279 | 0 | 0
2018-10-07 | 280 | 0 | 0
2018-10-08 | 281 | 1843 | MM
2018-10-09 | 282 | 0 | 0
2018-10-10 | 283 | 0 | 0
2018-10-11 | 284 | 0 | 0
2018-10-12 | 285 | 1124 | MM
2018-10-13 | 286 | 0 | 0
2018-10-14 | 287 | 0 | 0
2018-10-15 | 288 | 1800 | MM
2018-10-16 | 289 | 1750 | MM/TN
2018-10-17 | 290 | 0 | 0
2018-10-18 | 291 | 1755 | MM/TN
2018-10-19 | 292 | 1848 | MM/TN
2018-10-20 | 293 | 0 | 0
2018-10-21 | 294 | 1630 | TN
2018-10-22 | 295 | 1616 | TN
2018-10-23 | 296 | 1656 | TN
2018-10-24 | 297 | 1710 | TN
2018-10-25 | 298 | 1711 | TN
2018-10-26 | 299 | 1640 | TN
2018-10-27 | 300 | 1536 | TN
2018-10-28 | 301 | 1526 | TN
2018-10-29 | 302 | 1658 | TN
2018-10-30 | 303 | 1506 | TN
2018-10-31 | 304 | 1445 | TN
2018-11-01 | 305 | 1742 | TN
2018-11-02 | 306 | 1645 | TN
2018-11-03 | 307 | 1813 | TN
2018-11-04 | 308 | 1508 | TN
2018-11-05 | 309 | 1247 | TN
2018-11-06 | 310 | 1240 | TN
2018-11-07 | 311 | 1255 | TN
2018-11-08 | 312 | 1547 | TN
2018-11-09 | 313 | 1325 | TN
2018-11-10 | 314 | 1244 | TN
2018-11-11 | 315 | 1616 | TN
2018-11-12 | 316
2018-11-13 | 317 | 1240 | MG/NB
2018-11-14 | 318 | 1336 | TN
2018-11-15 | 319 | 1355 | TN
2018-11-16 | 320 | 1914 | TN
2018-11-17 | 321 | 1228 | TN
2018-11-18 | 322 | 1630 | TN
2018-11-19 | 323 | 1234 | TN
2018-11-20 | 324 | 1258 | TN
2018-11-21 | 325 | 1324 | TN
2018-11-22 | 326 | 1520 | TN
2018-11-23 | 327 | 1310 | TN
2018-11-24 | 328 | 1250 | TN
2018-11-25 | 329 | 1803 | TN
2018-11-26 | 330 | 1313 | TN
2018-11-27 | 331 | 1530 | TN
2018-11-28 | 332 | 1532 | TN
2018-11-29 | 333 | 1555 | TN
2018-11-30 | 334 | 1251 | TN
2018-12-01 | 335 | 1330 | TN
2018-12-02 | 336 | 1601 | TN
2018-12-03 | 337 | 1338 | TN
2018-12-04 | 338 | 1519 | TN
2018-12-05 | 339 | 1322 | TN
2018-12-06 | 340 | 1631 | TN
2018-12-07 | 341 | 1308 | TN
2018-12-08 | 342 | 1258 | TN
2018-12-09 | 343 | 1420 | TN
2018-12-10 | 344 | 1314 | TN
2018-12-11 | 345 | 1329 | TN
2018-12-12 | 346 | 1324 | TN
2018-12-13 | 347 | 1328 | TN
2018-12-14 | 348 | 1315 | TN
2018-12-15 | 349 | 1714 | TN
2018-12-16 | 350 | 1443 | TN
2018-12-17 | 351 | 1252 | TN
2018-12-18 | 352 | 1314 | TN
2018-12-19 | 353 | 1302 | TN
2018-12-20 | 354 | 1258 | TN
2018-12-21 | 355 | 1316 | TN
2018-12-22 | 356 | 1301 | TN
2018-12-23 | 357 | 0 | 0
2018-12-24 | 358 | #REF! | #REF!
2018-12-25 | 359 | 1354 | TN
2018-12-26 | 360 | 1322 | TN
2018-12-27 | 361 | 1305 | TN
2018-12-28 | 362 | 0 | 0
2018-12-29 | 363 | 1327 | TN
2018-12-30 | 364 | 0 | 0
2018-12-31 | 365 | 1254 | TN
| 366 | 0 | 0

## Met Tower
Date | Julian Day | Time (UTC) | Observer's Initials | 2m Temp | 2m TD | 10m SONIC | HATS Flask Inlet (inlet tower) | Sfc O3 (inlet tower) | CU GC | Check Enclosure Temp (~12 C) | Check date and time | Greenhouse Inlet Check | Comments (Climbing, removing instruments, etc)
2018-01-01 | 1 | 1719 | HA | rm lt frost | rm light frost | clr |  | rm lt frost | rm lt frost |  | * | rm frost
2018-01-02 | 2 | 0 | 0 | [8 empty] | rm frost | Could not visit TAWO due to weather
2018-01-03 | 3 | 0 | 0 | [8 empty] | rm frost | Could not visit TAWO due to weather
2018-01-04 | 4 | 0 | 0 | [8 empty] | rm frost | Could not visit TAWO due to weather
2018-01-05 | 5 | 1454 | HA | rm hvy frost | rm hvy frost | clr |  | rm hvy frost | rm hvy frost |  | * | rm frost
2018-01-06 | 6 | 0 | 0 | [8 empty] | rm frost
2018-01-07 | 7 | 0 | 0 | [8 empty] | rm frost
2018-01-08 | 8 | 1423 | HA | rm frost | rm frost | clr |  | rm frost | rm frost |  | * | rm frost
2018-01-09 | 9 | 0 | 0
2018-01-10 | 10 | 0 | 0
2018-01-11 | 11 | 0 | 0
2018-01-12 | 12 | 1440 | HA | rm frost | rm frost | clr |  | rm frost | rm frost |  | * | rm frost
2018-01-13 | 13 | 0 | 0
2018-01-14 | 14 | 0 | 0
2018-01-15 | 15 | 1413 | HA | rm frost | rm frost | clr |  | rm frost | rm frost |  | * | rm frost
2018-01-16 | 16 | 0 | 0
2018-01-17 | 17 | 0 | 0
2018-01-18 | 18 | 1407 | HA | rm frost | rm frost | clr |  | rm frost | rm frost |  | * | rm frost
2018-01-19 | 19 | 0 | 0
2018-01-20 | 20 | 1323 | HA | rm frost | rm frost | clr |  | rm frost | rm frost |  | * | rm frost | Unexpected power outage from 18:21-18:22Z and 19:00-19:04Z
2018-01-21 | 21 | 0 | 0
2018-01-22 | 22 | 1758 | HA | rm frost | rm frost | clr |  | rm frost | rm frost |  | * | rm frost
2018-01-23 | 23 | 0 | 0
2018-01-24 | 24 | 1409 | HA | rm frost | rm frost | clr |  | rm frost | rm frost |  | * | rm frost
2018-01-25 | 25 | 0 | 0
2018-01-26 | 26 | 0 | 0
2018-01-27 | 27 | 0 | 0
2018-01-28 | 28 | 0 | 0
2018-01-29 | 29 | 1432 | HA | rm frost | rm frost | clr |  | rm frost | rm frost |  | * | rm frost
2018-01-30 | 30 | 0 | 0
2018-01-31 | 31 | 1434 | HA | rm frost | rm frost | clr |  | rm frost | rm frost |  | * | rm frost
2018-02-01 | 32 | 0 | 0
2018-02-02 | 33 | 0 | 0
2018-02-03 | 34 | 1411 | HA | rm frost | rm frost | clr |  | rm frost | rm frost |  | * | rm frost
2018-02-04 | 35 | 0 | 0
2018-02-05 | 36 | 1424 | HA | rm frost | rm frost | clr |  | rm frost | rm frost |  | * | rm frost | Sonic anemometer wind speed and direction way off; will look into this matter further when winds die down.
2018-02-06 | 37 | 1428 | HA | [9 empty] | Climbed tower 1350-1415 UTC; cleaned sonic anemometer.
2018-02-07 | 38 | 1612 | HA | rm frost | rm frost | clr |  | rm frost | rm frost |  | * | rm frost
2018-02-08 | 39 | 0 | 0
2018-02-09 | 40 | 1314 | HA | rm frost | rm frost | clr |  | rm frost | rm frost |  | * | rm frost
2018-02-10 | 41 | 0 | 0
2018-02-11 | 42 | 0 | 0
2018-02-12 | 43 | 1458 | HA, SWD | rm frost | rm frost | clr |  | rm frost | rm frost |  | * | rm frost
2018-02-13 | 44 | 0 | 0
2018-02-14 | 45 | 1418 | SWD | rm frost | rm frost | clr | rm frost | rm frost | rm frost |  | * | rm frost
2018-02-15 | 46 | 0 | 0
2018-02-16 | 47 | 0 | 0
2018-02-17 | 48 | 1353 | SWD | rm hf | rm hf | clr | rm lf | rm lf | rm lf |  | * | rm lf
2018-02-18 | 49 | 0 | 0
2018-02-19 | 50 | 0 | 0
2018-02-20 | 51 | 0 | 0
2018-02-21 | 52 | 1428 | SWD | rm hf | rm hf | clr | rm lf | rm lf | rm lf |  | * | rm lf
2018-02-22 | 53 | 0 | 0
2018-02-23 | 54 | 0 | 0
2018-02-24 | 55 | 0 | 0
2018-02-25 | 56 | 1503 | SWD | rm hf | rm hf | clr | rm lf | rm lf | rm lf |  | * | rm lf
2018-02-26 | 57 | 1539 | SWD | clr | clr | clr | rm lf | rm lf | rm lf |  | * | clr
2018-02-27 | 58 | 1851 | SWD | rm hf | rm hf | clr | rm lf | rm lf | rm lf |  |  | clr
2018-02-28 | 59 | 0 | 0 | [8 empty] | rm lf
2018-03-01 | 60 | 0 | 0 | [8 empty] | rm lf
2018-03-02 | 61 | 1657 | SWD | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | rm lf
2018-03-03 | 62 | 0 | 0 | [8 empty] | rm lf
2018-03-04 | 63 | 0 | 0 | [8 empty] | rm lf
2018-03-05 | 64 | 1439 | SWD | rm lf | rm lf | clr | rm hf | rm hf | rm hf | * | * | rm lf
2018-03-06 | 65 | 0 | 0
2018-03-07 | 66 | 1517 | SWD | rm hf | rm hf | clr | rm lf | rm hf | rm hf | * | * | rm hf
2018-03-08 | 67 | 0 | 0
2018-03-09 | 68 | 0 | 0
2018-03-10 | 69 | 1243 | SWD | rm hf | rm hf | clr | rm lf | rm lf | rm hf | * | * | rm hf
2018-03-11 | 70 | 0 | 0 | [8 empty] | rm lf
2018-03-12 | 71 | 1454 | SWD | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | rm lf
2018-03-13 | 72 | 0 | 0 | [8 empty] | rm lf
2018-03-14 | 73 | 1529 | SWD | rm lf | rm lf | clr | rm lf | rm hf | rm hf | * | * | rm lf
2018-03-15 | 74 | 0 | 0 | [8 empty] | rm lf
2018-03-16 | 75 | 1244 | SWD | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | rm lf
2018-03-17 | 76 | 0 | 0 | [8 empty] | rm lf
2018-03-18 | 77 | 0 | 0 | [8 empty] | rm lf
2018-03-19 | 78 | 1459 | SWD | rm hf | rm hf | clr | rm hf | rm hf | rm hf | * | * | rm hf
2018-03-20 | 79 | 0 | 0 | [8 empty] | rm lf
2018-03-21 | 80 | 1320 | SWD | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-03-22 | 81 | 0 | 0 | [8 empty] | clr
2018-03-23 | 82 | 1300 | SWD/RN | rm lf | rm hf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-03-24 | 83 | 0 | 0 | [8 empty] | clr
2018-03-25 | 84 | 1227 | SWD | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-03-26 | 85 | 0 | 0
2018-03-27 | 86 | 0 | 0
2018-03-28 | 87 | 1503 | NB | rm hf | rm hf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-03-29 | 88 | 0 | 0 | [8 empty] | rm lf
2018-03-30 | 89 | 1722 | NB | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | rm lf
2018-03-31 | 90 | 0 | 0 | [8 empty] | rm lf
2018-04-01 | 91 | 1157 | NB | rm lf | rm lf | clr | rm hf | rm hf | rm hf | * | * | rm lf
2018-04-02 | 92 | 1306 | NB | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | rm lf
2018-04-03 | 93 | 0 | 0
2018-04-04 | 94 | 0 | 0
2018-04-05 | 95 | 1203 | NB | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | rm lf
2018-04-06 | 96 | 0 | 0
2018-04-07 | 97 | 1122 | NB | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | rm lf
2018-04-08 | 98 | 0 | 0
2018-04-09 | 99 | 1135 | NB | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-04-10 | 100 | 0 | 0
2018-04-11 | 101 | 0 | 0
2018-04-12 | 102 | 1130 | NB | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | rm lf
2018-04-13 | 103 | 0 | 0
2018-04-14 | 104 | 1314 | NB | clr | clr | clr | clr | clr | clr | * | * | clr
2018-04-15 | 105 | 0 | 0
2018-04-16 | 106 | 1127 | NB | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-04-17 | 107 | 0 | 0
2018-04-18 | 108 | 1145 | NB | rm hf | rm hf | clr | rm hf | rm hf | rm hf | * | * | rm lf
2018-04-19 | 109 | 0 | 0
2018-04-20 | 110 | 1141 | NB | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | rm lf
2018-04-21 | 111 | 0 | 0
2018-04-22 | 112 | 1207 | NB | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | rm lf
2018-04-23 | 113 | 1144 | NB | rm lf | rm lf | clr | rm hf | rm hf | rm hf | * | * | rm lf
2018-04-24 | 114 | 0 | 0
2018-04-25 | 115 | 0 | 0
2018-04-26 | 116 | 1130 | NB | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | rm lf
2018-04-27 | 117 | 1327 | NB/MB | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-04-28 | 118 | 0 | 0 | [8 empty] | clr
2018-04-29 | 119 | 0 | 0 | [8 empty] | clr
2018-04-30 | 120 | 1228 | NB/MB | clr | clr | clr | clr | clr | clr | * | * | clr
2018-05-01 | 121 | 1224 | MB | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-05-02 | 122 | 0 | 0 | [8 empty] | clr
2018-05-03 | 123 | 0 | 0 | [8 empty] | clr
2018-05-04 | 124 | 1227 | MB | rm snow | rm snow | clr | clr | clr | clr | * | * | clr
2018-05-05 | 125 | 1217 | MB | rm lf | rm lf | clr | rm hf | rm hf | rm hf | * | * | clr
2018-05-06 | 126 | 0 | 0 | [8 empty] | clr
2018-05-07 | 127 | 1224 | MB | clr | clr | clr | clr | clr | clr | * | * | clr
2018-05-08 | 128 | 1247 | MB | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-05-09 | 129 | 0 | 0 | [8 empty] | clr
2018-05-10 | 130 | 1247 | MB | rm lf | rm lf | clr | rm lf | rm lf | rm hf | * | * | clr
2018-05-11 | 131 | 0 | 0 | [8 empty] | clr
2018-05-12 | 132 | 1721 | MB | rm snow | rm snow | clr | clr | clr | clr | * | * | clr
2018-05-13 | 133 | 0 | 0 | [8 empty] | clr
2018-05-14 | 134 | 1153 | MB | clr | clr | clr | clr | rm lf | rm lf | * | * | clr
2018-05-15 | 135 | 1515 | MB | rm hf | rm hf | clr | clr | rm lf | rm lf | * | * | rm lf
2018-05-16 | 136 | 0 | 0 | [8 empty] | clr
2018-05-17 | 137 | 1205 | MB | rm lf | rm lf | clr | clr | rm lf | rm lf | * | * | clr
2018-05-18 | 138 | 0 | 0 | [8 empty] | clr
2018-05-19 | 139 | 1219 | MB | clr | clr | clr | clr | lf | lf | * | * | clr
2018-05-20 | 140 | 0 | 0 | [8 empty] | clr
2018-05-21 | 141 | 1658 | MB | rm lf | rm lf | clr | rm lf | rm hf | rm hf | * | * | rm lf
2018-05-22 | 142 | 1859 | MB | rm lf | clr | clr | clr | rm lf | rm hf | * | * | clr
2018-05-23 | 143 | 1243 | MB | rm lf | rm lf | clr | rm lf | rm lf | rm hf | * | * | clr
2018-05-24 | 144 | 1226 | MB | rm lf | clr | clr | clr | rm lf | rm lf | * | * | clr
2018-05-25 | 145 | 0 | 0 | [8 empty] | clr
2018-05-26 | 146 | 0 | 0 | [8 empty] | clr
2018-05-27 | 147 | 0 | 0 | [8 empty] | clr
2018-05-28 | 148 | 1220 | MB | rm lf | rm lf | clr | clr | rm hf | rm hf | * | * | clr | Power-cycled the Met box at 1257z.
2018-05-29 | 149 | 1347 | MB | rm lf | rm lf | clr | clr | rm lf | rm lf | * | * | clr
2018-05-30 | 150 | 1201 | MB | clr | clr | clr | clr | clr | clr | * | * | clr
2018-05-31 | 151 | 0 | 0 | [8 empty] | clr
2018-06-01 | 152 | 1129 | MB | clr | clr | clr | clr | clr | clr | * | * | clr
2018-06-02 | 153 | 0 | 0 | [8 empty] | clr
2018-06-03 | 154 | 1203 | MB | clr | clr | clr | clr | clr | clr | * | * | clr
2018-06-04 | 155 | 0 | 0 | [8 empty] | clr
2018-06-05 | 156 | 1302 | MB/MM | rm snow | rm snow | clr | clr | clr | clr | * | * | clr
2018-06-06 | 157 | 1430 | MB/MM | clr | rm snow | clr | clr | clr | clr | * | * | clr
2018-06-07 | 158 | 1209 | MB/MM | clr | clr | clr | clr | clr | clr | * | * | clr
2018-06-08 | 159 | 1220 | MM | clr | clr | clr | clr | clr | clr | * | * | clr
2018-06-09 | 160 | 1647 | MM | clr | clr | clr | clr | clr | clr | * | * | clr | Climbed to the top of the tower to remove snow/ice accumulation and photograph the tower top from approximately 1730 - 1800 UTC
2018-06-10 | 161 | 0 | 0
2018-06-11 | 162 | 1659 | MB/MM | rm lf | rm lf | clr | clr | rm hf | rm hf | * | * | clr
2018-06-12 | 163 | 1400 | MM | rm hf | rm hf | clr | rm lf | rm hf | rm hf | * | * | rm lf | Unplugged MET enclosure heater at 1400z
2018-06-13 | 164 | 0 | 0
2018-06-14 | 165 | 0 | 0
2018-06-15 | 166 | 1828 | MM | clr | clr | clr | clr | rm lf | rm lf | * | * | clr | MET instruments ran on the UPS during a planned power outage from 1700z - 1855z
2018-06-16 | 167 | 1746 | MM | clr | clr | clr | clr | clr | clr | * | * | clr
2018-06-17 | 168 | 0 | 0
2018-06-18 | 169 | [11 empty] | MET instruments ran on the UPS during two brief planned power outages from 0115z-0130z and 0645z-0650z
2018-06-19 | 170 | 1322 | MM | clr | clr | clr | clr | clr | clr | * | * | clr
2018-06-20 | 171 | 2023 | MM | clr | clr | clr | clr | clr | clr | * | * | clr
2018-06-21 | 172 | 0 | 0
2018-06-22 | 173 | 1358 | MM | clr | clr | clr | clr | clr | clr | * | * | clr
2018-06-23 | 174 | 0 | 0
2018-06-24 | 175 | 0 | 0
2018-06-25 | 176 | 1802 | MM | rm lf | rm lf | clr | clr | rm lf | rm hf | * | * | clr
2018-06-26 | 177 | 1849 | MM | clr | clr | clr | clr | clr | clr | - | * | clr | Internet outage due to Big House construction: 6/26 from 1515 -2136 and from 2345 to 0230 UTC on 6/27-- couldn't check enclosure temp
2018-06-27 | 178 | 1649 | MM | rm hf | rm hf | clr | clr | rm hf | rm hf | * | * | clr | Opened MET box and plugged in heater from 1622-1638 utc to find out current capacity for the TAWO tower
2018-06-28 | 179 | 1558 | MM | clr | clr | clr | clr | clr | clr | * | * | clr
2018-06-29 | 180 | 0 | 0
2018-06-30 | 181 | 2047 | MM | clr | clr | clr | clr | clr | rm hf | * | * | clr
2018-07-01 | 182 | 0 | 0
2018-07-02 | 183 | 0 | 0
2018-07-03 | 184 | 1603 | MM | clr | clr | clr | clr | clr | clr | * | * | clr
2018-07-04 | 185 | 0 | 0
2018-07-05 | 186 | 1424 | MM | clr | clr | clr | clr | rm lf | rm lf | * | * | clr
2018-07-06 | 187 | 0 | 0
2018-07-07 | 188 | 0 | 0
2018-07-08 | 189 | 0 | 0
2018-07-09 | 190 | 0 | 0
2018-07-10 | 191 | 1223 | MM | clr | clr | clr | clr | clr | clr | * | * | clr
2018-07-11 | 192 | 1811 | MM | clr | clr | clr | clr | clr | clr | * | * | clr
2018-07-12 | 193 | 0 | 0
2018-07-13 | 194 | 0 | 0 | [9 empty] | Raised 2m tower instruments, turned off power to the tower at 1258z
2018-07-14 | 195 | 0 | 0 | [9 empty] | Raised 2m tower instruments
2018-07-15 | 196 | 0 | 0 | [9 empty] | Power returned to tower, instruments powered back up at 01:47z; humidity sensor reading incorrectly ~35C
2018-07-16 | 197 | 2209 | MM | clr | clr | clr | clr | clr | rm lf | * | * | clr | Humidity sensor reading incorrectly ~35C
2018-07-17 | 198 | 0 | 0 | [9 empty] | Riggers plumbed the tower and tightened the guidelines; Power cycled the logger box while correcting the humidity sensor connector 2132-2205z
2018-07-18 | 199 | 1410 | MM | clr | clr | clr | clr | clr | clr | * | * | clr | Riggers plumbed the tower and tightened the guidelines 1100z-1400z
2018-07-19 | 200 | 0 | 0
2018-07-20 | 201 | 0 | 0
2018-07-21 | 202 | 1317 | MM | clr | clr | clr | clr | rm hf | rm hf | * | * | clr
2018-07-22 | 203 | 0 | 0
2018-07-23 | 204 | 0 | 0
2018-07-24 | 205 | 0 | 0
2018-07-25 | 206 | 2335 | MM | rm lf | rm lf | clr | rm lf | rm lf | rm hf | * | * | clr
2018-07-26 | 207 | 0 | 0
2018-07-27 | 208 | 204 | MM | clr | clr | clr | clr | clr | clr | - | * | clr | SRI is working on changing IP Addresses on station as they align the satellite for our internet, as a result, the ESRL webpage has not updated since July 25th at 1400z
2018-07-28 | 209 | 0 | 0
2018-07-29 | 210 | 258 | MM | rm hf | rm hf | clr | rm lf | rm hf | rm hf | - | * | - | SRI is working on changing IP Addresses on station as they align the satellite for our internet, as a result, the ESRL webpage has not updated since July 25th at 1400z
2018-07-30 | 211 | 1234 | MM | clr | clr | clr | clr | clr | clr | - | * | clr | IP address change needed to update the ESRL MET.all page
2018-07-31 | 212 | 0 | 0
2018-08-01 | 213 | 0 | 0
2018-08-02 | 214 | 2147 | MM | clr | clr | clr | clr | clr | clr | * | * | clr | Adjusted height of humidity and temperature sensor to 2.2m, climbed tower to measure height of sonic (25ft)
2018-08-03 | 215 | 1810 | MM | clr | clr | clr | clr | clr | clr | - | * | clr
2018-08-04 | 216 | 1430 | MM | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr | Power to TAWO tower shutdown during planned extended power outage 1040-1415z
2018-08-05 | 217 | 0 | 0
2018-08-06 | 218 | 0 | 0
2018-08-07 | 219 | 2037 | MM | rm hf | rm hf | clr | clr | rm hf | rm hf | * | * | clr | Power to TAWO tower shutdown during planned extended power outage 1028-1320z
2018-08-08 | 220 | 0 | 0
2018-08-09 | 221 | 1524 | MM | rm hf | rm hf | clr | rm lf | rm hf | rm hf | * | * | clr
2018-08-10 | 222 | 0 | 0
2018-08-11 | 223 | 0 | 0
2018-08-12 | 224 | 0 | 0
2018-08-13 | 225 | 2200 | MM | rm hf | rm hf | clr | rm lf | rm hf | rm hf | * | * | lf
2018-08-14 | 226 | 0 | 0
2018-08-15 | 227 | 0 | 0
2018-08-16 | 228 | 0 | 0
2018-08-17 | 229 | 1614 | MM | rm hf | rm hf | clr | rm lf | rm hf | rm hf | * | * | clr
2018-08-18 | 230 | 0 | 0
2018-08-19 | 231 | 0 | 0
2018-08-20 | 232 | 0 | 0
2018-08-21 | 233 | 1746 | MM | rm lf | rm lf | clr | rm lf | rm hf | rm hf | * | * | clr
2018-08-22 | 234 | 0 | 0
2018-08-23 | 235 | 0 | 0
2018-08-24 | 236 | 0 | 0
2018-08-25 | 237 | 1930 | MM | rm hf | rm hf | clr | rm hf | rm hf | rm hf | * | * | clr
2018-08-26 | 238 | 1749 | MM | clr | clr | clr | clr | clr | changed filter | * | * | clr | MET tower was powered down from 1502-1540 while the 10m Sonic was replaced with a new Lufft sensor
2018-08-27 | 239 | 1251 | MM | rm lf | rm lf | clr | clr | rm lf | rm lf | * | * | clr
2018-08-28 | 240 | 0 | 0
2018-08-29 | 241 | 0 | 0
2018-08-30 | 242 | 0 | 0
2018-08-31 | 243 | 1433 | MM | rm lf | rm lf | clr | clr | rm lf | rm lf | * | * | clr
2018-09-01 | 244 | 0 | 0
2018-09-02 | 245 | 0 | 0
2018-09-03 | 246 | 0 | 0
2018-09-04 | 247 | 1613 | MM | rm lf | rm lf | clr | rm lf | rm hf | rm hf | * | * | clr
2018-09-05 | 248 | 1700 | MM | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-09-06 | 249 | 0 | 0
2018-09-07 | 250 | 0 | 0
2018-09-08 | 251 | 0 | 0
2018-09-09 | 252 | 1632 | MM | rm lf | rm lf | clr | clr | rm lf | rm lf | * | * | clr
2018-09-10 | 253 | 0 | 0
2018-09-11 | 254 | 0 | 0
2018-09-12 | 255 | 1414 | MM | rm lf | clr | clr | clr | rm lf | clr | * | * | clr | Raised MET Tower power cables from TAWO to the tower @1447
2018-09-13 | 256 | 0 | 0
2018-09-14 | 257 | 0 | 0
2018-09-15 | 258 | 1950 | MM | rm hf | rm hf | clr | rm hf | rm hf | rm hf | * | * | clr
2018-09-16 | 259 | 0 | 0
2018-09-17 | 260 | 0 | 0
2018-09-18 | 261 | 1445 | MM | rm hf | rm hf | clr | rm hf | rm hf | rm hf | * | * | clr
2018-09-19 | 262 | 1319 | MM | rm hf | rm hf | clr | rm lf | rm lf | rm hf | * | * | clr
2018-09-20 | 263 | 0 | 0
2018-09-21 | 264
2018-09-22 | 265 | 0 | 0
2018-09-23 | 266 | 1650 | MM | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr | Ethernet cable connection lost from 2037 on 2018-09-22 to 1644Z on 2018-09-23. Re-seated ethernet cable connection and power cycled the Met system at 1644Z and turned logger box heater back on for the winter season
2018-09-24 | 267 | 1925 | MM | rm lf | rm lf | clr | clr | rm lf | rm lf | * | * | clr
2018-09-25 | 268 | 0 | 0
2018-09-26 | 269 | 0 | 0
2018-09-27 | 270 | 0 | 0
2018-09-28 | 271 | 2000 | MM | clr | clr | clr | rm lf | rm lf | rm lf | * | * | clr
2018-09-29 | 272 | 0 | 0
2018-09-30 | 273 | 0 | 0
2018-10-01 | 274 | 1930 | MM | rm hf | rm hf | clr | rm lf | rm lf | rm lf | * | * | clr | Tower work 1730-1930z: mounted 80/20 bar above sonic for RM Young windbird
2018-10-02 | 275 | 1937 | MM | rm lf | rm lf | clr | clr | rm lf | rm lf | * | * | clr
2018-10-03 | 276 | 0 | 0
2018-10-04 | 277 | 1653 | MM | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-10-05 | 278 | 1747 | MM | rm hf | rm lf | clr | lf | rm lf | rm lf | * | - | clr | Summit Met webpage down temporarily for data acquisition computer updates. Summit Wx display still operational.
2018-10-06 | 279 | 0 | 0
2018-10-07 | 280 | 0 | 0
2018-10-08 | 281 | 1843 | MM | rm lf | rm lf | clr | rm hf | rm hf | rm hf | * | * | clr
2018-10-09 | 282 | 0 | 0
2018-10-10 | 283 | 0 | 0
2018-10-11 | 284 | 0 | 0
2018-10-12 | 285 | 1124 | MM | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr | Summit Met webpage failed to update from 2018-10-10 1306z to 2018-10-11 at 1837z. After period of high wind, re-secured cables on tower at 1230z.
2018-10-13 | 286 | 0 | 0 | [9 empty] | Installed vane anemometer on met tower from 1400-1700z
2018-10-14 | 287 | 0 | 0
2018-10-15 | 288 | 1815 | MM | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-10-16 | 289 | 1750 | MM/TN
2018-10-17 | 290 | 0 | 0
2018-10-18 | 291 | 1755 | MM/TN
2018-10-19 | 292 | 1848 | MM/TN | rm lf | rm fl | clr | rm lf | rm lf | rm lf | * | * | clr | Power off to the CR1000 data logger for RM Young wiring modification at the logger 1900-1922z.
2018-10-20 | 293 | 1845 | MM | rm hf | rm hf | clr | rm hf | rm hf | rm hf | * | * | clr
2018-10-21 | 294 | 1750 | TN | rm hf | rm hf | clr | rm hf | rm hf | rm hf | * | * | clr
2018-10-22 | 295 | 1616 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-10-23 | 296 | 1656 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-10-24 | 297 | 1710 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr | Power off to the CR1000 data logger for RM Young wiring modification at the logger 1730-1844Z.
2018-10-25 | 298 | 1711 | TN | rm hf | rm hf | clr | rm hf | rm hf | rm hf | * | * | clr
2018-10-26 | 299 | 1640 | TN | rm hf | rm hf | clr | rm hf | rm hf | rm hf | * | * | n/a | afternoon/evening storm prevented Green House check
2018-10-27 | 300 | 1536 | TN | rm lf | rm lf | clr | clr | clr | rm lf | * | * | clr
2018-10-28 | 301 | 1526 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-10-29 | 302 | 1658 | TN | rm lf | rm lf | clr | rm hf | rm hf | rm hg | * | * | clr
2018-10-30 | 303 | 1506 | TN | clr | clr | clr | clr | clr | clr | * | * | clr
2018-10-31 | 304 | 1445 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-11-01 | 305 | 1742 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-11-02 | 306 | 1645 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-11-03 | 307 | 1813 | TN | rm hf | rm hf | clr | rm lf | rm hf | rm hf | * | * | clr
2018-11-04 | 308 | 1508 | TN | rm lf | rm lf | clr | rm lf | clr | clr | * | * | clr
2018-11-05 | 309 | 1247 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-11-06 | 310 | 1240 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-11-07 | 311 | 1255 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-11-08 | 312 | 1547 | TN | clr | clr | clr | rm lf | clr | clr | * | * | clr
2018-11-09 | 313 | 1325 | TN | rm lf | rm lf | clr | clr | clr | clr | * | * | clr
2018-11-10 | 314 | 1244 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-11-11 | 315 | 1616 | TN | rm hf | rm hf | clr | rm hf | rm hf | rm hf | * | * | clr
2018-11-12 | 316 | 1249 | TN | rm hf | rm hf | clr | rm hf | rm hf | rm hf | * | * | clr
2018-11-13 | 317 | 1320 | TN | rm hf | rm hf | clr | rm hf | rm hf | rm hf | * | * | clr
2018-11-14 | 318 | 1336 | TN | rm lf | rm lf | clr | rm hf | rm hf | rm hf | * | * | clr
2018-11-15 | 319 | 1355 | TN | rm hf | rm hf | clr | rm hf | rm hf | rm hf | * | * | clr
2018-11-16 | 320 | 1914 | TN | rm hf | rm hf | clr | rm hf | rm hf | rm hf | * | * | clr
2018-11-17 | 321 | 1228 | TN | clr | clr | clr | clr | clr | clr | * | * | clr
2018-11-18 | 322 | 1630 | TN | clr | clr | clr | clr | clr | clr | * | * | clr
2018-11-19 | 323 | 1234 | TN | rm hf | rm hf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-11-20 | 324 | 1258 | TN | clr | clr | clr | clr | clr | clr | * | * | clr
2018-11-21 | 325 | 1324 | TN | rm hf | rm hf | clr | rm lf | clr | clr | * | * | clr
2018-11-22 | 326 | 1520 | TN | clr | clr | clr | clr | clr | clr | * | * | clr
2018-11-23 | 327 | 1310 | TN | clr | clr | clr | rm lf | clr | clr | * | * | clr
2018-11-24 | 328 | 1250 | TN | rm lf | rm lf | clr | clr | clr | clr | * | * | clr
2018-11-25 | 329 | 1803 | TN | rm hf | rm hf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-11-26 | 330 | 1313 | TN | clr | clr | clr | clr | clr | clr | * | * | clr
2018-11-27 | 331 | 1530 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-11-28 | 332 | 1532 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-11-29 | 333 | 1555 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-11-30 | 334 | 1251 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-12-01 | 335 | 1330 | TN | rm lf | rm lf | clr | rm hf | rm hf | rm hf | * | * | clr
2018-12-02 | 336 | 1601 | TN | rm lf | rm lf | clr | rm hf | rm hf | rm hf | * | * | clr
2018-12-03 | 337 | 1338 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-12-04 | 338 | 1519 | TN | rm lf | rm lf | clr | rm hf | rm hf | rm hf | * | * | clr
2018-12-05 | 339 | 1322 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-12-06 | 340 | 1631 | TN | rm hf | rm hf | clr | rm hf | rm hf | rm hf | * | * | clr
2018-12-07 | 341 | 1308 | TN | rm lf | rm lf | clr | rm hf | rm hf | rm hf | * | * | clr | ESRL webpage data stopped logging at 2000 utc yesterday
2018-12-08 | 342 | 1258 | TN | rm lf | rm lf | clr | rm hf | rm hf | rm hf | * | * | clr
2018-12-09 | 343 | 1420 | TN | rm lf | rm lf | clr | rm hf | rm hf | rm hf | * | * | clr
2018-12-10 | 344 | 1314 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-12-11 | 345 | 1329 | TN | rm lf | rm lf | clr | rm hf | rm hf | rm hf | * | * | clr
2018-12-12 | 346 | 1324 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-12-13 | 347 | 1328 | TN | rm hf | rm hf | clr | rm hf | rm hf | rm hf | * | * | clr
2018-12-14 | 348 | 1315 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-12-15 | 349 | 1714 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-12-16 | 350 | 1443 | TN | clr | clr | clr | rm lf | rm lf | rm lf | * | * | clr
2018-12-17 | 351 | 1252 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-12-18 | 352 | 1314 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-12-19 | 353 | 1302 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-12-20 | 354 | 1258 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-12-21 | 355 | 1316 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-12-22 | 356 | 1301
2018-12-23 | 357 | 0 | 0
2018-12-24 | 358 | #REF! | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-12-25 | 359 | 1354 | TN | rm hf | rm hf | clr | rm lf | rm hf | rm hf | * | * | clr
2018-12-26 | 360 | 1322 | TN | rm hf | rm hf | clr | rm hf | rm hf | rm hf | * | * | clr | Climbed to the top of the tower to remove snow/ice accumulation from tower and instruments 1510 - 1605
2018-12-27 | 361 | 1305 | TN | rm hf | rm hf | clr | rm hf | rm hf | rm hf | * | * | clr
2018-12-28 | 362 | 0 | 0
2018-12-29 | 363 | 1327 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr
2018-12-30 | 364 | 0 | 0
2018-12-31 | 365 | 1254 | TN | rm lf | rm lf | clr | rm lf | rm lf | rm lf | * | * | clr

## O3 Analyzer
Date | Julian Day | Time (UTC) | Observer's Initials | Inlet | Panel Display (30-60 ppb) | Bench Temp (<45C) | Pressure (>620 hPA) | Flow rate (LPM) (>0.35) | Flow rate (LPM) (>.35) | Intensity Cell A (>45,000) | Intensity Cell B (>45,000) | Comments (Alarms, Calibration Levels...)
2018-01-01 | 1 | 1719 | HA | rm lt frost | 41.9 | 31.9 | 652.5 | 0.436 | 0.439 | 72 | 69 | calibration performed 1731- UTC; Zero: -.1, Level1: 50.3, Level2: 98.8
2018-01-02 | 2 | 0 | 0 | 0 | [7 empty] | Could not visit TAWO due to weather
2018-01-03 | 3 | 0 | 0 | 0 | [7 empty] | Could not visit TAWO due to weather
2018-01-04 | 4 | 0 | 0 | 0 | [7 empty] | Could not visit TAWO due to weather
2018-01-05 | 5 | 1454 | HA | rm hvy frost | 39.8 | 30.6 | 661.9 | 0.441 | 0.441 | 72 | 69
2018-01-06 | 6 | 0 | 0 | 0
2018-01-07 | 7 | 0 | 0 | 0
2018-01-08 | 8 | 1423 | HA | rm frost | 42.8 | 32.5 | 650.2 | 0.437 | 0.44 | 72 | 69 | calibration performed 1432z-1527. Zero: 15.9, Level 1: 15.9 , Level 2: 19.1. Instrument lost comms from 1527Z- Discovered unconnected 'Zero' tubing for calibration. Reconnected, and shutdown laptop and instrument from 1640-1645Z and restarted. Second calibration conducted from 1745-1830 , Zero: 0.0, Level 1:37.8, Level 2: 73.5. Comms not connected; PI alerted and troubleshooting in process.
2018-01-09 | 9 | 0 | 0 | 0 | [7 empty] | Comms down
2018-01-10 | 10 | 1531 | HA | 0 | 40.8 | 31.2 | 659.5 | 0.439 | 0.438 | 72 | 69 | Switched out comms cable, reestablished comms at 1530 UTC
2018-01-11 | 11 | 0 | 0 | 0
2018-01-12 | 12 | 1440 | HA | rm frost | 37.9 | 30.1 | 648.2 | 0.433 | 0.433 | 72 | 69
2018-01-13 | 13 | 0 | 0 | 0
2018-01-14 | 14 | 0 | 0 | 0
2018-01-15 | 15 | 1413 | HA | rm frost | 40.1 | 30.2 | 651 | 0.435 | 0.435 | 72 | 69 | could not perform calibration due to north winds
2018-01-16 | 16 | 0 | 0 | 0
2018-01-17 | 17 | 0 | 0 | 0
2018-01-18 | 18 | 1407 | HA | rm frost | 39.3 | 30.5 | 665 | 0.443 | 0.443 | 72 | 69 | calibration performed 1408-1455 UTC; Zero:-4, Level 1:35.7 , Level 2: 71.1
2018-01-19 | 19 | 0 | 0 | 0
2018-01-20 | 20 | 1323 | HA | rm frost | 43.3 | 31.3 | 663.8 | 0.443 | 0.445 | 72 | 69 | Unexpected power outage from 18:21-18:22Z and 19:00-19:04Z
2018-01-21 | 21 | 0 | 0 | 0
2018-01-22 | 22 | 1715 | HA | rm frost | 34.8 | 31.5 | 661.5 | 0.441 | 0.443 | 72 | 69 | calibration performed 1717-18:03 UTC; Zero: -.2, Level 1: 36.4 , Level 2: 71.5
2018-01-23 | 23 | 0 | 0 | 0
2018-01-24 | 24 | 1409 | HA | rm frost | 38.8 | 29.5 | 660.2 | 0.44 | 0.438 | 72 | 69
2018-01-25 | 25 | 0 | 0 | 0
2018-01-26 | 26 | 0 | 0 | 0
2018-01-27 | 27 | 0 | 0 | 0
2018-01-28 | 28 | 0 | 0 | 0
2018-01-29 | 29 | 1432 | HA | rm frost | 36 | 32.3 | 657.4 | 0.44 | 0.443 | 72 | 69 | calibration performed 1440-1527 UTC; zero: 0.0, level 1:37.3, Level 2: 72.7
2018-01-30 | 30 | 0 | 0 | 0
2018-01-31 | 31 | 1434 | HA | rm frost | 37.7 | 30.9 | 664.6 | 0.443 | 0.444 | 72 | 69
2018-02-01 | 32 | 0 | 0 | 0
2018-02-02 | 33 | 0 | 0 | 0
2018-02-03 | 34 | 1411 | HA | rm frost | 41.2 | 29.6 | 645.7 | 0.432 | 0.432 | 72 | 69
2018-02-04 | 35 | 0 | 0 | 0
2018-02-05 | 36 | 1424 | HA | rm frost | 40.1 | 31.1 | 641.4 | 0.431 | 0.432 | 72 | 69
2018-02-06 | 37 | 1428 | HA | 0 | 41 | 31.5 | 651 | 0.437 | 0.439 | 72 | 69 | calibration performed 1436-1523 UTC; Zero: 0.4, Level 1: 35.9, Level 2: 71.3
2018-02-07 | 38 | 1612 | HA | rm frost | 38.3 | 31.1 | 642.5 | 0.431 | 0.433 | 72 | 69
2018-02-08 | 39 | 0 | 0 | 0
2018-02-09 | 40 | 1314 | HA | rm frost | 39.2 | 30.6 | 639.3 | 0.43 | 0.43 | 72 | 69
2018-02-10 | 41 | 0 | 0 | 0
2018-02-11 | 42 | 0 | 0 | 0
2018-02-12 | 43 | 1458 | HA, SWD | rm frost | 38.4 | 29.8 | 649.4 | 0.435 | 0.433 | 72 | 69 | Calibration 15:43-16:31 UTC; Zero:-0.3,Level 1:35.0, Level 2: 69.5
2018-02-13 | 44 | 0 | 0 | 0
2018-02-14 | 45 | 1418 | SWD | rm frost | 39.8 | 30 | 651.7 | 0.435 | 0.435 | 71 | 68
2018-02-15 | 46 | 0 | 0 | 0
2018-02-16 | 47 | 0 | 0 | 0
2018-02-17 | 48 | 1353 | SWD | rm frost | 38.4 | 30.6 | 661.3 | 0.441 | 0.443 | 71 | 68
2018-02-18 | 49 | 0 | 0 | 0
2018-02-19 | 50 | 0 | 0 | 0
2018-02-20 | 51 | 0 | 0 | 0
2018-02-21 | 52 | 1420 | SWD | rm frost | 40.1 | 33.2 | 655.7 | 0.44 | 0.444 | 72 | 69 | Calibration 14:23-15:09Z; zero: -0.3, level 1: 37.5, level 2: 73.2
2018-02-22 | 53 | 0 | 0 | 0
2018-02-23 | 54 | 0 | 0 | 0
2018-02-24 | 55 | 0 | 0 | 0
2018-02-25 | 56 | 1503 | SWD | rm lf | 41.2 | 30.7 | 663 | 0.442 | 0.443 | 72 | 69
2018-02-26 | 57 | 1539 | SWD | rm lf | 39.3 | 34.5 | 665.7 | 0.446 | 0.452 | 72 | 69 | Calibration 15:41-16:30Z; zero: 0.0, level 1: 38.4 , level 2: 75.4
2018-02-27 | 58 | 1851 | SWD | rm lf | 45.5 | 34.1 | 675.8 | 0.45 | 0.458 | 72 | 69
2018-02-28 | 59 | 0 | 0 | 0
2018-03-01 | 60 | 0 | 0 | 0
2018-03-02 | 61 | 1657 | SWD | rm lf | 44.8 | 34.1 | 687.1 | 0.456 | 0.461 | 72 | 69
2018-03-03 | 62 | 0 | 0 | 0
2018-03-04 | 63 | 0 | 0 | 0
2018-03-05 | 64 | 1439 | SWD | rm hf | 42.8 | 33.9 | 667 | 0.445 | 0.451 | 72 | 69 | Calibration 14:44-15:31Z; zero: -0.3, level 1: 37.7 , level 2: 74.7
2018-03-06 | 65 | 0 | 0 | 0
2018-03-07 | 66 | 1517 | SWD | rm hf | 40.4 | 31.6 | 657.7 | 0.44 | 0.442 | 71 | 69
2018-03-08 | 67 | 0 | 0 | 0
2018-03-09 | 68 | 0 | 0 | 0
2018-03-10 | 69 | 1243 | SWD | rm lf | 37.2 | 31.5 | 664.7 | 0.444 | 0.442 | 71 | 69
2018-03-11 | 70 | 0 | 0 | 0
2018-03-12 | 71 | 1454 | SWD | rm lf | 40.9 | 32.9 | 655.5 | 0.44 | 0.445 | 71 | 69 | Calibration 14:55-15:46Z; zero: 0.0, level 1: 37.5, level 2: 72.6
2018-03-13 | 72 | 0 | 0 | 0
2018-03-14 | 73 | 1505 | SWD | rm hf | 42.1 | 31.8 | 654.5 | 0.439 | 0.441 | 71 | 69
2018-03-15 | 74 | 0 | 0 | 0
2018-03-16 | 75 | 1244 | SWD | rm lf | 45.3 | 32.6 | 658.7 | 0.441 | 0.446 | 71 | 69
2018-03-17 | 76 | 0 | 0 | 0
2018-03-18 | 77 | 0 | 0 | 0
2018-03-19 | 78 | 1459 | SWD | rm hf | 45.8 | 34.6 | 665.4 | 0.445 | 0.452 | 71 | 69 | Calibration 15:03-15:50Z; zero: 0.4, level 1: 38.3, level 2: 75.2
2018-03-20 | 79 | 0 | 0 | 0
2018-03-21 | 80 | 1320 | SWD | rm lf | 47 | 35.2 | 656.1 | 0.441 | 0.448 | 71 | 69
2018-03-22 | 81 | 0 | 0 | 0
2018-03-23 | 82 | 1300 | SWD/RN | rm lf | 41.7 | 34.7 | 649.3 | 0.436 | 0.444 | 71 | 69
2018-03-24 | 83 | 0 | 0 | 0
2018-03-25 | 84 | 1227 | SWD | rm lf | 36.9 | 32.1 | 659.9 | 0.442 | 0.444 | 71 | 68 | Calibration 12:28-13:17Z; zero: 0.0, level 1: 37.1, level 2: 73.0
2018-03-26 | 85 | 0 | 0 | 0
2018-03-27 | 86 | 0 | 0 | 0
2018-03-28 | 87 | 1503 | NB | rm lf | 42.5 | 33.9 | 663.3 | 0.443 | 0.45 | 71 | 68
2018-03-29 | 88 | 0 | 0 | 0
2018-03-30 | 89 | 1722 | NB | rm lf | 41.9 | 32.8 | 661 | 0.442 | 0.446 | 71 | 68
2018-03-31 | 90 | 0 | 0 | 0
2018-04-01 | 91 | 1157 | NB | rm hf | 36.6 | 35 | 662.7 | 0.444 | 0.452 | 71 | 69
2018-04-02 | 92 | 1306 | NB | rm lf | 37 | 33.8 | 655.5 | 0.439 | 0.445 | 71 | 68 | Replaced inlet filter 1240-1243z. Calibration 1250-1335z; zero: -0.4, level 1: 38.3, level 2: 74.5
2018-04-03 | 93 | 0 | 0 | 0
2018-04-04 | 94 | 0 | 0 | 0
2018-04-05 | 95 | 1203 | NB | rm lf | 32.5 | 31.3 | 650.1 | 0.435 | 0.437 | 71 | 68
2018-04-06 | 96 | 0 | 0 | 0
2018-04-07 | 97 | 1122 | NB | rm lf | 35.6 | 31.4 | 647.8 | 0.434 | 0.436 | 71 | 68
2018-04-08 | 98 | 0 | 0 | 0
2018-04-09 | 99 | 1135 | NB | rm lf | 42 | 32.3 | 656.7 | 0.44 | 0.447 | 71 | 68 | Calibration 1135-1224z. Zero: 0.0, Level 1: 39.0, Level 2: 74.8
2018-04-10 | 100 | 0 | 0 | 0
2018-04-11 | 101 | 0 | 0 | 0
2018-04-12 | 102 | 1130 | NB | rm lf | 43.9 | 34.2 | 660.9 | 0.443 | 0.449 | 71 | 68
2018-04-13 | 103 | 0 | 0 | 0
2018-04-14 | 104 | 1314 | NB | clr | 43.3 | 36.7 | 665.4 | 0.446 | 0.456 | 71 | 68
2018-04-15 | 105 | 0 | 0 | 0
2018-04-16 | 106 | 1127 | NB | rm lf | 53.1 | 36.7 | 667.5 | 0.447 | 0.457 | 71 | 68 | Calibration 1147-1233z. Zero: 0.0, Level 1: 39.7, Level 2: 77.1
2018-04-17 | 107 | 0 | 0 | 0
2018-04-18 | 108 | 1145 | NB | rm hf | 56.5 | 36.7 | 669.7 | 0.448 | 0.458 | 71 | 68
2018-04-19 | 109 | 0 | 0 | 0
2018-04-20 | 110 | 1141 | NB | rm lf | 53.6 | 35.5 | 670.6 | 0.448 | 0.458 | 71 | 68
2018-04-21 | 111 | 0 | 0 | 0
2018-04-22 | 112 | 1207 | NB | rm lf | 51.7 | 36.7 | 662.2 | 0.444 | 0.445 | 71 | 68
2018-04-23 | 113 | 1144 | NB | rm hf | 53.2 | 36.7 | 659 | 0.442 | 0.453 | 71 | 68 | Calibration 1146 - 1232z. Zero: 0.0, Level 1: 39.3, Level 2: 76.5
2018-04-24 | 114 | 0 | 0 | 0
2018-04-25 | 115 | 0 | 0 | 0
2018-04-26 | 116 | 1130 | NB | rm lf | 47.9 | 36 | 657.7 | 0.442 | 0.451 | 71 | 68
2018-04-27 | 117 | 1327 | NB/MB | rm lf | 39.9 | 36.2 | 664.1 | 0.445 | 0.454 | 70 | 68
2018-04-28 | 118 | 0 | 0 | 0
2018-04-29 | 119 | 0 | 0 | 0
2018-04-30 | 120 | 1228 | NB/MB | clr | 32.32 | 34.7 | 650.7 | 0.438 | 0.445 | 70 | 68 | Calibration 1230-1319z. Zero: 0.4, Level 1: 38.5, Level 2: 74.5
2018-05-01 | 121 | 1224 | MB | rm lf | 46.6 | 36.1 | 648.7 | 0.437 | 0.447 | 70 | 68
2018-05-02 | 122 | 0 | 0 | 0
2018-05-03 | 123 | 0 | 0 | 0
2018-05-04 | 124 | 1227 | MB | clr | 44.6 | 37 | 642.3 | 0.433 | 0.445 | 70 | 68
2018-05-05 | 125 | 1217 | MB | rm hf | 46.41 | 37 | 644.1 | 0.435 | 0.445 | 70 | 68
2018-05-06 | 126 | 0 | 0 | 0
2018-05-07 | 127 | 1224 | MB | clr | 42.22 | 36.6 | 650.9 | 0.439 | 0.449 | 70 | 68 | Calibration 1226-1313z. Zero: 0.0, Level 1: 38.8, Level 2: 76.1
2018-05-08 | 128 | 1247 | MB | rm lf | 41.48 | 36.7 | 652.3 | 0.438 | 0.449 | 70 | 68
2018-05-09 | 129 | 0 | 0 | 0
2018-05-10 | 130 | 1247 | MB | rm lf | 44.66 | 36.9 | 659.5 | 0.443 | 0.454 | 70 | 68
2018-05-11 | 131 | 0 | 0 | 0
2018-05-12 | 132 | 1721 | MB | clr | 44.2 | 36.5 | 665.8 | 0.446 | 0.456 | 70 | 68
2018-05-13 | 133 | 0 | 0 | 0
2018-05-14 | 134 | 1153 | MB | rm lf | 48.75 | 36.9 | 668.6 | 0.447 | 0.457 | 70 | 68 | Calibration 1154-1240z. Zero: 0.0, Level 1: 40.3, Level 2: 78.2
2018-05-15 | 135 | 1515 | MB | rm lf | 48.62 | 37.4 | 661.9 | 0.444 | 0.455 | 70 | 68
2018-05-16 | 136 | 0 | 0 | 0
2018-05-17 | 137 | 1205 | MB | rm lf | 46.39 | 37 | 662.2 | 0.444 | 0.455 | 70 | 68
2018-05-18 | 138 | 0 | 0 | 0
2018-05-19 | 139 | 1219 | MB | lf | 46.01 | 37 | 658.1 | 0.442 | 0.452 | 70 | 68
2018-05-20 | 140 | 0 | 0 | 0
2018-05-21 | 141 | 1658 | MB | rm hf | 40.72 | 37.5 | 661.5 | 0.444 | 0.456 | 70 | 68 | Calibration 1700-1747z. Zero: 0.0, Level 1: 40.1, Level 2: 78.4
2018-05-22 | 142 | 1859 | MB | rm lf | 36.97 | 37 | 663.7 | 0.445 | 0.456 | 70 | 68
2018-05-23 | 143 | 1243 | MB | rm lf | 50.09 | 36.7 | 661 | 0.444 | 0.454 | 70 | 68
2018-05-24 | 144 | 1226 | MB | rm lf | 35.91 | 37 | 663.1 | 0.445 | 0.456 | 70 | 68
2018-05-25 | 145 | 0 | 0 | 0
2018-05-26 | 146 | 0 | 0 | 0
2018-05-27 | 147 | 0 | 0 | 0
2018-05-28 | 148 | 1220 | MB | rm hf | 43.66 | 36.9 | 666.1 | 0.445 | 0.457 | 70 | 68 | Calibration 1221-1308z. Zero: -0.0, Level 1: 40.6, Level 2: 78.1
2018-05-29 | 149 | 1347 | MB | rm lf | 47.95 | 36.9 | 671.9 | 0.449 | 0.459 | 70 | 68
2018-05-30 | 150 | 1201 | MB | clr | 45.34 | 37.1 | 674.1 | 0.451 | 0.46 | 70 | 68
2018-05-31 | 151 | 0 | 0 | 0
2018-06-01 | 152 | 1129 | MB | clr | 40.74 | 37 | 668.1 | 0.448 | 0.457 | 70 | 68
2018-06-02 | 153 | 0 | 0 | 0
2018-06-03 | 154 | 1203 | MB | clr | 42.31 | 36.9 | 672.1 | 0.449 | 0.46 | 70 | 68
2018-06-04 | 155 | 0 | 0 | 0
2018-06-05 | 156 | 1302 | MB/MM | clr | 47.6 | 37.1 | 678.6 | 0.453 | 0.462 | 70 | 68 | Calibration 1309-1358z. Zero: -0.3, Level 1: 40.3, Level 2: 79.3
2018-06-06 | 157 | 1341 | MB/MM | clr | 44.8 | 37.5 | 678.3 | 0.452 | 0.463 | 70 | 68
2018-06-07 | 158 | 1209 | MB/MM | clr | 50.1 | 38.3 | 675.3 | 0.451 | 0.463 | 70 | 68
2018-06-08 | 159 | 1220 | MM | clr | 34.7 | 38.2 | 669.7 | 0.449 | 0.46 | 70 | 68
2018-06-09 | 160 | 1647 | MM | clr | 45.7 | 39.1 | 672.3 | 0.45 | 0.463 | 70 | 68
2018-06-10 | 161 | 0 | 0 | 0
2018-06-11 | 162 | 1659 | MB/MM | rm hf | 46.1 | 39 | 677.1 | 0.453 | 0.466 | 70 | 68 | Calibration 1701-1750z. Zero: -0.73, Level 1: 41.0, Level 2: 79.8
2018-06-12 | 163 | 1802 | MM | rm hf | 49.9 | 38.5 | 673.8 | 0.451 | 0.462 | 70 | 68
2018-06-13 | 164 | 0 | 0 | 0
2018-06-14 | 165 | 0 | 0 | 0
2018-06-15 | 166 | 1824 | MM | rm lf | 33.3 | 36.9 | 667.8 | 0.447 | 0.457 | 70 | 68 | O3 Analyzer was shut down during a planned power outage while the aerosol pump was off--1655z-1859z
2018-06-16 | 167 | 1746 | MM | clr | 36.6 | 37.4 | 666.3 | 0.447 | 0.457 | 70 | 68
2018-06-17 | 168 | 0 | 0 | 0
2018-06-18 | 169 | 0 | 0 | 0 | [7 empty] | O3 Analyzer was shut down while the aerosol pump was off during during two brief planned power outages from 0115z-0130z and 0645z-0650z
2018-06-19 | 170 | 1326 | MM | clr | 43.3 | 38 | 672.1 | 0.45 | 0.462 | 70 | 68 | Calibration 1321-1400z. Zero: -0.4, Level 1: 39.3, Level 2: 77.7
2018-06-20 | 171 | 2019 | MM | clr | 50 | 38.2 | 671.5 | 0.45 | 0.462 | 70 | 68
2018-06-21 | 172 | 0 | 0 | 0
2018-06-22 | 173 | 1358 | MM | clr | 40.4 | 37.5 | 665.4 | 0.446 | 0.457 | 70 | 68
2018-06-23 | 174 | 0 | 0 | 0
2018-06-24 | 175 | 0 | 0 | 0
2018-06-25 | 176 | 1802 | MM | rm lf | 39.9 | 38.8 | 663.9 | 0.446 | 0.459 | 70 | 68
2018-06-26 | 177 | 1849 | MM | clr | 36.7 | 38.9 | 669.4 | 0.448 | 0.461 | 70 | 68 | Calibration 1812-1855z. Zero: 0.7, Level 1: 40.8, Level 2: 78.7
2018-06-27 | 178 | 1649 | MM | rm hf | 41.3 | 37.7 | 675.7 | 0.452 | 0.462 | 70 | 68
2018-06-28 | 179 | 1558 | MM | clr | 41.7 | 37.4 | 679.7 | 0.453 | 0.463 | 70 | 68
2018-06-29 | 180 | 0 | 0 | 0
2018-06-30 | 181 | 2047 | MM | clr | 34.5 | 38.4 | 670.4 | 0.449 | 0.461 | 70 | 68
2018-07-01 | 182 | 0 | 0 | 0
2018-07-02 | 183 | 0 | 0 | 0
2018-07-03 | 184 | 1603 | MM | clr | 35.9 | 39.3 | 672.5 | 0.45 | 0.463 | 70 | 68 | Calibration 1546-1636z. Zero: -0.4, Level 1: 40.7, Level 2: 80.0
2018-07-04 | 185 | 0 | 0 | 0
2018-07-05 | 186 | 1424 | MM | rm lf | 42.7 | 37.7 | 669.8 | 0.449 | 0.459 | 70 | 68
2018-07-06 | 187 | 0 | 0 | 0
2018-07-07 | 188 | 0 | 0 | 0
2018-07-08 | 189 | 0 | 0 | 0
2018-07-09 | 190 | 0 | 0 | 0
2018-07-10 | 191 | 1223 | MM | clr | 45.6 | 38.3 | 670.2 | 0.449 | 0.461 | 70 | 68 | Calibration 1232-1317z. Zero: -0.0, Level 1: 40.0, Level 2: 78.3
2018-07-11 | 192 | 1811 | MM | clr | 41.4 | 37.5 | 673.5 | 0.451 | 0.461 | 70 | 68
2018-07-12 | 193 | 0 | 0 | 0
2018-07-13 | 194 | 0 | 0 | 0
2018-07-14 | 195 | 0 | 0 | 0
2018-07-15 | 196 | 0 | 0 | 0
2018-07-16 | 197 | 2209 | MM | clr | 37.9 | 37.5 | 667.8 | 0.448 | 0.458 | 70 | 68
2018-07-17 | 198 | 0 | 0 | 0
2018-07-18 | 199 | 1430 | MM | clr | 41.5 | 39 | 669.5 | 0.448 | 0.461 | 69 | 68 | Calibration 1317-1403z. Zero: -0.4, Level 1: 40.8, Level 2: 78.3
2018-07-19 | 200 | 0 | 0 | 0
2018-07-20 | 201 | 0 | 0 | 0
2018-07-21 | 202 | 1317 | MM | rm hf
2018-07-22 | 203 | 0 | 0 | 0
2018-07-23 | 204 | 0 | 0 | 0
2018-07-24 | 205 | 0 | 0 | 0
2018-07-25 | 206 | 2335 | MM | rm lf | 35.8 | 39.5 | 669.3 | 0.448 | 0.462 | 69 | 68 | Calibration 2255-2335z. Zero: -0.4, Level 1: 40.3, Level 2: 78.8
2018-07-26 | 207 | 0 | 0 | 0
2018-07-27 | 208 | 204 | MM | clr | 33.5 | 37.2 | 669.4 | 0.448 | 0.458 | 69 | 67
2018-07-28 | 209 | 0 | 0 | 0 | [7 empty] | TAWO planned power outage from 1218-1228z, TEI ozone instrument off from 1213-1229z
2018-07-29 | 210 | 258 | MM | rm hf | 34 | 37.3 | 679 | 0.453 | 0.463 | 69 | 67 | TEI instrument without flow during unplanned power outage from 0231-0246z
2018-07-30 | 211 | 1234 | MM | clr | 59.6 | 37.8 | 680.9 | 0.455 | 0.466 | 69 | 67
2018-07-31 | 212 | 0 | 0 | 0
2018-08-01 | 213 | 0 | 0 | 0
2018-08-02 | 214 | 2147 | MM | clr | 39.4 | 39.2 | 669.4 | 0.45 | 0.463 | 69 | 67 | Planned power outage from 1458-1646z. Calibration 1752-1852z. Zero: 1.0, Level 1: 40.1, Level 2: 77.9 Note: During the zero calibration, forgot to press RUN button, when realized mistake, pressed the RUN button and continued calibration
2018-08-03 | 215 | 1810 | MM | clr | 27.1 | 31.1 | 675.8 | 0.449 | 0.453 | 70 | 67 | Planned power outage from 1458-1740z. TEI, NEPH and CLAP and CPD computer powered off during the outage
2018-08-04 | 216 | 1430 | MM | rm lf | 33.2 | 23.3 | 680.1 | 0.453 | 0.455 | 69 | 67 | Planned power outage from 1030-1402z. TEI, NEPH and CLAP and CPD computer powered off during the outage.
2018-08-05 | 217 | 0 | 0 | 0
2018-08-06 | 218 | 0 | 0 | 0
2018-08-07 | 219 | 2037 | MM | rm hf | 40.5 | 37.6 | 679.5 | 0.454 | 0.464 | 69 | 67 | Planned power outage from 1028-1312z. TEI, NEPH and CLAP and CPD computer powered off during the outage. Calibration 2005-2050z. Zero: -0.3, Level 1: 39.2, Level 2: 77.7
2018-08-08 | 220 | 0 | 0 | 0
2018-08-09 | 221 | 1524 | MM | rm hf | 41.4 | 38.5 | 673.4 | 0.45 | 0.462 | 69 | 67 | Brief planned power outage during which the TEI was turned off from 1035-1038z
2018-08-10 | 222 | 0 | 0 | 0 | [7 empty] | Planned power outage from 1700-1702z, TEI and aerosol pumps off
2018-08-11 | 223 | 0 | 0 | 0
2018-08-12 | 224 | 0 | 0 | 0
2018-08-13 | 225 | 2200 | MM | rm hf | 43.8 | 37.8 | 671.8 | 0.45 | 0.461 | 69 | 70 | Calibration 2136-2224z. Zero: -0.4, Level 1: 40.4, Level 2: 77.5
2018-08-14 | 226 | 0 | 0 | 0
2018-08-15 | 227 | 0 | 0 | 0 | [7 empty] | Brief planned power outage during which the TEI was turned off from 1212-1228
2018-08-16 | 228 | 0 | 0 | 0
2018-08-17 | 229 | 1814 | MM | rm hf | 43.2 | 37.6 | 669.4 | 0.448 | 0.459 | 69 | 70
2018-08-18 | 230 | 0 | 0 | 0
2018-08-19 | 231 | 0 | 0 | 0
2018-08-20 | 232 | 0 | 0 | 0
2018-08-21 | 233 | 1746 | MM | rm hf | 41.3 | 38.1 | 668.9 | 0.449 | 0.46 | 69 | 67 | Calibration 1723-1811z. Zero: 0.3, Level 1: 39.8, Level 2: 78.4
2018-08-22 | 234 | 0 | 0 | 0
2018-08-23 | 235 | 0 | 0 | 0
2018-08-24 | 236 | 0 | 0 | 0
2018-08-25 | 237 | 1930 | MM | rm hf | 34 | 36.5 | 666.6 | 0.447 | 0.455 | 69 | 67
2018-08-26 | 238 | 1747 | MM | clr | 34.1 | 36.9 | 665.3 | 0.447 | 0.456 | 69 | 67
2018-08-27 | 239 | 1251 | MM | rm lf | 35.8 | 37.4 | 667.8 | 0.448 | 0.458 | 69 | 67 | Calibration 1210-1256z. Zero: -0.0, Level 1: 39.9, Level 2: 77.5. Two brief planned power outages during which the aerosol pump and the O3 Analyzer were turned off: 1050-1054 and 1257-1302.
2018-08-28 | 240 | 0 | 0 | 0
2018-08-29 | 241 | 0 | 0 | 0
2018-08-30 | 242 | 0 | 0 | 0
2018-08-31 | 243 | 1433 | MM | rm lf | 34.9 | 37.5 | 660.6 | 0.444 | 0.445 | 69 | 67
2018-09-01 | 244 | 0 | 0 | 0
2018-09-02 | 245 | 0 | 0 | 0
2018-09-03 | 246 | 0 | 0 | 0
2018-09-04 | 247 | 1613 | MM | rm hf | 41.4 | 37.7 | 672.3 | 0.451 | 0.461 | 69 | 67 | Calibration 1550-16z. Zero: 0.0, Level 1: 39.7, Level 2: 77.8.
2018-09-05 | 248 | 1700 | MM | rm lf | 42.2 | 36.8 | 677.1 | 0.452 | 0.462 | 69 | 70
2018-09-06 | 249 | 0 | 0 | 0
2018-09-07 | 250 | 0 | 0 | 0
2018-09-08 | 251 | 0 | 0 | 0
2018-09-09 | 252 | 1632 | MM | rm lf | 48.5 | 37.2 | 669.5 | 0.448 | 0.459 | 69 | 67
2018-09-10 | 253 | 1822 | MM | rm lf | 49.2 | 38.5 | 666.6 | 0.447 | 0.459 | 69 | 67 | Calibration 1701-1746z. Zero: 0.0, Level 1: 40.1, Level 2: 78.6.
2018-09-11 | 254 | 0 | 0 | 0
2018-09-12 | 255 | 1414 | MM | rm lf | 40.3 | 37.5 | 673.9 | 0.45 | 0.461 | 69 | 70
2018-09-13 | 256 | 0 | 0 | 0
2018-09-14 | 257 | 0 | 0 | 0
2018-09-15 | 258 | 1950 | MM | rm hf | 52.8 | 37 | 666.5 | 0.447 | 0.457 | 69 | 67
2018-09-16 | 259 | 0 | 0 | 0
2018-09-17 | 260 | 0 | 0 | 0
2018-09-18 | 261 | 1445 | MM | rm hf | 36.8 | 37.6 | 672.3 | 0.45 | 0.461 | 69 | 67 | Calibration 1413-1459z. Zero: 0.4, Level 1: 39.8, Level 2: 77.6.
2018-09-19 | 262 | 0 | 0 | rm lf
2018-09-20 | 263 | 0 | 0 | 0
2018-09-21 | 264 | 1900 | MM | rm lf | 40 | 36.7 | 667.3 | 0.447 | 0.456 | 69 | 67
2018-09-22 | 265 | 0 | 0 | 0
2018-09-23 | 266 | 0 | 0 | rm lf
2018-09-24 | 267 | 1925 | MM | rm lf | 38.9 | 37.6 | 657.1 | 0.442 | 0.453 | 69 | 67
2018-09-25 | 268 | 1809 | MM | rm lf | 39.5 | 37.2 | 657.1 | 0.442 | 0.453 | 69 | 67 | Calibration 1920-2008z. Zero: 0.4, Level 1: 39.3, Level 2: 76.8.
2018-09-26 | 269 | 0 | 0 | 0
2018-09-27 | 270 | 0 | 0 | 0
2018-09-28 | 271 | 1954 | MM | rm lf | 41.9 | 36.8 | 655 | 0.441 | 0.451 | 69 | 67
2018-09-29 | 272 | 0 | 0 | 0
2018-09-30 | 273 | 0 | 0 | 0
2018-10-01 | 274 | 0 | 0 | rm lf
2018-10-02 | 275 | 1937 | MM | rm lf | 34.2 | 37.1 | 667.7 | 0.448 | 0.458 | 69 | 67 | Calibration 1747-1832z. Zero: 0.4, Level 1: 39.4, Level 2: 77.1
2018-10-03 | 276 | 0 | 0 | 0
2018-10-04 | 277 | 0 | 0 | rm lf
2018-10-05 | 278 | 1747 | MM | rm lf | 42.5 | 35.8 | 660.2 | 0.444 | 0.452 | 69 | 67
2018-10-06 | 279 | 0 | 0 | 0
2018-10-07 | 280 | 0 | 0 | 0
2018-10-08 | 281 | 1843 | MM | rm hf | 40.1 | 36.8 | 653.9 | 0.44 | 0.45 | 69 | 67
2018-10-09 | 282 | 1819 | MM | rm hf | 44.2 | 37.3 | 649.3 | 0.437 | 0.449 | 69 | 67 | Calibration 1703-1755z. Zero: -1.1, Level 1: 38.7, Level 2: 75.2
2018-10-10 | 283 | 1658 | MM | rm lf | 34.7 | 36.3 | 650.6 | 0.438 | 0.447 | 69 | 67 | Replaced inlet filter 1631-1645z.
2018-10-11 | 284 | 0 | 0 | 0
2018-10-12 | 285 | 1124 | MM | rm lf | 37.4 | 33.1 | 657.5 | 0.441 | 0.446 | 69 | 66
2018-10-13 | 286 | 0 | 0 | 0
2018-10-14 | 287 | 0 | 0 | 0
2018-10-15 | 288 | 1800 | MM | rm lf | 36.5 | 37.1 | 647.3 | 0.437 | 0.448 | 69 | 67
2018-10-16 | 289 | 1750 | MM/TN | rm hf | 35.5 | 37.1 | 651.3 | 0.439 | 0.45 | 69 | 67 | Calibration 1703-1755z. Zero: 0.0, Level 1: 38.5, Level 2: 76.7
2018-10-17 | 290 | 0 | 0 | 0
2018-10-18 | 291 | 1755 | MM/TN | rm lf | 38.1 | 36.6 | 651.5 | 0.438 | 0.449 | 69 | 67
2018-10-19 | 292 | 1848 | MM/TN | rm lf | 36.3 | 36.2 | 648.7 | 0.437 | 0.448 | 69 | 67 | Planned power outage from 1538-1541, and again from 1602-1606, TEI surface ozone analyzer and aerosol pumps off
2018-10-20 | 293 | 0 | 0 | rm hf
2018-10-21 | 294 | 1630 | TN | rm hf | 31.5 | 34.3 | 652.5 | 0.439 | 0.445 | 68 | 66
2018-10-22 | 295 | 1616 | TN | rm lf | 32.4 | 34.1 | 652.1 | 0.438 | 0.445 | 68 | 66
2018-10-23 | 296 | 1656 | TN | rm lf | 31.9 | 35.8 | 661.3 | 0.443 | 0.452 | 69 | 66 | Calibration: 1742-1834z. Zero: -0.8, Level 1: 38.4, Level 2: 75.3
2018-10-24 | 297 | 1710 | TN | rm lf | 31.8 | 35.8 | 664.1 | 0.445 | 0.453 | 69 | 66
2018-10-25 | 298 | 1711 | TN | rm hf | 36.8 | 34.7 | 670.5 | 0.448 | 0.455 | 68 | 66
2018-10-26 | 299 | 1640 | TN | rm hf | 35.8 | 33.7 | 676.3 | 0.45 | 0.455 | 68 | 66
2018-10-27 | 300 | 1536 | TN | clr | 38.5 | 34.4 | 665.1 | 0.445 | 0.452 | 68 | 66
2018-10-28 | 301 | 1526 | TN | rm lf | 33.9 | 37 | 653.4 | 0.44 | 0.45 | 69 | 67
2018-10-29 | 302 | 1658 | TN | rm hf | 31.8 | 36.8 | 658.6 | 0.442 | 0.452 | 69 | 66
2018-10-30 | 303 | 1506 | TN | clr | 37.3 | 34.4 | 661.4 | 0.444 | 0.452 | 69 | 66 | Calibration: 1700-1753 . Zero: -0.8, Level 1: 37.8, Level 2: 75.3
2018-10-31 | 304 | 1445 | TN | rm lf | 34.2 | 33.2 | 663.4 | 0.444 | 0.449 | 68 | 66
2018-11-01 | 305 | 1742 | TN | rm lf | 31.4 | 35 | 664.1 | 0.445 | 0.453 | 68 | 66
2018-11-02 | 306 | 1645 | TN | rm lf | 34.1 | 35.8 | 662.9 | 0.445 | 0.453 | 68 | 66
2018-11-03 | 307 | 1813 | TN | rm hf | 35.9 | 35 | 655.8 | 0.44 | 0.448 | 68 | 66
2018-11-04 | 308 | 1508 | TN | clr | 35 | 33.8 | 656.3 | 0.44 | 0.446 | 68 | 66
2018-11-05 | 309 | 1247 | TN | rm lf | 33.5 | 34 | 657.4 | 0.441 | 0.447 | 68 | 66
2018-11-06 | 310 | 1240 | TN | rm lf | 36.2 | 34.8 | 667.4 | 0.447 | 0.454 | 68 | 66 | Calibration: 1245-1333 Zero: -0.4; Level 1: 38.6; Level 2: 75.1
2018-11-07 | 311 | 1255 | TN | rm lf | 36.6 | 33.8 | 668.1 | 0.447 | 0.452 | 68 | 66
2018-11-08 | 312 | 1547 | TN | clr | 33.8 | 32.8 | 665.3 | 0.444 | 0.448 | 68 | 66
2018-11-09 | 313 | 1325 | TN | clr | 34.8 | 32.3 | 659.8 | 0.441 | 0.445 | 68 | 66
2018-11-10 | 314 | 1244 | TN | rm lf | 36.5 | 37 | 660.9 | 0.444 | 0.454 | 69 | 66
2018-11-11 | 315 | 1616 | TN | rm hf | 37.7 | 36.1 | 665.5 | 0.446 | 0.455 | 68 | 66
2018-11-12 | 316 | 1233 | TN | rm hf | 36.7 | 36.9 | 664.5 | 0.446 | 0.456 | 68 | 66
2018-11-13 | 317 | 1320 | TN | rm hf | 38.8 | 37 | 663.5 | 0.446 | 0.455 | 69 | 66 | Calibration: 1339-1425. Zero: -0.7; Level 1: 39.3; Level 2: 76.2
2018-11-14 | 318 | 1336 | TN | rm hf | 38.6 | 37.2 | 659.1 | 0.442 | 0.453 | 68 | 66
2018-11-15 | 319 | 1355 | TN | rm hf | 35.6 | 36.8 | 657.9 | 0.443 | 0.452 | 68 | 66
2018-11-16 | 320 | 1914 | TN | rm hf | 34.9 | 36.1 | 664.6 | 0.446 | 0.455 | 68 | 66
2018-11-17 | 321 | 1228 | TN | clr | 38 | 33 | 663 | 0.443 | 0.448 | 68 | 66
2018-11-18 | 322 | 1630 | TN | clr | 38.4 | 37.1 | 660.1 | 0.443 | 0.454 | 68 | 66
2018-11-19 | 323 | 1234 | TN | rm lf | 38.3 | 37 | 675.7 | 0.451 | 0.461 | 68 | 66
2018-11-20 | 324 | 1258 | TN | clr | 36.5 | 33.1 | 681 | 0.453 | 0.457 | 68 | 66 | Calibration: 1304-1350. Zero: 0.0; Level 1: 38.1; Level 2: 74.7
2018-11-21 | 325 | 1324 | TN | clr | 37.7 | 32.7 | 668.7 | 0.447 | 0.45 | 68 | 66
2018-11-22 | 326 | 1520 | TN | clr | 34.9 | 34.8 | 667.9 | 0.447 | 0.454 | 68 | 66
2018-11-23 | 327 | 1310 | TN | clr | 36.2 | 36.3 | 674.7 | 0.451 | 0.46 | 68 | 66
2018-11-24 | 328 | 1250 | TN | clr | 41.1 | 37 | 677.7 | 0.453 | 0.462 | 68 | 66
2018-11-25 | 329 | 1803 | TN | rm lf | 38.2 | 36.8 | 678.2 | 0.453 | 0.462 | 68 | 66
2018-11-26 | 330 | 1313 | TN | clr | 36.1 | 36.7 | 669.1 | 0.449 | 0.457 | 68 | 66
2018-11-27 | 331 | 1530 | TN | rm lf | 38.1 | 34.3 | 665.3 | 0.445 | 0.452 | 68 | 66 | Calibration: 1720-1807. Zero: -0.3; Level 1: 37.5; Level 2: 74.0
2018-11-28 | 332 | 1532 | TN | rm lf | 42.2 | 30.9 | 659.4 | 0.44 | 0.441 | 68 | 66
2018-11-29 | 333 | 1555 | TN | rm lf | 37.3 | 36 | 652.3 | 0.44 | 0.448 | 68 | 66
2018-11-30 | 334 | 1251 | TN | rm lf | 35.5 | 37.1 | 654.6 | 0.441 | 0.451 | 68 | 66
2018-12-01 | 335 | 1330 | TN | rm hf | 33 | 33.8 | 654.7 | 0.439 | 0.445 | 68 | 66
2018-12-02 | 336 | 1601 | TN | rm hf | 38 | 38.5 | 653.7 | 0.44 | 0.453 | 69 | 66
2018-12-03 | 337 | 1338 | TN | rm lf | 35.7 | 39.3 | 653.5 | 0.441 | 0.454 | 68 | 66
2018-12-04 | 338 | 1519 | TN | rm hf | 41.4 | 41.9 | 657 | 0.443 | 0.459 | 69 | 67 | Calibration: 1542 - 1628 : Zero: -1.1; Level 1: 41.0; Level 2: 78.6
2018-12-05 | 339 | 1322 | TN | rm lf | 37.9 | 34 | 655.3 | 0.44 | 0.446 | 68 | 66
2018-12-06 | 340 | 1631 | TN | rm hf | 42 | 39.4 | 657.3 | 0.443 | 0.455 | 68 | 66
2018-12-07 | 341 | 1308 | TN | rm hf | 42.9 | 39.3 | 659.9 | 0.445 | 0.457 | 68 | 66
2018-12-08 | 342 | 1258 | TN | rm hf | 42.8 | 40.1 | 660.2 | 0.444 | 0.458 | 68 | 66
2018-12-09 | 343 | 1420 | TN | rm hf | 40.8 | 37.7 | 660.3 | 0.444 | 0.455 | 68 | 66
2018-12-10 | 344 | 1314 | TN | rm lf | 43.5 | 36.8 | 660.1 | 0.444 | 0.453 | 68 | 66
2018-12-11 | 345 | 1329 | TN | rm hf | 44.1 | 40.4 | 662.2 | 0.446 | 0.46 | 68 | 66 | Calibration: 1341 - 1417 UTC. Zero: -0.3; Level 1: 38.4; Level 2: 75.9;
2018-12-12 | 346 | 1324 | TN | rm lf | 36.6 | 34.1 | 653 | 0.438 | 0.445 | 68 | 66
2018-12-13 | 347 | 1328 | TN | rm hf | 36 | 36.3 | 659.1 | 0.442 | 0.452 | 68 | 66
2018-12-14 | 348 | 1315 | TN | rm lf | 38.6 | 36.7 | 653.5 | 0.439 | 0.45 | 68 | 66
2018-12-15 | 349 | 1714 | TN | rm lf | 41.5 | 36.4 | 655.1 | 0.441 | 0.45 | 68 | 66
2018-12-16 | 350 | 1443 | TN | rm lf | 37.5 | 36.5 | 658.3 | 0.442 | 0.452 | 68 | 66
2018-12-17 | 351 | 1252 | TN | rm lf | 36.4 | 36.1 | 658.3 | 0.441 | 0.451 | 68 | 66
2018-12-18 | 352 | 1314 | TN | rm lf | 28.2 | 36.6 | 664.7 | 0.446 | 0.456 | 68 | 66 | Calibration: 1312 - 1357 UTC . Zero: -0.3; Level 1: 38.6; Level 2: 76.9
2018-12-19 | 353 | 1302 | TN | rm lf | 34.1 | 37 | 661.3 | 0.444 | 0.455 | 68 | 66
2018-12-20 | 354 | 1258 | TN | rm lf | 36.7 | 38 | 661.4 | 0.445 | 0.456 | 68 | 66
2018-12-21 | 355 | 1316 | TN | rm lf | 41.6 | 37.5 | 659.9 | 0.443 | 0.454 | 68 | 66
2018-12-22 | 356 |  |  | 0
2018-12-23 | 357 | 0 | 0 | 0
2018-12-24 | 358 | 1301 | TN | rm lf | 39.6 | 37.3 | 650.2 | 0.438 | 0.45 | 68 | 66
2018-12-25 | 359 | 1354 | TN | rm hf | 36.4 | 38.2 | 654.7 | 0.44 | 0.454 | 68 | 66 | Calibration: 1408 - 1455 . Zero: 0.4; Level 1: 39.8; Level 2: 77.8
2018-12-26 | 360 | 1322 | TN | rm hf | 39.9 | 38 | 653.4 | 0.44 | 0.452 | 68 | 66
2018-12-27 | 361 | 1305 | TN | rm hf | 36.9 | 37.2 | 651.7 | 0.438 | 0.45 | 68 | 66 | Replaced inlet filter 1432-1445Z .
2018-12-28 | 362 | 0 | 0 | 0
2018-12-29 | 363 | 1327 | TN | rm lf | 36 | 37.1 | 655.9 | 0.441 | 0.451 | 68 | 66
2018-12-30 | 364 | 0 | 0 | 0
2018-12-31 | 365 | 1254 | TN | rm lf | 35.2 | 36.8 | 663.1 | 0.445 | 0.455 | 68 | 66
| 366 | 0 | 0 | 0

## DRUM Sampler
Date | Julian Day | Time (UTC) | Observer's Initials | Inlet Clear? | Pump running/temp OK? | (Weekly) Vacuum (in HG) | (Weekly) "Flow" (WC) | Comments
2018-01-01 | 1 | 1719 | HA | rm lt frost | * | 13 | 8
2018-01-02 | 2 | 0 | 0 |  |  |  |  | Could not visit TAWO due to weather
2018-01-03 | 3 | 0 | 0 |  |  |  |  | Could not visit TAWO due to weather
2018-01-04 | 4 | 0 | 0 |  |  |  |  | Could not visit TAWO due to weather
2018-01-05 | 5 | 1454 | HA | rm hvy frost | * | 13 | 8
2018-01-06 | 6 | 0 | 0
2018-01-07 | 7 | 0 | 0
2018-01-08 | 8 | 1423 | HA | rm frost | * | 13 | 8
2018-01-09 | 9 | 0 | 0
2018-01-10 | 10 | 0 | 0 | rm frost | * | 13 | 8
2018-01-11 | 11 | 0 | 0
2018-01-12 | 12 | 1440 | HA | rm frost | * | 13 | 8
2018-01-13 | 13 | 0 | 0
2018-01-14 | 14 | 0 | 0
2018-01-15 | 15 | 1413 | HA | rm frost | * | 13 | 8
2018-01-16 | 16 | 0 | 0
2018-01-17 | 17 | 0 | 0
2018-01-18 | 18 | 1407 | HA | rm frost | * | 13 | 8
2018-01-19 | 19 | 0 | 0
2018-01-20 | 20 | 1323 | HA | rm frost | * | 13 | 8 | Unexpected power outage from 18:21-18:22Z and 19:00-19:04Z
2018-01-21 | 21 | 0 | 0
2018-01-22 | 22 | 1758 | HA | rm frost | * | 13 | 8
2018-01-23 | 23 | 0 | 0
2018-01-24 | 24 | 1409 | HA | rm frost | * | 13 | 8
2018-01-25 | 25 | 0 | 0
2018-01-26 | 26 | 0 | 0
2018-01-27 | 27 | 0 | 0
2018-01-28 | 28 | 0 | 0
2018-01-29 | 29 | 1432 | HA | rm frost | * | 13 | 8
2018-01-30 | 30 | 0 | 0
2018-01-31 | 31 | 1434 | HA | rm frost | * | 13 | 8
2018-02-01 | 32 | 0 | 0
2018-02-02 | 33 | 0 | 0
2018-02-03 | 34 | 1411 | HA | rm frost | * | 12 | 8
2018-02-04 | 35 | 0 | 0
2018-02-05 | 36 | 1424 | HA | rm frost | * | 12 | 8
2018-02-06 | 37 | 1428 | HA
2018-02-07 | 38 | 1612 | HA | rm frost | * | 13 | 8
2018-02-08 | 39 | 0 | 0
2018-02-09 | 40 | 1314 | HA | rm frost | * | 12 | 8
2018-02-10 | 41 | 0 | 0
2018-02-11 | 42 | 0 | 0
2018-02-12 | 43 | 1458 | HA, SWD | rm hvy frost | * | 12.5 | 8.5
2018-02-13 | 44 | 0 | 0
2018-02-14 | 45 | 1418 | SWD | rm frost | * | 12.5 | 8.5
2018-02-15 | 46 | 0 | 0
2018-02-16 | 47 | 0 | 0
2018-02-17 | 48 | 1351 | SWD | rm frost | * | 13 | 8.5
2018-02-18 | 49 | 0 | 0
2018-02-19 | 50 | 0 | 0
2018-02-20 | 51 | 0 | 0
2018-02-21 | 52 | 1428 | SWD | rm frost | * | 13 | 8.5
2018-02-22 | 53 | 0 | 0
2018-02-23 | 54 | 0 | 0
2018-02-24 | 55 | 0 | 0
2018-02-25 | 56 | 1503 | SWD | rm frost | * | 13.5 | 9
2018-02-26 | 57 | 1539 | SWD | rm frost | * | 13.5 | 9
2018-02-27 | 58 | 1851 | SWD | rm frost | * | 13.5 | 9.5
2018-02-28 | 59 | 0 | 0
2018-03-01 | 60 | 0 | 0
2018-03-02 | 61 | 1657 | SWD | rm frost | * | 14 | 9.5
2018-03-03 | 62 | 0 | 0
2018-03-04 | 63 | 0 | 0
2018-03-05 | 64 | 1439 | SWD | rm frost | * | 13 | 9.5
2018-03-06 | 65 | 0 | 0
2018-03-07 | 66 | 1517 | SWD | rm frost | * | 13 | 9.5
2018-03-08 | 67 | 0 | 0
2018-03-09 | 68 | 0 | 0
2018-03-10 | 69 | 1243 | SWD | rm frost | * | 13 | 9.5
2018-03-11 | 70 | 0 | 0
2018-03-12 | 71 | 1454 | SWD | rm frost | * | 13 | 10
2018-03-13 | 72 | 0 | 0
2018-03-14 | 73 | 1529 | SWD | rm frost | * | 12.5 | 9.5
2018-03-15 | 74 | 0 | 0
2018-03-16 | 75 | 1244 | SWD | rm frost | * | 13 | 10
2018-03-17 | 76 | 0 | 0
2018-03-18 | 77 | 0 | 0
2018-03-19 | 78 | 1459 | SWD | rm frost | * | 14 | 11
2018-03-20 | 79 | 0 | 0
2018-03-21 | 80 | 1320 | SWD | rm frost | * | 14 | 12
2018-03-22 | 81 | 0 | 0
2018-03-23 | 82 | 1300 | SWD/RN | rm frost | * | 13.5 | 13
2018-03-24 | 83 | 0 | 0
2018-03-25 | 84 | 1227 | SWD | rm frost | * | 14 | 13
2018-03-26 | 85 | 0 | 0
2018-03-27 | 86 | 0 | 0
2018-03-28 | 87 | 1503 | NB | rm frost | * | 14 | 13.5
2018-03-29 | 88 | 0 | 0
2018-03-30 | 89 | 1722 | NB | rm frost | * | 13.5 | 13
2018-03-31 | 90 | 0 | 0
2018-04-01 | 91 | 1157 | NB | rm frost | * | 13.9 | 13.5
2018-04-02 | 92 | 1306 | NB | rm frost | * | 13.8 | 14
2018-04-03 | 93 | 0 | 0
2018-04-04 | 94 | 0 | 0
2018-04-05 | 95 | 1203 | NB | rm frost | * | 13 | 13.5
2018-04-06 | 96 | 0 | 0
2018-04-07 | 97 | 1122 | NB | rm frost | * | 13 | 13.5
2018-04-08 | 98 | 0 | 0
2018-04-09 | 99 | 1135 | NB | clr | * | 14 | 14.5
2018-04-10 | 100 | 0 | 0
2018-04-11 | 101 | 0 | 0
2018-04-12 | 102 | 1130 | NB | rm frost | * | 14 | 14.5
2018-04-13 | 103 | 0 | 0
2018-04-14 | 104 | 1314 | NB | clr | * | 14.5 | 15
2018-04-15 | 105 | 0 | 0
2018-04-16 | 106 | 1127 | NB | rm frost | * | 14.5 | 14.5
2018-04-17 | 107 | 0 | 0
2018-04-18 | 108 | 1145 | NB | rm frost | * | 15 | 14
2018-04-19 | 109 | 0 | 0
2018-04-20 | 110 | 1141 | NB | rm frost | * | 14.5 | 13.5
2018-04-21 | 111 | 0 | 0
2018-04-22 | 112 | 1207 | NB | rm frost | * | 14.5 | 14
2018-04-23 | 113 | 1144 | NB | rm frost | * | 14.5 | 13.5
2018-04-24 | 114 | 0 | 0
2018-04-25 | 115 | 0 | 0
2018-04-26 | 116 | 1130 | NB | clr | * | 14.5 | 13.5
2018-04-27 | 117 | 1327 | NB/MB | rm frost | * | 14.5 | 13.5
2018-04-28 | 118 | 0 | 0
2018-04-29 | 119 | 0 | 0
2018-04-30 | 120 | 1228 | NB/MB | clr | * | 14 | 13.5
2018-05-01 | 121 | 1223 | MB | rm frost | * | 13.8 | 13.5
2018-05-02 | 122 | 0 | 0
2018-05-03 | 123 | 0 | 0
2018-05-04 | 124 | 1227 | MB | rm snow | * | 14 | 13
2018-05-05 | 125 | 1217 | MB | clr | * | 14 | 13
2018-05-06 | 126 | 0 | 0
2018-05-07 | 127 | 1224 | MB | rm snow | * | 14 | 13.2
2018-05-08 | 128 | 1247 | MB | clr | * | 14 | 13
2018-05-09 | 129 | 0 | 0
2018-05-10 | 130 | 1247 | MB | rm frost | * | 14.2 | 13.1
2018-05-11 | 131 | 0 | 0
2018-05-12 | 132 | 1721 | MB | rm snow | * | 14.3 | 13.3
2018-05-13 | 133 | 0 | 0
2018-05-14 | 134 | 1153 | MB | rm snow | * | 14.5 | 13.3
2018-05-15 | 135 | 1515 | MB | clr | * | 14.2 | 13.1
2018-05-16 | 136 | 0 | 0
2018-05-17 | 137 | 1205 | MB | rm frost | * | 14.3 | 13.2
2018-05-18 | 138 | 0 | 0
2018-05-19 | 139 | 1219 | MB | rm frost | * | 14.2 | 13.1
2018-05-20 | 140 | 0 | 0
2018-05-21 | 141 | 1658 | MB | clr | * | 14.2 | 13.2
2018-05-22 | 142 | 1859 | MB | clr | * | 14.3 | 13.3
2018-05-23 | 143 | 1243 | MB | clr | * | 14.2 | 13.3
2018-05-24 | 144 | 1226 | MB | rm frost | * | 14.3 | 13.3
2018-05-25 | 145 | 0 | 0
2018-05-26 | 146 | 0 | 0
2018-05-27 | 147 | 0 | 0
2018-05-28 | 148 | 1220 | MB | clr | * | 14.3 | 13.5
2018-05-29 | 149 | 1347 | MB | clr | * | 14.3 | 13.6
2018-05-30 | 150 | 1201 | MB | clr | * | 14.5 | 13.6
2018-05-31 | 151 | 0 | 0
2018-06-01 | 152 | 1129 | MB | clr | * | 14.3 | 13.6
2018-06-02 | 153 | 0 | 0
2018-06-03 | 154 | 1203 | MB | clr | * | 14.3 | 13.6
2018-06-04 | 155 | 0 | 0
2018-06-05 | 156 | 1302 | MB/MM | clr | * | 14.8 | 13.8
2018-06-06 | 157 | 1341 | MB/MM | clr | * | 14.8 | 14
2018-06-07 | 158 | 1209 | MB/MM | clr | * | 14.9 | 14
2018-06-08 | 159 | 1220 | MM | clr | * | 14.8 | 14
2018-06-09 | 160 | 1647 | MM | clr | * | 14.8 | 14
2018-06-10 | 161 | 0 | 0
2018-06-11 | 162 | 1659 | MB/MM | clr | * | 14.9 | 13.5
2018-06-12 | 163 | 1802 | MM | clr | * | 14 | 13.5
2018-06-13 | 164 | 0 | 0
2018-06-14 | 165 | 0 | 0
2018-06-15 | 166 | 1624 | MM | clr | * | 14.4 | 13.5
2018-06-16 | 167 | 1746 | MM | clr | * | 14.4 | 13.5
2018-06-17 | 168 | 0 | 0
2018-06-18 | 169 | 0 | 0
2018-06-19 | 170 | 1326 | MM | clr | * | 14.5 | 13.8
2018-06-20 | 171 | 2023 | MM | clr | * | 14.8 | 13.5
2018-06-21 | 172 | 0 | 0
2018-06-22 | 173 | 1358 | MM | clr | * | 14.3 | 13.8
2018-06-23 | 174 | 0 | 0
2018-06-24 | 175 | 0 | 0
2018-06-25 | 176 | 1802 | MM | clr | * | 14.5 | 13.5
2018-06-26 | 177 | 1849 | MM | clr | * | 14.5 | 13.8
2018-06-27 | 178 | 1648 | MM | clr | * | 14.5 | 14
2018-06-28 | 179 | 1558 | MM | clr | * | 14.5 | 13.8
2018-06-29 | 180 | 0 | 0
2018-06-30 | 181 | 2046 | MM | clr | * | 14.5 | 13.5
2018-07-01 | 182 | 0 | 0
2018-07-02 | 183 | 0 | 0
2018-07-03 | 184 | 1603 | MM | clr | * | 14.8 | 13.5
2018-07-04 | 185 | 0 | 0
2018-07-05 | 186 | 1424 | MM | clr | * | 14.5 | 13.5
2018-07-06 | 187 | 0 | 0
2018-07-07 | 188 | 0 | 0
2018-07-08 | 189 | 0 | 0
2018-07-09 | 190 | 0 | 0
2018-07-10 | 191 | 1217 | MM | clr | * | 14.5 | 14
2018-07-11 | 192 | 1811 | MM | clr | * | 14.5 | 14
2018-07-12 | 193 | 0 | 0
2018-07-13 | 194 | 0 | 0
2018-07-14 | 195 | 0 | 0
2018-07-15 | 196 | 0 | 0
2018-07-16 | 197 | 2209 | MM | clr | * | 14.5 | 14
2018-07-17 | 198 | 0 | 0
2018-07-18 | 199 | 1430 | MM | clr | * | 14.5 | 14
2018-07-19 | 200 | 0 | 0
2018-07-20 | 201 | 0 | 0
2018-07-21 | 202 | 1317 | MM | clr | * | 14.5 | 14
2018-07-22 | 203 | 0 | 0
2018-07-23 | 204 | 0 | 0
2018-07-24 | 205 | 0 | 0
2018-07-25 | 206 | 2335 | MM | clr | * | 14.5 | 14
2018-07-26 | 207 | 0 | 0
2018-07-27 | 208 | 204 | MM | clr | * | 14.5 | 14
2018-07-28 | 209 | 0 | 0
2018-07-29 | 210 | 258 | MM | clr | * | 14.5 | 14.5
2018-07-30 | 211 | 1234 | MM | clr | * | 15 | 14
2018-07-31 | 212 | 0 | 0
2018-08-01 | 213 | 0 | 0
2018-08-02 | 214 | 2147 | MM | clr | * | 14.8 | 14
2018-08-03 | 215 | 1810 | MM | clr | * | 14.5 | 14 | 1716-1740z DRUM sampler shutdown during power outage
2018-08-04 | 216 | 1430 | MM | clr | * | 14.6 | 14 | 1030-1402z DRUM sampler shutdown during power outage
2018-08-05 | 217 | 0 | 0
2018-08-06 | 218 | 0 | 0
2018-08-07 | 219 | 2037 | MM | clr | * | 14.8 | 14 | 1028-1312z DRUM sampler shutdown during power outage
2018-08-08 | 220 | 0 | 0
2018-08-09 | 221 | 1524 | MM | clr | * | 14.5 | 14
2018-08-10 | 222 | 0 | 0
2018-08-11 | 223 | 0 | 0
2018-08-12 | 224 | 0 | 0
2018-08-13 | 225 | 2200 | MM | clr | * | 14.5 | 14
2018-08-14 | 226 | 0 | 0
2018-08-15 | 227 | 0 | 0
2018-08-16 | 228 | 0 | 0
2018-08-17 | 229 | 1614 | MM | clr | * | 14.5 | 6.8 | Can't feel any rotations in pelican case
2018-08-18 | 230 | 0 | 0
2018-08-19 | 231 | 0 | 0
2018-08-20 | 232 | 0 | 0
2018-08-21 | 233 | 1746 | MM | clr | * | 14.5 | 6.8
2018-08-22 | 234 | 0 | 0
2018-08-23 | 235 | 0 | 0
2018-08-24 | 236 | 2110 | MM | clr | * | 14.8 | 6.5 | Turned DRUM sampler off at 2110z to replace it but didn't have all of the right parts so turned the sampler back on at 2148z
2018-08-25 | 237 | 1930 | MM | clr | * | 13.2 | - | Retro'd old DRUM sampler and pump. Last readings were recorded at 1752z, HG pressure - 14.8 and WC pressure - 6.5. The slotted position on the 1st disc is at ~30 degrees
2018-08-26 | 238 | 1747 | MM | clr | * | 13.8 | -
2018-08-27 | 239 | 1251 | MM | clr | * | 13.5 | -
2018-08-28 | 240 | 0 | 0
2018-08-29 | 241 | 0 | 0
2018-08-30 | 242 | 0 | 0
2018-08-31 | 243 | 1430 | MM | clr | * | 13.2 | -
2018-09-01 | 244 | 0 | 0
2018-09-02 | 245 | 0 | 0
2018-09-03 | 246 | 0 | 0
2018-09-04 | 247 | 1613 | MM | rm fr | * | 14 | - | The slotted position on the 1st disc is at ~40 degrees
2018-09-05 | 248 | 1700 | MM | rm lf | * | 13.9
2018-09-06 | 249 | 0 | 0
2018-09-07 | 250 | 0 | 0
2018-09-08 | 251 | 0 | 0
2018-09-09 | 252 | 1632 | MM | rm lf | * | 13.9 | -
2018-09-10 | 253 | 0 | 0
2018-09-11 | 254 | 0 | 0
2018-09-12 | 255 | 1414 | MM | rm lf | * | 13.9
2018-09-13 | 256 | 0 | 0
2018-09-14 | 257 | 0 | 0
2018-09-15 | 258 | 1950 | MM | rm lf | * | 13.9
2018-09-16 | 259 | 0 | 0
2018-09-17 | 260 | 0 | 0
2018-09-18 | 261 | 1400 | MM | rm hf | * | 14
2018-09-19 | 262 | 0 | 0
2018-09-20 | 263 | 0 | 0
2018-09-21 | 264 | 1900 | MM | rm lf | * | 14
2018-09-22 | 265 | 0 | 0
2018-09-23 | 266 | 0 | 0
2018-09-24 | 267 | 1925 | MM | clr | * | 13.6
2018-09-25 | 268 | 0 | 0
2018-09-26 | 269 | 0 | 0
2018-09-27 | 270 | 0 | 0
2018-09-28 | 271 | 2000 | MM | clr | * | 14
2018-09-29 | 272 | 0 | 0
2018-09-30 | 273 |  | 0
2018-10-01 | 274 | 1938 | MM | rm hf | * | 14
2018-10-02 | 275 | 1937 | MM | rm hf | * | 13.8
2018-10-03 | 276 | 0 | 0
2018-10-04 | 277 | 0 | 0
2018-10-05 | 278 | 1747 | MM | clr | * | 13.8
2018-10-06 | 279 | 0 | 0
2018-10-07 | 280 | 0 | 0
2018-10-08 | 281 | 1843 | MM | rm hf | * | 13.7
2018-10-09 | 282 | 0 | 0
2018-10-10 | 283 | 0 | 0
2018-10-11 | 284 | 0 | 0
2018-10-12 | 285 | 1232 | MM | rm lf | * | 13.2
2018-10-13 | 286 | 0 | 0
2018-10-14 | 287 | 0 | 0
2018-10-15 | 288 | 1815 | MM | rm lf | * | 13
2018-10-16 | 289 | 1810 | MM/TN
2018-10-17 | 290 | 0 | 0
2018-10-18 | 291 | 1836 | MM/TN | rm lf | * | 12.8
2018-10-19 | 292 | 1848 | MM/TN | rm lf | * | 12
2018-10-20 | 293 | 0 | 0
2018-10-21 | 294 | 1630 | TN | rm hf | * | 12
2018-10-22 | 295 | 1616 | TN | rm lf | * | 12.4
2018-10-23 | 296 | 1656 | TN | rm lf | * | 12.2
2018-10-24 | 297 | 1710 | TN | rm lf | * | 12.8
2018-10-25 | 298 | 1711 | TN | rm hf | * | 12.8
2018-10-26 | 299 | 1640 | TN | rm hf | * | 13
2018-10-27 | 300 | 1536 | TN | clr | * | 13
2018-10-28 | 301 | 1526 | TN | rm lf | * | 13
2018-10-29 | 302 | 1658 | TN | rm lf | * | 13
2018-10-30 | 303 | 1506 | TN | rm lf | * | 13
2018-10-31 | 304 | 1445 | TN | rm lf | * | 13
2018-11-01 | 305 | 1742 | TN | rm lf | * | 13
2018-11-02 | 306 | 1645 | TN | rm lf | * | 12.9
2018-11-03 | 307 | 1813 | TN | rm lf | * | 13
2018-11-04 | 308 | 1508 | TN | rm lf | * | 13
2018-11-05 | 309 | 1247 | TN | rm lf | * | 13
2018-11-06 | 310 | 1240 | TN | rm lf | * | 13
2018-11-07 | 311 | 1255 | TN | rm lf | * | 13.1
2018-11-08 | 312 | 1547 | TN | rm lf | * | 13.1
2018-11-09 | 313 | 1325 | TN | rm lf | * | 13
2018-11-10 | 314 | 1244 | TN | rm lf | * | 13
2018-11-11 | 315 | 1616 | TN | rm hf | * | 13
2018-11-12 | 316 | 1233 | TN | rm lf | * | 12.9
2018-11-13 | 317 | 1320 | TN | rm lf | * | 13.2
2018-11-14 | 318 | 1336 | TN | rm lf | * | 12.8
2018-11-15 | 319 | 1355 | TN | rm hf | * | 12.8
2018-11-16 | 320 | 1914 | TN | rm hf | * | 13
2018-11-17 | 321 | 1228 | TN | clr | * | 12.8
2018-11-18 | 322 | 1630 | TN | rm lf | * | 13
2018-11-19 | 323 | 1234 | TN | rm lf | * | 13.1
2018-11-20 | 324 | 1258 | TN | clr | * | 13.2
2018-11-21 | 325 | 1324 | TN | clr | * | 13.2
2018-11-22 | 326 | 1520 | TN | clr | * | 13.1
2018-11-23 | 327 | 1310 | TN | clr | * | 13.1
2018-11-24 | 328 | 1250 | TN | clr | * | 13.1
2018-11-25 | 329 | 1803 | TN | rm lf | * | 13.4
2018-11-26 | 330 | 1313 | TN | rm lf | * | 13.3
2018-11-27 | 331 | 1530 | TN | rm lf | * | 13
2018-11-28 | 332 | 1532 | TN | rm lf | * | 12.6
2018-11-29 | 333 | 1555 | TN | clr | * | 12.3
2018-11-30 | 334 | 1251 | TN | rm lf | * | 12.5
2018-12-01 | 335 | 1330 | TN | rm lf | * | 12.9
2018-12-02 | 336 | 1601 | TN | rm lf | * | 12.8
2018-12-03 | 337 | 1338 | TN | rm lf | * | 13.1
2018-12-04 | 338 | 1519 | TN | rm lf | * | 13
2018-12-05 | 339 | 1322 | TN | rm lf | * | 12.6
2018-12-06 | 340 | 1631 | TN | rm hf | * | 12.7
2018-12-07 | 341 | 1308 | TN | rm lf | * | 12.4
2018-12-08 | 342 | 1258 | TN | rm lf | * | 13.1
2018-12-09 | 343 | 1420 | TN | rm lf | * | 13
2018-12-10 | 344 | 1314 | TN | rm lf | * | 12.9
2018-12-11 | 345 | 1329 | TN | rm hf | * | 12.7
2018-12-12 | 346 | 1324 | TN | rm lf | * | 12.5
2018-12-13 | 347 | 1328 | TN | rm hf | * | 13.3
2018-12-14 | 348 | 1315 | TN | rm lf | * | 13
2018-12-15 | 349 | 1714 | TN | rm lf | * | 13
2018-12-16 | 350 | 1443 | TN | clr | * | 13
2018-12-17 | 351 | 1252 | TN | rm lf | * | 13.1
2018-12-18 | 352 | 1314 | TN | rm lf | * | 13
2018-12-19 | 353 | 1302 | TN | rm lf | * | 13
2018-12-20 | 354 | 1258 | TN | rm lf | * | 13.2
2018-12-21 | 355 | 1316 | TN | rm lf | * | 13.2
2018-12-22 | 356
2018-12-23 | 357 | 0 | 0
2018-12-24 | 358 | 1301 | TN | rm lf | * | 13
2018-12-25 | 359 | 1354 | TN | rm lf | * | 13.1
2018-12-26 | 360 | 1322 | TN | rm lf | * | 13
2018-12-27 | 361 | 1305 | TN | rm hf | * | 13
2018-12-28 | 362 | 0 | 0
2018-12-29 | 363 | 1327 | TN | rm lf | * | 12.8
2018-12-30 | 364 | 0 | 0
2018-12-31 | 365 | 1254 | TN | rm lf | * | 12.5
| 366 | 0 | 0