GEOSUMMIT

Greenland Environmental Observatory Summit Station

2023 Protocol

Year-Round Measurements and Observations

Greenland Environmental Observatory - Science Coordination Office

General Inquiries:

Science Coordination Office (SCO)

sco@geo-summit.org

Individual members of SCO:

Jack Dibb

University of New Hampshire, Durham

Morse Hall, 39 College Rd.

Durham, NH 03824-3525

(603) 862-3063 (office), (603) 862-1915 (fax)

jed@guero.sr.unh.edu and jack.dibb@unh.edu

Robert Hawley

Dartmouth College

122 Fairchild Hall

Hanover, NH 03755

(603) 646-1425

robert.l.hawley@dartmouth.edu

Zoe Courville

Cold Regions Research and Engineering Laboratory (CRREL)

72 Lyme Rd

Hanover, NH 03755

(603) 646-4425

zoe.courville@us.army.mil

Von Walden

Washington State University

Paccar Clean Tech Building

Pullman, WA 99164

(509) 335-5645 (office), (208)-301-4699 (cell)

v.walden@wsu.edu

Message to science technicians…

The success of previous year-round measurement campaigns at Summit Station has been demonstrated both logistically and scientifically. During the previous field seasons much was learned to help create a more well-defined framework for the success of Summit Station. The concept of developing the station as a year-round long-term observatory came out of the successes and failures of the previous seasons.

GEOSummit, the Greenland Environmental Observatory at Summit Station, has been established as a long-term observatory under the National Science Foundation’s Arctic Sciences Program. The observatory is run through the Science Coordination Office, which consists of a team of past and present scientists from the Summit Station research community. Through a contract with the NSF, Battelle ARO provides logistical support and station maintenance. The purpose of the observatory is to provide baseline atmospheric, meteorological, and chemical data to the general scientific community for use in modeling, field studies, and future planning, as well as to generate a record of baseline measurements for future comparison and evaluation of climate change.

The success of the observatory lies primarily in the hands of the science technicians in the field. Good data will result from good sampling. However, it is not meant to be a substitute for communication. Please do not hesitate to contact any of the principal investigators for each task or any individuals directly at the Science Coordination Office.

Finally, we appreciate your hard work and efforts in – what we all appreciate is at a minimum – a challenging environment. Hopefully the reward is in the unique experience, but nonetheless our thanks go out to you!

Table of Contents

ICECAPS – Polar Atmospheric Emitted Radiance Interferometer (PAERI) – MSF	7

ICECAPS – Micro-Pulse Lidar (MPL) – MSF	19

ICECAPS – Laser Ceilometer (Ceilometer) – MSF	23

ICECAPS – Total Sky Imager (TSI) – MSF	27

ICECAPS – Millimeter Cloud Radar (MMCR) – MSF	33

ICECAPS – Data Management (Dataman) – MSF	41

ICECAPS – Ice Particle Image Camera (IcePIC) – MSF	45

ICECAPS – Precipitation Occurrence Sensor System (POSS) – MSF	49

ICECAPS - Multi-Angle Snowflake Camera (MASC) – MSF	53

ICECAPS – SOnic Detection And Ranging (SODAR) – MSF	57

ICECAPS – IceCAM – MSF	61

ICECAPS – Microwave Radiometers (MWR) – MSF	63

ICECAPS-ACE – BIRAL – MSF	67

ICECAPS-ACE - Optical Particle Counter (OPC) – MSF	70

ICECAPS – Albedo Stand and Solar Tracker – AWO	73

ICECAPS-ACE – POPS – AWO	77

ICECAPS-ACE – SKYOPC – AWO	79

ICECAPS-ACE – Fluxtower – AWO	81

ICECAPS-ACE – SIMBA – AWO	86

ICECAPS-ACE – Enclosure – AWO	92

ICECAPS-ACE - HMP (T/RH sensors in aspirated shields) – AWO	98

ICECAPS-ACE - KT15 (Infrared radiation thermometer) – AWO	101

ICECAPS-ACE - LICOR (LI-7500 CO2/H2O Analyzer) – AWO	104

ICECAPS-ACE - Metek (3D sonic anemometer) – AWO	107

ICECAPS-ACE - SnD (Sonic distance sensor) – AWO	110

ICECAPS-ACE - Ventus (2D sonic anemometer) – AWO	113

ICECAPS-ACE – Water-Based Condensation Particle Counter (CPC) – AWO	117

ICECAPS – Radiosonde Campaign	120

ICECAPS – Significant Weather Observations	131

NOAA – Carbon Cycle Greenhouse Gases (CCGG) Flasks – AWO	133

NOAA – Halocarbons & Atmospheric Trace Species (HATS) Flasks – AWO	136

NOAA – Surface Ozone Monitoring – AWO	140

NOAA – Meteorology Suite – AWO	143

NOAA – Firn Air – AWO	146

NOAA – Nephelometer – AWO	149

NOAA – Continuous Light Absorption Photometer (CLAP) – AWO	155

GFZ Potsdam / Strollo – Seismometer – AWO	160

NASA – Bamboo Forest Accumulation	163

NASA – ICESat Accumulation	164

NASA – ICESat Transect	165

DTU – Magnetometer – AWO	184

Vieregg - RNO-G	185

Dartmouth / SCO – GIS data management	191

SCO – Clean Air / Science Sector Reporting	194

NOAA / Battelle-ARO – Weather Station Displays	196

Battelle-ARO – Elevation and Altimeter Reporting	198

Battelle-ARO – Weather Observations	199

Battelle ARO - Mobile Science Facility Protocol	204

CRREL - Station Accumulation/Drift Survey	208

# ICECAPS – Polar Atmospheric Emitted Radiance Interferometer (PAERI) – MSF

Primary contact: 	Von Walden, v.walden@wsu.edu, (509) 335-5645 (office), (208)-301-4699 (cell)

How often:		Daily: Regular checks

Weekly: None

Clean requirements:		None

Supplies needed:		Soft-bristled brush

Introduction:

The Polar Atmospheric Emitted Radiance Interferometer (PAERI) measures downwelling infrared radiation from 3.3 - 21 μm (480 - 3000 cm-1) at less than 10-minute temporal

resolution with a spectral resolution of one wavenumber. The spectral data gathered by the PAERI can be used for temperature/moisture profiling, cloud phase, cloud optical depth, emissivity, and effective particle size.

The PAERI comprises a Bomem interferometer, whose detector is cooled with the Stirling cooler, one ambient temperature (ABB) and one hot temperature (HBB) calibration blackbody, a motor-driven scene mirror that alternately directs atmospheric or blackbody radiation at the detector, and the control electronics and PC. The instrument penetrates the south wall of the MSF, with the ABB, HBB, scene mirror and motor in a protective enclosure outside the building, and the interferometer, Stirling cooler, and control systems inside the building. A fan blows air upwards from the bottom of the exterior enclosure to try to maintain the air near the scene mirror at ambient temperature and to help keep the mirror snow free. If snow or rime builds up on the enclosure, do NOT brush it off with the fan running; it will draw the snow up and obscure the optical path, resulting in LW/SW Responsivity faults.  Instead, remove the enclosure as instructed in “Servicing the Enclosure Fan” below, bring it inside, clean it, and then reinstall it. The top of the enclosure is at roof level and has a motor-driven hatch.  After testing during the first winter of 2010/2011, it was decided to leave the hatch open 100% of the time because this actually minimizes snow collecting inside the PAERI enclosure.

Data from the PAERI are transferred to the local and Boulder data archives daily by the Dataman archiving scripts running on the ICECAPS Linux PC. The PAERI PC will reboot itself once per day at 24Z; the instrument will remain powered up throughout this reboot.

Note: The status of the flags and other specifics can be viewed via the PAERI Quicklooks page, located at: https://labs.wsu.edu/lar-sensing/icecaps-experiment/

Procedure:

Daily Checks

- Outside the MSF, on the south wall, check the operation of the enclosure fan

  - If the fan is not running, verify that it is plugged into its extension cord, then check the fuse located between the fan and the extension cord

  - Clear any rime off of the fan blades using the brush on the bamboo pole. If more than a small amount is present, the enclosure should be brought inside for cleaning. This is to prevent liberated snow from being blown inside the PAERI and causing faults.

  - If the fan is excessively noisy or has seized up (i.e., will not spin and/or blows a fuse each time it is plugged in), see “Servicing the Enclosure Fan” below

- On the MSF roof, clear all snow, ice, and rime from the railings above and to either side of the PAERI hatch, being careful to direct snow away from the instrument. Clear snow from the deck around the PAERI hatch

- At the PAERI PC, verify that the PC time is synced to the MSF time server

- Review the spectra and time series to verify that they are reasonable for the conditions

  - Scroll backwards through the spectra for the day and verify that they are reasonable and free of errors

    - The “|<<” and “>>|” buttons will take you to the first or last records of the day, “<<” and “>>” will scroll through individual records, “<>” will allow you to select a specific record, and “!<<” will display the next record with an error flag.

  - The upper of the three time series graphs shows the brightness temperatures of the window (red), 3μm (green), and 10 μm (blue). The most variable of this graph should be the 10 μm brightness, which will vary with the degree of “cloudiness” that PAERI is viewing

  - The middle time series shows the interferometer detector temperature (red) and its cooler current (black). These should remain very stable; sudden changes are an immediate problem

  - The lower time series shows the hot black body temperature (red) and the interferometer second port temperature (yellow); these should also remain very stable. Sudden changes are an indication of a problem

- Verify that the Ingest script is running properly in a terminal window

  - Ingest should cycle through a continuous sequence of eight roughly 20-second atmospheric scans (“Sn” – “Zn”) followed by two 20-second blackbody scans (“Hn” and “An”)

  - The Local Mirror Monitor window should show the mirror at zero degrees during the eight atmospheric scans, at 60 degrees for the “H” blackbody scan, and at 120 degrees for the “A” blackbody scan

- Verify that the Local Mirror Monitor is responding consistently with Ingest

  - As of 2013, the “Enable”, “Safe” and “Clean” buttons on the Local Mirror Monitor can be clicked, but have no effect on the actual mirror operation

- Occasionally the time series graphs do not update, but everything else is running as normal.  If this is the case, do not restart the software – data are being logged and it will likely return to normal after the standard nightly reboot. To verify, open a Cygwin window, navigate to /cygdrive/c/ftp/AEyymmdd and make sure that the yymmddC1.RNC and yymmddC2.RNC files have been created (and have some bytes in them). If those two files are present, then everything is fine except for the plotting software.

- Verify that the daily data directories have been created and that files are being saved

  - A directory titled “AEyymmdd” should be located in C:\; where “yymmdd” is today’s date

  - Data files should be created roughly every minute in that directory

  - Additionally, a directory titled “AEyymmdd” should be located in C:\ftp

- Check all data flags for yellow, red, or blue flags. You can check this here: statusLights.png (640×480) (icecaps-paeri-quicklooks.s3.us-west-2.amazonaws.com)

  - Flag colors:

    - Green: Value is OK

    - Yellow: Value is outside of ideal, instrument operation may be impacted

    - Red: Value is too far out of tolerance; instrument is no longer operating properly

    - Blue: There is no data for this flag

  - If a flag has a stripe of another color on its left side, it indicates that that parameter has been in that other flag state at some point since 0000 UTC on the current day

  - To see the present value of the parameter, scroll the mouse over the flag. To see a brief description of the parameter, and the threshold values, double-click the flag

  - Important Flags, meanings, responses

    - ABB & HBB Max Temp Diff

      - These mean the blackbodies are experiencing large temperature gradients. They’re not a big issue, just an indication for Von to be careful with calibration data.

      - If these go yellow or red, and do not resolve in a few hours, contact Von.

    - HBB Temp, HBB Temp Stability

      - If either of these Hot Black Body flags go yellow or red, contact Von.

    - LW/SW Responsivity

      - This is an indicator of the responsivity of the detector in the long or shortwave bands

      - If these go yellow or red, it is typically an indicator that the scene mirror has accumulated snow or frost

      - This flag often clears itself; if it does not do so in a few hours, proceed to “Cleaning the Scene Mirror” below

      - Note that the LW & SW Responsivity flags often turn red briefly overnight on heavy rime days. Von is aware of this and because they are only red for a brief period, there is nothing to be done. If the flag is red for longer than an hour then let Von know or if the flag is red during the day, go check on the mirror from the roof to see if any rime on the mirror can be spotted.

    - Outside Air Temp, Air Temp Near BBs

      - These are the ambient temperatures outside the enclosure, and inside the enclosure near the blackbodies. Ideally, these temperatures should be the same

      - If the BB temp has drifted more than 3-4 K from the outside air temperature, it often indicates that the enclosure fan has failed or is excessively rimed

    - Mirror Motor Temp

      - This is the temperature of the motor that drives the scene mirror

      - This temperature will often go yellow or red when ambient temperatures are consistently < -50°C. While the motor may continue to operate at these temps, these temperatures will often lead to the motor’s rotary encoder malfunctioning and the spectra becoming erroneous (see Encoder Scene Confirm)

      - In consistently cold temperatures, the solution to this flag is to unplug and remove the enclosure fan. Leave the lower hatch uncovered and store the fan and tube inside the MSF

    - Interferometer 2nd Port Temp, Cooler Comp Temp, and Cooler Expander Temp.

      - These are temperatures associated with the performance of the Stirling cooler

      - If they are yellow or red, it is often associated with the MSF main room being too hot

      - These are serious; lower room temperature and contact Von

    - Encoder Scene Confirm, Max Sample Std Dev

      - Encoder Scene Confirm monitors the rotary encoder on the scene mirror motor, and is a measure of how reliably the controller knows the position of the mirror.

      - These will often go yellow or red when ambient temps are <-50C. In this case, see Mirror Motor Temp above.

      - If these flags appear when ambient temps are warmer, contact Von. A software reboot may be necessary.

    - Detector Temp

      - This is an indicator of the health of the Stirling cooler

      - This flag will often go yellow or red when the MSF temperature is high, meaning that the Stirling cooler is working hard. Cool the MSF.

      - Report to Von if red. No need to report if yellow.

    - Cooler Current

      - This is a better indicator of the health of the Stirling cooler (see Detector Temp).

      - Cool the MSF

      - If this flag goes yellow or red, contact Von immediately.

Mirror Cleaning Procedure

This procedure should be followed anytime that the scene mirror needs cleaning.  The scene mirror is located outside in the PAERI shelter.  This will probably need to be done after a storm with high winds or a lot of snowfall.

- If this is your first time cleaning the PAERI scene mirror, please contact Von Walden (v.walden@wsu.edu) to review this procedure with him. Otherwise, proceed to step 2.

- Perform the Normal Shutdown Procedure for the PAERI software (all steps prior to rebooting the computer).

- Do not forget to rename the date folders.

Open a cygwin terminal window on the PAERI computer. Then type the following commands:

cd /cygdrive/c/aeri_run

python Mirror.py program

python Mirror.py ready

python Mirror.py go_angle 0

Locate the PAERI Scene Mirror cleaning brush in the MSF vestibule.

Go onto the MSF roof and clean the mirror by reaching into the hatch and cleaning the mirror from above.

Reboot the PAERI computer to start normal data acquisition.

Mirror Safing Procedure

If you ever feel that the PAERI mirror is in danger of being damaged (for instance, due to an incoming severe storm), then it may be prudent to “safe” the PAERI mirror. The procedure stops normal data acquisition and sets the position of the PAERI mirror at the hot blackbody (HBB) position. This position will protect the mirror from snow and ice accumulation.

- Perform the Normal Shutdown Procedure for the PAERI software.

- Do not forget to rename the date folders.

Open a cygwin terminal window on the PAERI computer. Then type the following commands:

cd /cygdrive/c/aeri_run

python Mirror.py program

python Mirror.py ready

python Mirror.py go_angle 180

Once the danger to the mirror has passed, reboot the PAERI computer to start normal data acquisition.

PAERI Software Reboot

- In the Windows quick launch menu (lower left corner, to right of Start menu), select “STKILL”. This will run a script that stops Ingest, the AERI window, and the mirror monitor

- In the same menu, select “wdtdis.bat”; this will disable the watchdog timer and its reboot sequence

- Go to C:\ and locate the folder labelled “AEyymmdd”, where “yymmdd” is today’s date. Rename this folder “AEyymmdd_1”; this prevents overwriting of data when the PC is rebooted. If multiple reboots are required during a day, increment the “_1” upwards with every reboot.

- Go to C:\ftp and rename the “AEyymdd” folder in that directory as in the step above

- Reboot the PC using the standard Windows restart

- NOTE: no other PAERI component should be powered down during a reboot

- Upon reboot, the PAERI data acquisition will start up automatically. It will take at least one full measurement cycle for the data window to open and begin displaying valid data

PAERI Emergency Shutdown

- Perform all steps in “PAERI Software Reboot”, except that you will shut the PC down from within Windows, instead of restarting it

- Ideally, contact Von prior to powering down any other PAERI components besides the PC. He will advise you on any other precautions that you must take before powering down the instrument

  - The Stirling cooler contains a liquid nitrogen dewar that can be damaged during the warm-up that occurs after the instrument is powered down; Von may advise you to vent the dewar just after turning it off by loosening the 4 screws that secure the cooler’s “cold finger” to the detector dewar

- If there is immediate danger to personnel or property that requires shutting the instrument down before consulting Von, shut the power down to the instrument using the power strip mounted on the left back corner of the PAERI rack.

PAERI Power Failure Recovery

- If station and UPS power has failed completely, then recovered, contact Von immediately to determine the status of the Stirling cooler

- At a minimum, the PAERI PC will have booted itself back up. Because of the uncontrolled nature of the shutdown, the “AEyymmdd” folders will not have been renamed prior to shutdown and hence will create problems upon the reboot. Execute all steps in the “PAERI Software Reboot” to correct this folder issue

Servicing the Enclosure Fan

Wooden plate

Fuse assembly was moved inside 2018-12)

Assembly slides out this way

Sheet metal fan tube

- If the enclosure fan must be shut down for any more than 30 seconds, remove the entire fan tube assembly from the enclosure and bring it inside the MSF. Restarting a cold fan outdoors may damage the fan motor

- Power the enclosure fan down by unplugging it from the extension cord outside

- Remove the entire fan assembly by sliding it horizontally southwards, out of its slot. Take the assembly indoors

- To check the fuse, open up the tube of foam insulation. A 20mm fuse holder is inside the tube. Von asks that we use the lowest amp-rated fuse that won’t blow under normal operating conditions. While 250mA, 500mA, and 750mA fuses are available, in winter conditions, only the 750mA fuses seem to work without blowing.

Fuse assembly with 20mm fuse holder

- Replacing the fan:

  - Remove the drywall screws that secure the sheet metal tube to the wooden plate; remove the tube

  - Unplug the electrical pigtail from the fan

  - Remove the four bolts and nuts that secure the fan to the wooden plate

  - Be sure that the replacement fan is retrofitted with cold temperature bearings

  - Be sure that the replacement fan is oriented such that air will blow upwards, through the wooden plate

  - Installation is the reverse of removal

- Replacing the fan bearings:

  - Place the fan, label-side-down, on a table and remove three Phillips machine screws from the face of the fan; remove the fan

  - This will expose the rotor and part of the windings. The rotor is held on the shaft by a snap ring; remove this ring with the expanding snap ring pliers

  - Remove the washer, spring, and second washer below the snap ring

  - Draw the rotor off the shaft. If it is difficult to pull off, you can temporarily re-attached the fan to the rotor and use the fan to improve your grip

  - There will be two bearings; one each pressed into the bottom and top sides of the rotor; check them for stiffness or seizure

  - If a bearing needs to be replaced:

    - Drive the old bearing out using the aluminum drift punch and light taps from a hammer

    - Press the new bearing in gently using the bearing driver and a hammer until it contacts that bearing seat

    - Be sure that the replacement bearing is a cold temperature rated bearing (608ZZ, A1M5SML Y223)

Rotor (l), bearing driver (c), drift punch (r)

  - Installation is reverse of removal. Note that the components installed on the shaft above the rotor, from bottom to top, are: washer, spring, washer, retaining ring

- Re-assembly and re-installation of the fan is reverse of removal. Be sure re-start the fan as quickly as possible once it is outside, before the fan motor cools off too much

The PAERI Hatch

- The PAERI hatch is currently not used as it creates snow impaction during wind events. It remains in a fully open position at all times. The hatch mechanism, sensor and controller should be left undisturbed.

Display errors indicative of serious issues: Notify PI as soon as possible.

Retro:

None

Consumable Resupply Items:

None

# ICECAPS – Micro-Pulse Lidar (MPL) – MSF

Primary contact: 	Von Walden, v.walden@wsu.edu

Secondary contact:	Matt Shupe, matthew.shupe@noaa.gov

How often:		Daily: Perform regular checks of instrument operation

Monthly: Perform after-pulse calibration of instrument

Clean requirements:		None

Supplies needed:		Soft-bristled brush

Lidar telescope cover

Introduction:

The Sigma Space Micro-Pulse Lidar measures cloud height, thickness, and particle type utilizing backscatter from a vertically oriented 532 nm laser with a polarized beam. The instrument measures backscatter in both the co- and cross-polarized directions to help distinguish ice and liquid water targets. The instrument is located entirely within the MSF, with the telescope aimed out a window on the south side of the MSF roof. Data are transferred to and displayed by the SigmaMPL software running on a dedicated PC; data are transferred and archived daily by the ICECAPS Linux server.

SAFETY NOTE: The MPL utilizes a Class 2 laser and is specified by the manufacturer to be eye safe. Nonetheless, good sense dictates not to look straight into the beam and not to bring reflective tools or clothing into the path of the beam.

Procedure:

Daily Checks

- Verify that the time on the computer is properly sync’ed to the MSF time server

- In the “Raw Data” and “R2 Corrected” tabs of the SigmaMPL software, verify that the data are updating regularly and that the outputs match prevailing conditions, the MMCR and ceilometer

- In the “Hardware” tab, verify that the “Bin Resolution” is 15m and “Averaging Time” is 5 seconds. “Save Data” should be checked, and the current file name should appear as “DATA\YYYYMMDDHH00.mpl”, where “HH” is the current UTC hour

- Check for any error messages in the “Hardware Messages” box at the left side of the SigmaMPL window

  - The error messages ‘missing calibration file’ and VantageVue communications errors are normal and acceptable.

- Check the Laser Energy value in the “Profile Information” box at the left side of the SigmaMPL window.  The value should be around 5.0 uJ; note the value in the daily log

- Check the MPL Data folder (C:\Program Files\SigmaMPL\DATA) to ensure that files are updating hourly, and note in daily log. File sizes (except current file) should be about 11 MB. After the beginning of each new day, the ICECAPS Linux archive system pulls a copy of the previous day's files and then moves them into the C:\Program Files\SigmaMPL\DATA\Archive directory as a local backup archive.

- Check that the window is clean and free of condensation or frost, that the interior window fan is running, and note in daily log.  Clean exterior of window if needed with a soft bristled brush

  - NOTE: To prevent damage to the instrument from reflections, never pass a reflective tool or clothing in the path of the beam

- Check that the laser current (Ia) on the laser power supply is roughly 0.98A. The SHUTTER, Q-SW, and LDD green lights should be on. On the right side of the laser telescope, the red POWER light and green ACQ ACTIVE lights should be on. The green SYNC/USB light should be flashing.  Green light should be emanating from the top of the telescope; this can be visible outside when it is dark, and can often be seen inside as a green spot on the window.

- Record all findings in the log file (C:\Program Files\SigmaMPL\MPL log.txt)

Monthly Afterpulse Calibration

NOTE – It is best not to start the Afterpulse calibration within the 20 minutes before the hour (i.e. between xx:40 – xx:59) as the file you will create will change over at the top of the hour.

- Click Stop Collecting in the Hardware window of SigmaMPL, then close the software down. This will allow for a new file to be written

- Carefully thread the lid on top of the lidar telescope to shut out all outside light.

- Start up SigmaMPL

- Go to: File à Realtime Hardware Control

- Set the resolution to 30m and the time averaging to 30s. To make the system accept this information, you may need to click Stop Collection.

- Check ‘Save Data’.

- Click ‘Start Collection’.

- Let the system collect data with the lid on for 8 minutes. PI’s suggestion is 8 minutes.  In the MPL log notebook, note:

  - Time and date of this afterpulse calibration

  - Filename to which the system is saving data

  - Co Pol Bkgd Ave and Co Pol Bkgd Std values, from the Profile Information box to the left side of SigmaMPL

- After 5-10 minutes, hit Stop Collection.

- Ensure that the proper .mpl file has been saved to the C:\Program Files\SigmaMPL\DATA directory

- Repeat all above steps, this time setting the resolution to 15m and the time averaging to 30s. Un-check and re-check the ‘Save Data’ box to create a new filename

- Remove the telescope lid carefully when the second 8 min period is complete.

- Close SigmaMPL, then restart it

- Select File > Real Time Hardware Control

- Ensure that the collection parameters are at 15 m vertical resolution and 5 sec time averaging. If they were not at these from the start, you will need to click Stop Collection, then Start Collection to incorporate these parameters

- "Control Laser" should NOT be checked. Click "Save Data", and the instrument will resume measuring

Software Reboot

- Record the time and error type that caused SigmaMPL to freeze up, note it in the log file

- Shut the SigmaMPL software down, then restart the PC from within Windows. The power supply and telescope should remain powered up

- Once restarted, open the SigmaMPL software

- Select File > Real Time Hardware Control

- In the Hardware Tab, set:

  - Bin Resolution to 15m

  - Averaging Time to 5 seconds

  - Check Save Data, at which time the current file name should appear as “DATA\YYYYMMDDHH00.mpl”, where “HH” is the current UTC hour

  - Measurements will begin

- Record the time that measurements restarted in the log file

- Perform the regular daily checks

- Open the UltraVNC Server in the start menu. No window will open, but the MPL will now accept VNC connections.

System Startup

- Verify all cables are plugged into the power strip.

- Leave the lidar power supply plugged in for at least 5 minutes before powering up any part of the system; this allows the laser to warm up

- Turn on lidar power switch, just behind where the power cord connects to the bottom of the lidar telescope.  It can be difficult to get reach.

- Turn on the red main power switch on the front of the laser power supply.

- Turn the key on the front of the laser power supply to the ON position.  Q-SW LED will illuminate, but others will not

- Press the LDD button, the LDD LED should illuminate

- Verify that the laser is set for external trigger. On the power supply, press HOME, MENU, select Pulse Settings, press ENTER.  It should say Freq = Ext

- Verify that the current (Ia) is at roughly 1.0A, otherwise set it to that

  - Press MENU, select Diode Settings, press ENTER

  - If Ia is not near 1.0, then select Ia and adjust the current with the knob on the front of the power supply

  - Press ENTER to accept

- Press SHUTTER button on front of power supply; its LED should illuminate

- Check that laser light is coming from the top of the telescope.

  - Remove the laser cover if it is on

- Power up the Windows computer and start the SigmaMPL software

- Select File > Real Time Hardware Control

- In the Hardware window, verify that:

  - Averaging time is 5 seconds

  - Vertical resolution is 15 m

  - Communications are on COM1

  - "Control Laser" should NOT be checked, otherwise software will crash

  - “Save Data” is checked

  - If any of these values were initially different and needed to be changed, click the Stop Collection button, then Start Collection in order for the new values to be recognized

- Click Start Collection, if not already started.

- Perform the Daily Checks to verify that the instrument is working normally

- Open the Ultra VNC Server from the Start menu. No window will open, but the MPL will now accept VNC connections.

System Shutdown

- In the Hardware window on the SigmaMPL software, click Stop Collection

- Close the laser shutter by pressing the Shutter button on the laser power supply

- Turn of the laser current by pressing the LDD button on the laser power supply

- Exit the SigmaMPL program

- Turn off the key switch on the laser power supply

- Turn off the red power switch on the laser power supply

- Shut down the PC and monitor

- Turn off the lidar power switch on the bottom of the lidar

- Put the lid on the telescope if the system will be down for a long time

Retro:  None

Consumable Resupply Items:  None

# ICECAPS – Laser Ceilometer (Ceilometer) – MSF

Primary contact: 	Matt Shupe, matthew.shupe@noaa.gov

Secondary contact:	Victor Morris, victor.morris@pnl.gov

How often:		Daily: Daily checks

Weekly: None

Clean requirements:		None

Supplies needed:		Soft-bristled brush

Introduction:

The Vaisala CT25K laser ceilometer is a lidar that transmits a 16W, 905 nm beam and measures optical backscatter from cloud particles to determine cloud thickness and interpret cloud base heights.  The instrument is located at the east railing of the MSF roof and communicates data to a dedicated PC inside the MSF running Vaisala’s CT-View 2.0 software. The software stores hourly data files that are transferred by Dataman to the archives and to Boulder once per day.

Because the instrument is an aviation instrument, the altitude outputs are set for feet rather than meters.

The instrument is also used by Summit operations staff via VNC Viewer for flight weather observations, so it is important that the cloud detection screen be left open when not in the MSF.

SAFETY NOTE: The CT25K uses a Class 1M laser, meaning that it is eye-safe unless viewed through optics (i.e., binoculars, microscope, etc.). Nonetheless, good sense dictates not to look directly down at the lens when the instrument is energized.

Procedure:

Daily Checks

- Review the current CT-View data screen to ensure that the instrument is operating and that data are being taken continuously

- Check that current data display matches outside conditions as well as outputs from the MPL and MMCR

- A warning will always exist on this system for ‘low battery’ as well as ‘low charge’. Request a complete Status message (View > Status > Get Status Message); review the message for unusual readings or alarm conditions aside from this known warning.

  - Known errors: “Voltage high or low (W)” and “Receiver optical cross-talk compensation poor (V)” These usually occur during storm or high wind events. On the roof, check if there is snow or rime built up inside the Ceilometer opening. If the snow or rime persists for longer than a day, then might need to clear it out.

- Check the data directory (C:\CT-View2.0\data) to ensure that hourly files have been saved since midnight of the current day. All previous days’ files have been moved to the archive folder by a Dataman script.

  - Note:  As of January 2020, due to the age of the ceilometer software, hourly data file are being saved with the prefix for year 2010 (i.e.: 10MMDDhh.dat).  This issue is corrected by the by the ICECAPS dataman computer during the nightly data transfers such that you should see all files listed with the correct date in the ArchiveBoulder log (but not in the GetData section of the ICECAPS log).  If this correction is not occurring notify Matt Shupe immediately as this could result in data from 2010 getting overwritten.

- On the roof, listen to verify that the blower fan is operating

- If the blower is not running, and ambient temperatures are below -50˚C, check the complete Status message (see above). If the ceilometer’s outside temperature sensor is at < -50˚C, the “OUTSIDE” temperature reading will be out of range (“////C”) and the “BLOWER” and “OUTHEATER” will automatically be set to off. This is normal and will necessitate manual cleaning of the window. Open the north-side, upper door on the instrument; check that the window is snow- and frost-free. Clear accumulated snow with a soft-bristled brush. Consider clearing liquid water with kim wipes.

System Startup

- Power the ceilometer up by plugging in the power cable inside the MSF (UPS Circuit 4)

- Open the north-side, lower door of the ceilometer and verify that all three switches are ON (Figure 1). No key should be required to open this door; it is held fast by friction.

Figure 1: Ceilometer Lower Enclosure (switches circled)

- Power up the ceilometer PC

- If CT-View does not start automatically, start it from the desktop icon

- Click the “Open Communication Port” icon

- Start logging of data by setting menu Log > On

System Shutdown

- Close the CT-View software

- Power down the ceilometer PC, if desired

- Power down the ceilometer using the three switches inside the ceilometer, or by unplugging the power cable inside the MSF (UPS Circuit 4)

Retro:

None

Consumable Resupply Items:

None

# ICECAPS – Total Sky Imager (TSI) – MSF

Primary contact: 	Matt Shupe, matthew.shupe@noaa.gov

Secondary contact:	Von Walden, v.walden@wsu.edu

How often:		Daily: Perform regular checks of instrument operation

Early summer (~May 1): Install the TSI >-20

Late summer (~Oct 1): Remove the TSI

Clean requirements:		None

Supplies needed:		None

Introduction:

The Total Sky Imager (TSI) is a digital-imaging system used to monitor full-sky cloud conditions. Collected images are processed to determine the percentage of the sky with cloud cover.

The TSI manager software runs on the ceilometer computer. Data are saved locally to the C:\TSI directory. Images are initially saved in the “captured” folder and are subsequently moved to the “processed” folder. The “processed” folder is emptied nightly by dataman.

Procedure:

Daily Checks

- Within the TSI Manager software (on the ceilometer computer), open the date directory corresponding to the current date (YYYYMMDD), and scroll through cloud cover % to confirm this is realistic given current weather conditions.

- Check that the number of unprocessed images is < 2. This value appears to the right of the countdown to the next image capture. This number can rise when the instrument continues to capture images but the photos don’t get processed due to an issue in TSI Manager. If a backlog exists, perform the following action:

  - To clear the backlog, open the executable ‘tsifix_backup’ on the desktop; this is a Windows batch script. This will close and restart the software every 10 minutes to allow processing to proceed without hang-up. When the backlog has been cleared, close the tsifix_backup script; doing so avoids a window opening every 10 minutes that affects the ability of VNC users to view the ceilometer window.

- Make sure the shadowband is properly aligned with the sun. If it is not positioned correctly, overexposure to direct sunlight can permanently damage the CCD in the imager. White lines on an image indicate that the CCD is being overdriven and the shadowband needs to be realigned. See alignment notes below.

- Clean mirror or imager lens if visibly dirty (in this environment, they very rarely need to be cleaned) using soft, clean, lint-free cloth and nonabrasive glass cleaner (like Windex).

  - On heavy rime days, the mirror will often get covered and the image processing system will interpret the rime stuck on the mirror as cloud cover. Clear off as soon as possible by lightly sweeping the paintbrush over the mirror, being careful not to scratch the mirror.

  - On warm but also heavy rime days, the rime can get stuck onto the mirror. If this is the case, then use ethanol or alcohol solution to help remove the rime.

To inspect current image

- View the current image by clicking today’s folder (YYYYMMDD) and scrolling to most current image at the bottom of the list. NOTE: Do not click “Current” (under “Log” tab) in TSI Manager, as this will cause the system to stall.

Installation

- The TSI network and power cables remain installed in place through the winter. Confirm that the cables are present and locate the ends of the cables to prepare for this project.

- Raise the TSI to the MSF roof, using the main lifting handle and taking the utmost care to prevent damage to the imager arm and mirrored surface. The TSI can be lifted by the main handle and enclosure body; DO NOT pick up the TSI by the imager support arm!

- Place the TSI on the stand according to the orientation section below.

- Carefully feed the network cable into the grommet on the back of the TSI. , and plug the grommet with foam. Connect the network cable inside the instrument to Port 2, labeled “Computer.” Note that the lightning arrestor is not necessary, and should not be used.

- The grey cable coiled inside the enclosure is a for connection to a backup 12V battery. This is not used in the Summit installation and should remain securely coiled.

- Inside the MSF, plug in the TSI power cord to a UPS wall receptacle. This is the yellow extension cord that enters the building through the south wall pass-through near the MPL.

- Turn on the TSI using the power switch inside the TSI enclosure.

- The TSI should connect to the network.

- Once the TSI is powered and communicating on the network, check the settings (see images below) with TSI Manager software. It is likely that the TSI will need to be aligned with each seasonal installation. See instructions below.

- To initiate data collection, follow the startup procedures below.

Orientation

The instrument was relocated to the SW corner of the MSF roof in 2022. The imager arm is oriented to the north to align with the antenna support for the Vieregg project.  Contrary to the above screenshot, Imager Arm Orientation in the Image Capture tab of TSI Setup should be set to “North”.  All other software settings should remain as shown.

Aligning

To align the TSI, a serial-to-USB cable is used to communicate with the instrument. Issues were historically caused by using an off-brand serial-to-USB adapter. The Moxa adapter did work. Consult the manual section 2-5 for realignment procedures:

C:\Users\ARO-Summit-Technicia\Polar Field Services\PFS-ARC-Summit Techs - Documents\Instrument Manuals and Supporting Documents\ICECAPS\TSI

Uninstallation

The TSI should be uninstalled and stored inside MSF when temperatures are below -30°C for 6 hours or more per day. This is usually around mid-September through late May. Moving the TSI on or off the MSF roof requires two people.

- Removal: Turn off power switch (inside instrument), unplug power and network cables. Pull network cables out carefully through grommet on instrument.

- Protect the ends of the network and power cables with a plastic bag and zipties, and secure to the TSI stand for the winter season.

- Unscrew TSI from stand (4 bolts) and secure hardware in a bag. Attache hardware to TSI handle once placed into storage. Lower TSI instrument from roof. Take utmost care to protect the mirrored dome and fragile imager arm. As noted before, the TSI can be lifted by the main handle and enclosure body; DO NOT pick up the TSI by the imager support arm!

- Unplug the TSI extension cord inside the MSF.

- Unplug the TSI network cable inside the MSF.

- Store in the MSF with a cloth covering the mirror.

Shutdown:

- Stop image acquisition (press 'stop' button on TSI Manager) and close software.

- Perform shutdown procedures for ceilometer software following the protocol for that instrument.

- Shut down the ceilometer computer.

- Power off TSI by main power switch inside instrument, or by unplugging at the instrument stand (on roof).

Startup:

- Make sure TSI is plugged in (at instrument stand on roof) and power on via switch inside instrument.

- Turn on ceilometer computer.

- Start TSI Manger software ; make sure settings are correct (see above). Start image acquisition (press 'play' button in TSI Manager)

- Proceed with restoring operation of other science systems supported by this computer, such as the ceilometer.

Network error:

If an error saying “java.net.NoRouteToHostException: Operation timed out: no further information” this is indicative of a network issue.

The first fix to try is reseating the network cables to the TSI on the instrument side (open up the mirror of the TSI and look at the Ethernet Hub). The network cable should be connected to Port 2 which is labeled “Computer” in Sharpie. The Rx column should have a green light on both Port 1 and 2 and no lights in the PAR column. If there is a red light on the PAR side, then try reseating the cable.

On the left is what the Ethernet Hub looks like when this network error occurs, on the right what the Ethernet Hub should look like while the TSI is operating and collecting data.

If this doesn’t fix the error, on the ethernet switch look for the TSI switch (should be switch 21) and look for a blinking green light in the top row and it’s okay if the bottom light is off.

If none of these troubleshooting tricks work, try switching ports on the ethernet switch side and also use an Ethernet tester found in the Big House office.

Retro:

None

Consumable Resupply Items:

None

# ICECAPS – Millimeter Cloud Radar (MMCR) – MSF

Primary contact: 	Matt Shupe, matthew.shupe@noaa.gov

How often:		Daily: Daily checks

Quarterly: Replace desiccant

Clean requirements:	None

Supplies needed:	None

Introduction:

The MMCR is a NOAA 8mm (34 GHz, or Ka-band) vertically-pointing cloud radar that measures reflectivity and Doppler velocity of cloud particles. Measurements are made in 45m and 90m bins with a range of up to 13km, and data are used to determine cloud heights and thicknesses, particle size, cloud phase, and structure and movement within the cloud.

The MMCR antenna is located on the northeast corner of the MSF roof, and all control electronics are below it inside the MSF. The radar and its data are managed by two PC’s: mmcr-radar runs Vaisala’s LapXM software to analyze the returns from the receiver, while mmcr-dm runs a LabView script that handles the data from LapXM, controls the transmitter output, instrument self-calibrations, environmental and instrument monitoring, and controls the UPS that supports the radar. Data are transferred nightly from mmcr-dm to the ICECAPS Linux archive and to Boulder by the Dataman scripts.

Two manuals exist for the instrument: the original 1996 manual and an addendum describing the later upgrades to the control systems. Additionally, there are helpful troubleshooting guides that describe some particular behaviors of this radar, and possible solutions.

SAFETY NOTE: The radar emits short pulses of energy to the atmosphere through its antenna.  These pulses are weak and dispersed enough to not be dangerous to human health unless a person were to be in close proximity to the radar feed horn.  This horn is located at the center of the antenna underneath the radome.  Thus, significant efforts would be required for a human to put themselves in harm’s way.  Nonetheless, to minimize any exposure, it is recommended that people should not put any part of their body over the top of the antenna while it is in operation.  Moreover, direct reflections off of objectives just above the radar antenna could cause damage to the radar system itself.  The radar transmitter should be turned off if there will be any object (human or other) above the radar antenna.

Pulse controller (visible from ladder)

Radar UPS

Radian Interface

Radian Receiver/Modulator

Mmcr-dm and mmcr-radar PC’s

Monitor, keyboard, mouse, KVM switch

Up/Down Controller

TWTA (door on left)

Daily checks

- On the Pulse Controller:

  - Verify that all switches are DOWN except for the reset and pre-amp switches

  - Verify that no error lights are on

- On the TWTA (Travelling Wave Tube Amplifier), verify that:

  - The “Remote” light is ON

  - The “Fault” light is OFF

    - If the “Fault” light is ON, check inside the door, or the “TWTA Panel” window on mmcr-dm, for the fault. Report the fault to the PI

    - Attempt to reset the fault using the “Reset” button on the front of the TWTA

  - Check the Helix Current; it should be roughly 0.60 mA

    - Contact the PI if it is out of tolerance

- On the Radian Receiver/Modulator, verify that the Range Resolution is changing regularly from 50m to 90m

  - If the Resolution is not changing, this is typically an indicator that the radar needs to be rebooted (see below)

- On the Radian Interface, verify that “Tx” is blinking and that “Antenna Position” is indicating “0”.  Occasional flashing to “32” is normal.

  - If the antenna position is not “0”, this is typically an indicator that the radar needs to be rebooted (see below).

  - It is normal for “32” to flash about every 12 seconds.

- On the UPS, verify that the “AC Line” and “Ready” lights are green and that the UPS is not in alarm (i.e., beeping, or “Alarm” light on)

  - If the UPS shows an alarm, check the mmcr-dm Labview UPS Panel window for a description of the alarm, and be aware that the UPS may have already shut the radar down, or may begin shutting the radar down shortly

- On the LapXM screen of mmcr-radar, verify that:

  - The PC is sync’ed to the MSF time server

  - The reflectivity/Doppler velocity profiles are cycling through different range gates and modes every few seconds

  - The profiles are free of noise and consistent with prevailing conditions

  - The error warning indicator is green (“OK”).

    - “NetCDF” errors occur daily. As long as these errors can be cleared in LapXM, a reboot is not necessary.

  - The numeric display on the left side of the LapXM window may be wiped out by overlying windows, and will not return until the radar is rebooted. This is normal

- Switch to radar-dm using the KVM switch

  - On the main Labview screen, verify that there are no red alarm lights

  - Verify that the Transmitter Power is 180W +/- 10W

  - Check the 24-hour time series of transmitter output power; verify that it remains within tolerance except for a drop during the daily calibration at 1230Z

  - Open the “TWTA Panel” window, check for any red alarms in that window

  - Open the “UPS Panel” window, verify that:

    - The input and output voltages are roughly 120VAC

    - The UPS load is roughly 50%

    - The UPS has not logged any events since the last check

  - In C:\MMCR\data\ARM, verify that MMCRMom and MMCRSpecMom files, as well as a Health file, and an ELog file have been written in the last hour

    - Open the recent Health file and note any values that are out of tolerance (“Out” > 0), contact the PI if any values are out of tolerance

      - As of 2014, the “TWTA helix current” and“TWTA cathode voltage”regularly have a small percentage of values out of tolerance. These are known issues.

      - If pitch > 0.2 this indicates that the MSF is not level, contact PI.

    - If possible, open the 1300Z Health file and verify that the radar passed its 1230Z calibration cycle

    - Open the recent Elog file and verify that no errors are present. If they are, contact the PI

      - As of winter 2014, “netCDF” errors will appear in the Elog’s; they will typically appear continuously until the next time the system is rebooted. See the LapXM section above.

- On the roof of the MSF, check that the antenna is not covered with wet snow or ice

  - Dry snow and rime coverage does not necessitate cleaning

  - If the antenna needs cleaning, switch the transmitter off (see below) and carefully clean the antenna surface with a brush or soft mop

Switching the Transmitter Off/On-clarify with Matt when to do this (not necessary for reboot)

- Open the small door on the left side of the TWTA, and move the red switch to “Local”. This shuts off the transmitter and prevents the PC’s from remotely switching it back on

  - The LapXM display should go blank

- When ready to switch the transmitter back on:

  - Press the “Operate” button on the front of the TWTA, then move the red switch inside the TWTA door back to “Remote”

  - If the “Operate” button does not work, then move the red switch inside the TWTA door back to “Remote” and try  clicking the “Operate” button within the “TWTA Panel” window of mmcr-dm

  - If all else fails, reboot the radar (see below)

Rebooting the radar

- On the main Labview screen of mmcr-dm, click the large red button labelled “Shutdown and Reboot PC’s”. After confirmation, a countdown until the reboot will begin

- The reboot process is automatic and lasts roughly between 10 and 20 minutes, which includes the cool down and warm up periods for the TWTA

- There are a number of faults that can occur during the reboot process that require one to repeat the reboot; see the Troubleshooting Guides on the FTP for possible scenarios. It is probable that multiple (3, 4, 5….) reboots are needed.

Shutting down the radar (automatic)

- On the main Labview screen of mmcr-dm, click the large red button labelled “Shutdown and Reboot PC’s”. After confirmation, a countdown until the reboot will begin

- At a certain point in the shutdown process, after the radar components and PC’s have been completely powered down, the UPS will shut off, as indicated by its green ‘ready’ light turning off. At this point, switch the key on the UPS to OFF to prevent the radar from restarting. Note that, after a short time (~1min), the system will automatically start booting back up, so it is important to time the switch accordingly

Shutting down the radar (manual)

- Switch the TWTA to “Local” (see above)

- Switch the TWTA to standby using the “Standby” button on the front panel of the TWTA

- If the radar is going to be off for <4 hours and power to the TWTA will be maintained, leave the TWTA powered up. Otherwise, power the TWTA off

- In the main Labview screen of mmcr-dm, click “Stop Radar Monitor”

- When the arrow in the top left corner goes white, close the Labview screen

- Shut mmcr-dm down using the standard Windows shutdown

- Switch to mmcr-radar and select the “Services” shortcut in the lower left corner of Windows

- Right-click “Lapxm_main” and select “Stop”

- When the LapXM screen has closed, shut down the PC using the normal Windows shutdown

Cold starting the radar

- Switch the key on the UPS to ON, the automatic startup process will begin and last roughly 10-15 minutes

- Depending on circumstances under which the radar shut down, there are a number of faults that can occur during the boot process that require one to reboot; see the Troubleshooting guides for possible scenarios

- One issue to be aware of that happens if an automatic shutdown is started and the computers are shutdown manually or lose power before it can complete:  The computers may come up and then shortly shutdown again.  This is from shutdown commands that have not been cleared and are waiting to be executed.  The best thing is to shut down the full radar system and start over again. It will normally come up on the second try; do not be alarmed if this happens.

Changing the TWTA Dry Air Desiccant

- The desiccant cartridge is located on the back, or east, side of the TWTA

- Unthread the thumb screws from the base of the tube, remove the tube and the two cartridges

- Retrieve a pair of recharged cartridges from the grey Hardigg spare parts case and install them in the tube

  - The silica gel cartridge (Beach Filters EL65 “Silica Gel”), which is slightly blue, and may be stamped “Silica Gel” or “-SG”, is installed first, as the upper cartridge

  - The clay pellet cartridge (Beach Filters EL65), which is slightly brown, is installed second, as the lower cartridge

  - Verify that the air flow arrows on the side of the cartridge are from bottom to top

- Re-install the base

- Save the used cartridges in the spare parts case, noting on their bag how many times they have been used

  - Each cartridge can be re-charged up to three times, recharge them in a conventional oven at 200F for 2 hours

Assembled TWTA Desiccant Tube

Adjusting watts and pitch:

There is a cable attached via a zip tie that is sensitive to any movement. If the pitch is significantly over 1 and/or the watts are significantly below 180, you may need to adjust this cable. Left and right movement affect the pitch. Up and down movement affect the watts.

Common LabVIEW error: Notify the PI; you will be instructed to click continue button in the error window.

Another common LabVIEW error: notify Matt; you will be instructed to click Cancel. The processing should catch back up, let Matt know if it doesn’t and then you may have to restart.

Retro:

None

Consumable Resupply Items:

None

# ICECAPS – Data Management (Dataman) – MSF

Primary contact: 	Matt Shupe, matthew.shupe@noaa.gov

How often:		Daily: Regular checks

Weekly: GPS time server check

Clean requirements:		None

Supplies needed:		None

Introduction:

Dataman is the data management user on the ICECAPS Linux server.  This user runs a collection of data management scripts that are responsible for retrieving data from all of the ICECAPS instrument computers, performing certain quality control and processing tasks, archiving the data locally, and transferring the data to the NOAA servers in Boulder. The Dataman architecture and scripts are explained in great detail in the Dataman manual.

Procedure:

Daily Checks

- Verify that the time on the ICECAPS Linux server is consistent with the GPS time server

  - The system is set to sync automatically. Simply verify this is working.

- Verify that all subfolders in the “To Boulder” directory (/home/dataman/search/ToBoulder) are empty. If not, the overnight data transfer did not finish. Note this in the ICECAPS weekly report; contact Matt if the transfer process does not catch up the following day

- Verify that the Linux mounts to all of the instrument PCs, the RAID drive, and the USB disk; that they are active and are not approaching being full (i.e., are <95%)

  - To do so, in a terminal window, type “df –k” or “df –h”

  - The PAERI and LoggerPC mounts will often be >95%; this is not of concern unless it rises to 100%

  - Note that a pause of several seconds during the mount listing process usually indicates an issue locating one of the expected mounts.

  - If any mount is missing, see “Remounting Drives” below

- Run spot_check_archive.sh script in a terminal window

  - In the first tab:

    - Verify that the GetData script has moved all of the previous day’s data files from all instruments to the ICECAPS “work” folder (/home/dataman/search/work). Missing files could indicate a drive unmounting or an instrument ceasing to record data

    - Verify that LinkData has linked data from the archive to the ToBoulder folder

    - Verify that ArchiveData has successfully moved all of the data from /home/dataman/search/work to /data/archive

  - In the second tab, verify that ArchiveToBoulder has moved all files from the first tab to Boulder

    - Note the time of completion in the ICECAPS weekly report, as well as any abnormalities.

    - To confirm time of completion for previous-day’s-transfers execute the following in a Terminal Window:

      - tail -n 1 /data/archive/logs/Ar*YYYYMMDD*

      - where YYYYMMDD reflects day of interest

- Check the front of the ICECAPS Linux server for warning lights on any of the RAID or USB hard drives; contact Matt immediately in the event of a hard drive failure

Weekly Checks

- Check the status of the GPS time server on the MPL computer.

  - Log into the time server at 192.168.1.170.

    - Use Internet Explorer. Chrome does not run Java and will not display all information.

    - NOTE: Do not update JAVA.

  - Check the TIME tab; the Example time should be consistent with the current date and time as well as that displayed on the front of the server rack

    - For reference, one can check a website such as time.gov for an approximate check of the server’s accuracy; they may not match to the exact second due to network latency

  - Check the REFERENCE tab; the reference should be set to “GPS”, after waiting about 15 seconds, a prompt will come up asking to run or update a Java script.  Click “Run.” (Important: do not to update JAVA.)

    - The page should update and display “Reference Stable” as “TRUE”, and GPS should indicate “LOCKED”.

    - If the Reference Stable is “FALSE”, contact Matt immediately

  - If the GPS is “ACQUIRING”, click “Open Satellite Status” window to check how many satellites the server is acquiring (three are required)

  - If the GPS remains in “ACQUIRING”, check the condition of the satellite antenna cable and the antenna at the northeast corner of the MSF roof.

Powering Down the System

- If voluntarily powering down the system, ensure that you are doing so when no major programs are running.  Most of the onsite data archival processes occur from about 0100-0500 UTC time; avoid this time for powering down or rebooting the system

- Stop the POSS data acquisition by closing the terminal window running the posacq process. Note that POSS acquisition will need to be manually restarted when the computer system is repowered. Consult POSS protocol.

- To shut the system down you can either do so through the operating systems menu (“System” on the right-hand toolbar) or you can use the “shutdown” command at a terminal window command line.  There are various options to “shutdown” if needed; to see these, type “man shutdown” at a command prompt to see the manual page

- To simply reboot the system, use “shutdown –r” or reboot via the operating system menu.  See Important Note below.

Starting up the System

- Ensure that the computer system is powered (note there are two power supplies, both of which should be connected to line power). The computer is located in the MSF server rack to the right of the tech desktop and is labeled “Dataman.”

- Important Note:  As of February 2020, it appears that the CMOS battery on the motherboard is dead/dying and is not holding the correct boot settings.  As a result, the correct boot drive must be manually selected during startup by preforming the following steps:

  - On startup, when presented with the “Supermicro” logo, press the F11 key to be directed to the boot menu.

  - When the menu comes up, select the Samsung drive from the list and press enter.

  - The system should then boot up normally.

- Press the power button on the front of the system.  Standard boot up messages will scroll to the terminal.  These scroll quickly; if you see any irregularities or error messages, report these to Matt (in some cases it may be very hard to read the messages as things scroll quickly)

- After a minute or so the system should be booted up.  Log in as the “dataman” user (see password sheet for credentials)

- It is common for a couple of the mounts to not mount correctly.  Follow the procedure for “Re-Mounting the Drives”.  Ensure that the mount command comes back cleanly

- Since the POSS data are logged on this computer, open a terminal window and follow the procedures for starting the POSS data acquisition

Re-Mounting the Drives

- Type “sudo mount –a”.  This will ask for the dataman password. This command should come back cleanly (i.e., no error messages)

  - Errors can indicate missing mounts, so if you know a system is missing, turned off, or otherwise inactive, errors can be expected. However, any unexpected pauses, errors or omissions should be investigated.

- Type “exit” to remove the root permissions

- Check the status of the mounts using “df –k” or “df –h”

Swapping USB disks

- As the USB disk fills, or at Matt’s request, the disk will need to be swapped out and carried back to Boulder

- This procedure is best done in the morning (Summit time) on a weekday so that Matt can update the diskpath file and check the disk before the overnight archiving run begins. Verify that Matt is available that day before swapping the disk.

- Identify the full disk to be swapped; as of winter 2014 there was only one disk installed

- Power down the full disk, remove all cable connections to the disk

- Install the cables on the empty disk; power it up

- Using “df –k” or “df –h”, verify that ICECAPS has mounted the new disk. If needed, re-mount the drives as above

- Store the full disk in its Pelican case, carry it back to the U.S., and ship it to:

Matthew Shupe

R/PSD3

325 Broadway

Boulder, CO 80305

Retro:

See USB disk swap above.

Consumable Resupply Items:

None

# ICECAPS – Ice Particle Image Camera (IcePIC) – MSF

Primary contact: 	Von Walden, v.walden@wsu.edu

Secondary contact:	Chris Cox, Christopher.Cox-2@colorado.edu

How often:		As appropriate: Sampling and uploading of photos

Once per phase, at turnover: Clean upper and lower objective lens

Clean requirements:		None

Supplies needed:		Nikon D50 camera

Camera battery charger

Light source

Isopropyl alcohol

Glass microscope slides

Optical wipes

Introduction:

The primary goal of the Ice Particle Imaging Camera project is to document ice particle shapes and sizes from sources including precipitation at the surface, diamond dust, and blowing snow. It is unlikely that a comprehensive and statistically representative sample dataset of ice precipitation at Summit can be collected. The focus, therefore, will be on documenting different precipitation events. IcePIC data can be a valuable tool in setting the context of remote sensing measurements made continuously by ICECAPS.

For example, IcePIC data can be used as an interpretation tool in the analysis of particle orientation as seen by ICECAPS lidars during precipitation events. This will help span the gap between the lidar measurement and a given set of meteorological conditions. Also, a longer-term dataset of events considered to be of a similar type (such as diamond dust) can be analyzed with reference to radiance measurements by the ICECAPS Polar Atmospheric Emitted Radiance Interferometer (PAERI) to find the radiative effect of such events.

Procedure:

- Clean two microscope slides with alcohol and kim wipes. Dry thoroughly to avoid smudges on the glass.

- Store the cleaned microscope slides in the IcePIC shed, or leave the slides in a box outdoors to equilibrate to ambient temperatures prior to exposing the slides to any precipitation

- Store the camera and light in the computer rack in the MSF

- The IcePIC stage and camera mount remain in the IcePIC shed

- Sampling can occur in a number of different situations:

  - Short duration samples during active snowfall

  - Short duration samples during blowing/drifting snow events

  - Longer duration samples during diamond dust events

  - Overnight samples of accumulation

- When ready to sample, place two slides, at least one with a scale, face up on the T-plate

  - Place the T-plate in the sampling location such that one slide is upwind of the T-plate divider, and the other downwind

  - The sampling location should be far enough from the MSF to reduce the introduction of snow from the building into the sample. Place the T-plate on the portable IcePIC sampling table.

  - In certain cases (i.e., stronger winds), sampling immediately in the lee of the MSF can allow collection of particles that otherwise would be blown away. This is OK, but carefully describe the sampling location in the log file

  - Note the time that sampling starts.

- If the sampling location is a distance from the IcePICS shed, carry the slides to the shed in a box to prevent crystals from being blown off the slides in transit

- Install the camera on the microscope mount, place the slide on the stage, and place the light below the stage, directed towards the mirror.

- Adjust the positioning and brightness of the light, the focus (if necessary) and ISO settings for optimal results.

  - It can be helpful to take multiple photos of each slide view at different ISO settings, accomplished by holding the shutter button down for 3 seconds.

    - This is called “bracketing” and helps ensure that at least one of your pictures has good brightness and contrast.

    - The camera is set up to automatically cycle through three ISO settings.

- Your first and last photos should be of the slide scale

  - This helps identify situations where the slide is contaminated while taking pictures – it is surprising how many crystals can be knocked off of the microscope by the shutter’s action or how fast crystals can grow/melt/change

- Take representative photos of different crystals on each of the two slides, noting which set of photos is the upwind versus downwind slide.  Photos should not focus on interesting particles, but on capturing a representative survey of all particles.

- When finished, bring the slides, camera, and light back into the MSF. Clean the slides with alcohol and wipes, and place them back in their storage location

- Download the photos from the camera to a science technician laptop.

- Save each sampling session in its own folder titled “YYYYMMDDhhmmZ”

  - The file names CANNOT have spaces in them or Dataman will hang indefinitely when it tries to copy them to the ToBoulder folder. NO spaces in file names!

  - The folder name is case-sensitive as well.

- Create a *.txt log file for each sampling session in the folder with the photos:

  - Time of sampling, and duration if it is a longer duration sample

  - Temperature, RH, and winds at time of sampling

  - Location of sampling

  - Type of crystals

  - Any other comments or notes about the sample

  - A template for the log file is available on the Tech OneDrive:

/Manual and Supporting Documents/ICECAPS/IcePIC/YYYYMMDDhhmmZ.txt

- Review all of the photos, deleting improperly exposed or unfocused photos, and note the crystal types in the log file

- Transfer the entire sampling session folder to the ICECAPS Linux PC:

  - Open the WinSCP FTP client.

  - Select the existing profile for ‘dataman’ and click ‘Login’. If the profile does not exist, create a new profile:

    - File Protocol: SFTP

    - Host name: 192.168.0.73

    - Port: 22

    - Username: dataman

    - Password: see KeePass password manager

    - Save this profile to facilitate future access.

  - With the connection established, drag-and-drop the folder you’d prepared with IcePic data into this directory:

    - home/dataman/search/work/icepic

Cleaning the objective (once per phase, at turnover):

- Bring the microscope into the MSF, defrost, and remove all snow and rime

- Remove the lower objective from the microscope

  - The lower objective segment is the lowest 1” of the tube, just above the stage; it unthreads from the knurled section of tube just above it

- Clean the upper and lower lens of the objective very carefully using compressed air, then alcohol and an optical wipe

- Re-install the lower objective segment

- Using the camera, check the focus and adjust as needed

Retro:

None

Consumable Resupply Items:

Isopropyl alcohol

Optical wipes

Glass microscope slides

# ICECAPS – Precipitation Occurrence Sensor System (POSS) – MSF

Primary contact: 	Matt Shupe – matthew.shupe@noaa.gov

Secondary contact:	David Hudak – david.hudak@ec.gc.ca (do not contact directly; Matt will forward messages if necessary)

How often:		Daily: Regular sensor checks

Weekly: None

Clean requirements:		None

Supplies needed:		6’ Ladder

Soft-bristled brush

Introduction:

The POSS is on loan from Environment Canada; it measures precipitation occurrence, precipitation type, rain drop size distribution, and snow water equivalent using an X-band (10.5 Ghz) bi-static continuous wave Doppler radar with a range of 3m. The POSS system is mounted on a guyed mast to the southeast of the MSF.  On the mast are two small enclosures that contain parts of the instrument system and the data logging system.  Data are streamed from the data logging system to the POSACQ script running on the ICECAPS Linux computer, which records the data and generates error logs.

The transmitted power from the system is nominally 43 mW and thus is not a hazard to human operators.

Procedure:

Daily Checks

- On the ICECAPS Linux computer, verify that the POSACQ window is open and that a new data message arrives from the POSS every minute. Verify that the UTC time of the current data message matches that of the MSF time server

- DO NOT CLOSE the POSACQ window; as data recording will cease. Minimizing the window is OK

  - If the POSACQ window is closed accidentally, see System Startup procedure below

- In the POSS data folder (home/dataman/POSS/posacq/data):

  - Open the MESPOS and MAINLOG error log text files. The following errors are normal and acceptable:

    - transient noise

    - air temp

    - heater current

    - min noise limit

    - Max noise limit

    - DC offset (occasionally)

  - Report any other errors to the PI

  - Verify that there is a data file labeled “POSYYYYMMDD.txt” in the folder and that it is updating every minute

- At the instrument, verify that the instrument is clear of snow and ice, particularly the faces of the transmitter and receiver (the top of the slanted heads).  If snow or rime exists, use a ladder or brush attached to bamboo to access the instrument; with a soft brush, clear the transmitter and receiver windows and the divider plate of the snow and rime.  Attempt to minimize the blockage of the signals (i.e., work from the side as much as possible instead of the top). If water ice exists on the system, please communicate with the primary contact.

- Make a comment in the POSACQ data stream about the time and success of the instrument cleaning

    - In the POSACQ screen, between data messages, type ‘comment’ and then a brief description of actions, then enter

      - i.e. ”comment <insert time>daily checks perform all is well”

      - i.e. “comment <insert time>ice buildup on window, was cleaned”

    - One cannot type multi-line comments; type in short phrases

    - A sensor data message may arrive while typing a comment, this will displace what has been typed; just keep typing

System Startup

- Verify that the POSS instrument power cable is plugged into the outdoor power pole.

- Verify that there is NOT another instance of POSACQ running on the computer:

  - In a terminal window, type “ps –edf | grep posacq”

  - “psSPACE-edf|gregSPACEposacq”

  - If you see a line ending in “./posacq POSACQ1.PAR”, then the program is running

  - Typing “kill -9 <PID#>” will kill that instance of the posacq program, where <#PID> is the process identification number associated with the process found using the ps command above.

- In the very unlikely event that the POSACQ1 parameter file needs to be changed or verified, do so now

  - See POSACQ1.PAR Notes below

- To start data acquisition, change to the POSS acquisition directory by typing into a terminal window: “cd /home/dataman/POSS/posacq” then start acquisition by typing “./posacq1 POSACQ1.PAR”.  The system will start acquiring and data will spool to the window

- Perform the daily checks

System Shutdown

- On the ICECAPS Linux computer, close the POSACQ window

- Power the POSS instrument down by unplugging it from the outdoor power pole.

System Notes

- There is a data logger and other system hardware in the two white weather-tight enclosures just north and east of the base of the mast.  For troubleshooting system problems, it may be necessary to access the equipment inside of these boxes.

- Inside one of the boxes is a light bulb on a thermostat.  The objective of the bulb is to provide a small amount of heat to keep the electronics warmer than the outside ambient temperatures.  On occasion it may be necessary to change this light bulb.  Typically this system will be checked out by the investigator team during summer maintenance.

Raising the POSS

The POSS periodically needs to be excavated and raised to the surface. This is performed on an annual basis in June/July, or as needed due to accumulation. This is done both to maintain access to the POSS for service, and because the POSS data quality becomes increasingly impacted by blowing snow as the sensors approach the surface. If a POSS raise is scheduled, or if it appears necessary due to accumulation, take a photo of the current state of burial, and contact the PI and science PM to coordinate.

- To the extent possible, wait for a mild temperature day so that the electrical cables are not overly brittle. This is a good task to perform with a partner, due to the need to lift the POSS after excavation and to stabilize the system with guy wires once at the surface.

- Shut down and disconnect power from the POSS, following the standard protocol.

- Identify a new site for the POSS. This should be ~1 meter from the existing site and selected to minimize the change in distance from the MSF and change in electrical cable length. The relocation allows for the POSS to be placed just away from what will soon be an excavation pit. At the new site, compact the snow surface with your boots, adding snow as needed to create a surface-level platform. This is intended to reduce initial settling and minimize rounds of guy line adjustments required.

- The POSS has three guy lines, which are terminated in plywood anchors. Dig out each of the anchors, loosen or disconnect from the POSS stand, and re-bury anchors 50 cm below the surface, radially about the new site.

- The vertical POSS pole is mounted on a steel plate of 30 cm x 30 cm. Dig out the stand base, and with a partner, lift the POSS to the surface and move it onto the compacted snow of the new site. Cover and tamp snow over the base to anchor it in place.

- Confirm the POSS pole is plumb, then reattach and tighten guy lines.

- Secure electrical cables moved during relocation.

- Reconnect power and restart the POSS, following the standard protocol. Notify PIs.

- Monitor guy lines and retighten as the system settles.

POSACQ1.PAR Notes

- This file sets the parameters for the acquisition, recording, and error logging from the POSS. This file is read by POSACQ each time it is run

- Under normal circumstances, there should be no need to view or verify the contents of this file. If the file as found is not consistent with what is listed below, contact the PI for further instruction

- As of late 2013, the contents of the .PAR were as below:

  - POSACQ1.PAR_updated=27July2010/vs

  - site=SUM

  - maildest=steven.brady@ec.gc.ca;vlado.stojanovic@ec.gc.ca;matthew.shupe@noaa.gov

  - possendmail=1

  - auxsendmail=0

  - poscom=/dev/ttyS0

  - auxcom =/dev/ttyS1

  - pospresent=1

  - auxdevpresent=0

  - path=/home/dataman/POSS/posacq/

- These lines correspond to the following:

  - Line 1: Used for basic recordkeeping

  - Line 2: Used to designate the sensor location (i.e., SUM for Summit)

  - Line 3: Defines the email addresses to which to send fault messages

  - Lines 4&5: Used to define whether fault emails should be sent

  - Lines 6&7: Used to define the device address for the present instruments (requires program restart)

  - Lines 8&9: Used to define which instruments are present

  - Line 10: Used to define the “program path” (requires program to restart)

Retro:

None

Consumable Resupply Items:

None

# ICECAPS - Multi-Angle Snowflake Camera (MASC) – MSF

Primary contact: 	John Rausch, john.rausch@vanderbilt.edu

How often:	Daily: Regular checks

Weekly: Inspect instrument & clean window if needed

Annually: Reposition MASC

Clean requirements:

Supplies needed:		Soft cotton cloth, isopropyl alcohol

Introduction:

The Multi-Angle Snowflake Camera (MASC) is a device for imaging snowflakes from three angles, to better capture the three-dimensional structure of ice crystals.  Additionally, the device records the fall speed of the hydrometeors.  When used in conjunction with mass measurement devices, it allows an improved understanding of ice habit and equivalent liquid water content.

The instrument is located approximately 40 meters from the MSF in the SSW direction.  The outdoor electronics enclosure is mounted on a pole <1 meter from the tripod. Communication between the instrument and the managing computer is through USB and FireWire 800 (IEEE 1394b). Category 6 cable converters in the electronics enclosure are used for each interface to extend the range of the interfaces to the computer in the MSF.

Images are stored locally on the interfaced laptop under C:\MASC_Output\Summit. Once per hour, an IDL script executes to reduce the size of the images acquired in the previous hour. At 00:15Z the images from the previous day are transferred to Z:\MASC_Backup. Data from Z:\MASC_Backup are archived nightly to ICECAPS disk array.

Procedure:

Daily Checks:

Check for rime buildup on the IR sensors, the flashes, and the lenses. Brush any snow off and note the time you did so in the C:\MASC_Output\Summit\MASC_Summit_notes.txt file. A shortcut is on the desktop.

Verify that the MASC data acquisition program is running on the PC. The watchdog display updates once per minute. Verify that there are no reset events and that all three cameras are detected.  If everything is working properly, you should see this:

If “# dropped images <bad frames>” exceeds 10, alert the PIs.

Verify that images are being recorded if a snowfall event occurs.  Images are stored locally under C:\MASC_output\Summit\YYYY.MM.DD_hhZ.  Note that if the ice crystals are too small, the MASC’s infrared light gates may not detect them and therefore not photograph them.

Verify that the data reduction script is executing. Open the data_reduction_log.txt shortcut on the desktop. At the tail of the file, verify the script has been run hourly since the last check. If images were captured, the number processed within the hour will be displayed, otherwise a message of No images to process will appear.

Sometimes the mousepad on the computer will stop working. To troubleshoot this, press alt-tab and then type ‘stop’ in the MASC terminal window to shut down the program. Restart Windows by pressing the “Command” key (Equivalent to the Windows key) and type “cmd” to get a new terminal window. From there type “shutdown /r” to finish the reboot. If that does not work, use the power button.

Weekly Checks:

Verify the MASC head and electronics box are approximately 1 m off the ground. Clear any snow built up around the base. If they are less than ½ m off the ground, please raise the mounting poles for the MASC and electronics box.

Verify the camera is rigidly mounted on the pole, make sure it doesn’t spin easily.  If the MASC rotates, tighten the nut on the flange, attached to the base of the MASC, with an Allen key.  The nut is located on the section parallel to the main pipe.

Check the windows for each camera lens.  If they are dirty or water spotted, clean them with a soft cotton cloth.  If this is not effective, use isopropyl alcohol with as high a purity as possible.  DO NOT use ethanol/methanol/denatured alcohol as it may damage the silicon seals.  Also, use only a soft cotton cloth – DO NOT use paper towels as they tend to scratch the polycarbonate windows.

Verify the heating elements are functional on the unit.  The MASC body and imaging area should be free of snow and ice.

If any testing needs to be done (dropping snow through the decagon), please note the date and time of testing on the MASC_Summit_notes files on the desktop.

Annual Checks:

Raise the instrument and electronics enclosure if necessary.

System Startup:

Start the laptop in the MSF. The system will automatically login, but if needed, the username is johnra and the password is icecaps.

Warm Start (MASC has been off for less than an hour): Make sure the repeater units in the enclosure and MASC heat and power cords are plugged in and live.

Cold Start (MASC has been off for longer than an hour): Make sure only the heater is on inside the electronics box and MASC. Let the heaters run alone for 1 hour. After 1 hour, plug in the instrument power to the MASC and repeater units inside the electronics box.

Execute the RunMASC batch file on the desktop. The watchdog for the instrument will start. Verify there are no communication errors, and all three cameras appear in the watchdog.

System Shutdown:

Type stop and press return in the MASC terminal window.

Shutdown the laptop.

Inside the enclosure turn the power off to all devices.

Retro:

None

Consumable Resupply Items:

None

# ICECAPS – SOnic Detection And Ranging (SODAR) – MSF

Primary contact: 	 Matt Shupe, matthew.shupe@noaa.gov

Secondary contact:	 Bill Neff, william.neff@noaa.gov

Scott Abbott, scott.abbott@noaa.gov

How often:		Daily: Regular checks

Weekly, plus as needed after wind: Dish inspect and clean

Quarterly: Level check

As needed: Relocate to mitigate drifting

Clean requirements:		None

Supplies needed:		Long-handled windshield brush

Soft-bristled brush

Level

Flat-head screwdriver

Introduction:

This NOAA custom-built bi-static ‘sonic detection and ranging’ (SODAR) device emits audible acoustic pulses (chirps) and measures the backscatter from density variations in the lowest several hundred meters above the instrument. The returns are used to infer the depth and stability of the boundary layer.

The instrument is located on a stand to the east of the MSF and has a dedicated interface, power amplifier, and PC located in the MSF. Data are streamed to the PC and archived daily via a mount on the ICECAPS Linux archive system.

Procedure:

Daily Checks:

- On the SODAR PC, verify that the instrument is measuring and updating the onscreen display at least once per minute

  - The SODAR PC shares a monitor with the MWR PC, switchable via the KVM

  - Type “Scroll Lock-Scroll Lock-0-5” to switch from the MWR screen to the SODAR screen

- Verify that the measured data match prevailing conditions

- Verify that .jpg files are being regularly captured and stored in D:\EtlData\Sodar\JPG; review the sequence of .jpg files for consistency with prevailing conditions, and indications of noise interference

- In the SODAR log (D:\EtlData\Sodar\SODAR Site Log.txt), note the nature and time of any significant noise-generating events in camp that may be visible in the data (aircraft movements, heavy equipment near MSF, cleanings, etc.)

- When finished, leave the SodarPC.exe software open and maximized on the desktop with no overlying windows; this ensures that the screen captures will provide clear images

- Outside, verify that the SODAR is chirping at a consistent rate and acoustic frequency

Inspection and cleaning (weekly, plus as needed after wind):

- Note the time of inspection in the SODAR log

- Turn off power to the QSC amplifier.

- Open the SODAR front hatches and visually inspect the dishes for accumulation

- Cleaning is required if more than ½” inch of snow has accumulated in the SODAR dishes, or if the data are no longer consistent with prevailing conditions

- Open each compartment and, using the long-handled brush, carefully brush the snow off of the surface of the dishes

- On every visit, before closing the hatch, use a flathead screwdriver to clean snow from the grooves into which the hatch seats, and the upper groove in the hatch, as shown in this photograph:

- The top groove of the hatch can also become packed with snow after high wind events. If the hatch cannot slide high enough to gain adequate clearance on the bottom jamb, consider using a ladder to access the top of the SODAR box and using a thin piece of metal to clear packed snow from the upper groove.

- Fully close each hatch after cleaning.

- Check the height of the instrument cables running from the MSF to the SODAR; raise the bamboo supports and cables if the snow level is burying the cables.

- Restore power to the QSC amplifier.

- Note the duration of the cleaning in the log file

Leveling:

- Clean the SODAR dishes of snow or ice

- On the receiver (north) dish:

  - Carefully remove the microphone head from the stand and place it carefully, microphone up, on the dish

  - Place a level on the microphone mounting flange in the east-west and north-south directions

  - Shift the dish in its cradle so that the flange is level in both directions

  - Carefully replace the microphone

- On the transmitter (south) dish:

  - Procedure is the same as for the receiver dish except that the speaker need not be removed; place the level on top of the speaker

- Note the time and result of the leveling in the log file

System Startup

- Turn on the SODAR power strip in the back of the MSF computer rack and the black switch on the SODAR rack-mounted computer (located behind the keyed panel). This should power up the PSDSOD5 interface, the sodar computer, and the signal conditioner, but NOT the QSC power amplifier

  - The system may take up to 5 minutes to stabilize once powered up

- Upon reboot, an error is displayed: “Pri Master Hard Disk: S.M.A.R.T. status bad”.  This is normal; press F1 to continue.

- After logging in to Windows, uncheck the option to “protect your system…”

- On the SODAR PC, start the PatchMix DSP program using the icon on the desktop

  - The left column in PatchMix (WAVE 1/2) is the output from the power amp to the transmitter – inactive until the amp is powered up

  - The right column in PatchMix (PCI Card in L) is the signal from the receiver microphone. To test this signal, have an assistant clap their hands near the receiver compartment; the claps should be visible

- Start the SodarPC.exe software

- Within SodarPC.exe, click the sodar icon (dish with red vertical waves) in the upper left corner of the window. From the menu bar, select “Sodar” > “Auto Display”. The display window should begin updating, although without data

- Turn on the QSC PLX1202 power amplifier, located on the shelf below the table with the SODAR and ICECAPS monitors

  - The CH1 lights will start blinking

  - Assuming the amplifier levels have not been changed or need adjusting, no further action is needed

  - If the amplifier levels require adjusting, turn the “CH1” level knob clockwise until the top red LED, labelled “CLIP”, flickers. Then turn the knob counter-clockwise back one notch so that the “CLIP” LED nether flickers nor illuminates. The ideal setting for the “CH1” level knob is 28.

- Data onscreen should now be updating and valid; be sure to leave the SODAR screen maximized and in front so that the screen capture functions

System Shutdown

- In the SodarPC.exe screen, shut the data acquisition off by clicking the dish icon without red vertical lines in the upper left corner of the screen

- Turn off the QSC power amplifier

- Close the PatchMix DSP software

- Shut down the SODAR PC using the standard Windows shutdown

- Once the SODAR PC is off, turn off the SODAR powerstrip in the back of the MSF computer rack

Retro:

None

Consumable Resupply Items:

None

# ICECAPS – IceCAM – MSF

Primary contact: 	Von Walden, v. walden@wsu.edu

Secondary contact:	Matt Shupe, matthew.shupe@noaa.gov

How often:		Daily: Regular checks

Weekly: None

Clean requirements:		None

Supplies needed:		Soft-bristled brush

Introduction:

IceCAM is an Ubiquiti Networks airCam IP-based security camera mounted on the west railing of the MSF roof, just north of the ladder, facing to the east. It photographs the roof every fifteen minutes, with the objective of providing a photographic record of the MSF roof, and specifically the ICECAPS microwave radiometers and cloud radar. Photos are stored on the ICECAPS Linux server (/home/dataman/search/icammount) every fifteen minutes by a chron job. Those files are archived on the ICECAPS server and sent to Boulder by the regular daily archiving scripts.

Procedure:

Daily Checks

- Review the days’ photos on the ICECAPS Linux system (/home/dataman/search/icammount), verifying that a photo is saved every 15 minutes and that the camera is aimed properly

- On the roof, using a soft-bristled brush, gently brush any snow, ice, or rime off of the camera lens if needed.

Configuration Check

- The camera view and configuration are accessible by browser at 192.168.1.118; see keepass for username and password

- One can view the current image, remotely reboot the device, or reconfigure the settings from this window

- If the settings are lost, the configuration file as of November 2013 is stored at Server\ftp\science\ICECAPS\IceCAM. This config file can be uploaded to the camera from the bottom of the “DEVICE” tab on the webpage (Upload…..).  Default login: ubnt/ubnt

To change the time:

- Go to 192.168.1.118 and log in

- Change the timezone and save

- Change the timezone back to GMT and save

- The webcam will now connect to a time server and refresh the date/time

Power Up/Power Down

- To power up, check that the Ethernet cable from the camera is plugged into the Power over Ethernet (PoE) injector behind the computer rack and that the second Ethernet cable from the PoE injector is plugged into the main ICECAPS network switch. Plug the PoE injector into its wall outlet. Verify that the ICECAPS Linux server begins acquiring an image every 15 minutes.

- To power down, unplug the PoE injector from the wall outlet.

Retro:

None

Consumable Resupply Items:

None

# ICECAPS – Microwave Radiometers (MWR) – MSF

Primary contact: 	Olson, Erik, erik.olson@ssec.wisc.edu

How often:		Daily: Daily checks

Weekly: None

Semi-annually: Liquid nitrogen calibration

Clean requirements:		None

Supplies needed:		Soft-bristled brush

Calibration supplies:

Liquid nitrogen and transfer dewars

LN2 target well

Target frame and mirror

Introduction:

The MWR installation comprises two RPG microwave radiometers: the HATPRO and “HF,” or 150-90. The HATPRO, or “primary,” measures brightness temperatures on eight frequencies each on the flanks of the water (22.2 – 31.4 GHz) and oxygen (50.8 – 58.8 GHz) absorption lines. The RPG-150-90, or “secondary” radiometer, measures cloud absorption at 90 and 150 GHz. Data from these instruments are used to determine column liquid water path, cloud liquid water path, integrated vapor content, and temperature and humidity profiles through the troposphere. The instruments self-calibrate every 4 minutes by measuring radiation from an ambient temperature calibration target, and are calibrated semi-annually with a liquid nitrogen target.

The radiometers are mounted side-by-side on the north side of the MSF roof, with the HATPRO to the east and the 150-90 to the west; data from both are streamed to a single dedicated computer in the MSF. Data are archived locally and transferred to Boulder daily by the Dataman scripts running on the ICECAPS Linux PC.

Procedure:

Daily Checks

- Check the RPG-HATPRO Radiometer time (in the GPS-Clock section) and compare to the MSF time server time. Notify Erik if time difference is 5 seconds or greater. Note that the radiometer time in the RPG-150-90 section is not important and will drift since the GPS antenna has been removed from that instrument. In the current master / slave configuration that produces combined data files, the time in the data comes from the master (HATPRO) instrument. The local computer time is not important to the data, but will likely be pretty close to the MSF time server.

- In the “Status and Config.” tab of RPG-HATPRO:

  - Check the status of the GPS signal, the environmental measurements, and the timestamp

  - Verify that the software displays “MEASUREMENT RUNNING”

  - In the “Controllers” subsection of the Master and Slave, verify that “Main Contr.”, “Elev. Stepper”, and “Cal. Stepper” are responding

- In the “Brightness temps” tab of RPG-HATPRO:

  - Verify that the data are updating

    - Check the three real-time spectra on the right side of the screen. These take about 30 seconds to update, and the changes can be quite small.

  - Check the time series data for each of the eighteen channels; verify that periodic calibration spikes are present, that the 90° data are free of spurious noise, and that the data are consistent with the prevailing conditions

    - As of winter 2014, the 51.26 GHz band has known noise issues

    - Erik notes that various channels will drop to reading 0.00 in colder temps. This is a known issue.

- In the “Housekeeping” tab for both the HATPRO and 150-90, verify that there are no red status indicators. Report any red alarms to the PI immediately

  - As of winter 2014, the 150-90 “Receiver 2 Temperature Stability” will periodically go red, especially when one is on the roof of the MSF. This is a known issue, it typically clears itself, and is not worth reporting

  - The status indicator bars may not completely fill the window space. This can look like a data dropout, but is normal per Erik.

- Verify that five data files have been written to C:\Program Files\HATPRO\ data within the last hour

- On the roof of the MSF:

  - Carefully clear the radiometer main windows of any frost or rime with a soft-bristled brush

  - If there is significant accumulation (> 25 cm2), photograph the window and email to the PI.

  - Clear any rime away from the intake to the main blower

  - Verify that the main blower heater is functional – it should sublimate any snow or frost tossed or ingested into the blower

    - DO NOT TOUCH THE HEATER; it will burn instantly

  - The computer fan that aspirates the ambient temperature sensor, between the heater intake and the instrument body, is often frozen and not operational. The PIs are aware of this.

- Note anything unusual, any window cleanings, and any outages in the log file (C:\Program Files\RPG-HATPRO\HATPRO_field_notes.doc) on the MWR PC. Log entries for both radiometers should be placed in the HATPRO file; the older HF log file is no longer used

Common Error

GPS timeclock inaccurate

Notify Erik and he will instruct you to restart the MWR. No need to unplug and fully shut down the instrument. Just terminate the measurement and restart it. For further instruction see: C:\Users\ARO-SummitScienceTec\Polar Field Services\PFS-ARC-Summit Techs - Documents\Instrument Manuals and Supporting Documents\ICECAPS\MWR and document titled “GPS Timeclock Restart.docx”

Radiometer Shutdown

- In the RPG-HATPRO software, click the “Terminate Running Batch” button in the top toolbar. This button is a right-pointing blue triangle with a vertical line to its right

- Close the software

- Switch off the radiometers’ instrument power using the power strip in the ICECAPS computer rack. This strip is on the right side of the shelf above the NOAA-Sodar computer, below the drawers and the MWR computer

  - Note that this will not switch the heaters and blowers off, just the instrument. The heaters and blowers should be plugged into a separate, non-UPS protected outlet

  - If the instrument will be shut down for a long period, shut down the computer using the standard Windows shutdown, and unplug the two power cables to the heaters/blowers

Radiometer Startup/Restart

- Power the instruments back up using the power strip (see above)

  - If the instrument has been powered down for a long period, also power up the MWR computer and the heater/blowers. If the instruments have completely cooled to ambient temperature during non-summer season, the internal computers may have problems starting. Contact Dave and Erik if this happens for instructions.

- Re-open the RPG-HATPRO software at the PC. Software name is R2CH.

- Select Communication > Serial Configuration; verify that the software connects to both instruments. As of 2016-07-15, the HATPRO (Master) is set to COM4, and the 150-90 (Slave) is set to COM2. Click “Connect.”

- Select the “Send Measurement Batch File to Radiometer” button; this is two buttons to the right of the sun, or four buttons to the left of the “Terminate” button used earlier

- Choose the current batch file

  - As of 2014, this file is “SUMMIT_VERSION_11_SCAN.mbf”.  The default is to use the largest version number. (Load SUMMIT_VERSION_11_SCAN_HF.mbf for 150-90 radiometer only—ie. when HATPRO is out for repair/calibration.)

- Measurements will begin; conduct the daily checks and note the reboot/startup in the log file

Radiometer Liquid Nitrogen Absolute Calibration

- See “ MWR LN2 Calibration” instructions among instrument SOPs.

- Additional protocols updated by T and Sierra are located at C:\Users\ARO-SummitScienceTec\Polar Field Services\PFS-ARC-Summit Techs - Documents\Instrument Manuals and Supporting Documents\ICECAPS\MWR\LN2 calibration photos

Retro:

None

Consumable Resupply Items:

None

# ICECAPS-ACE – BIRAL – MSF

Contact: 	Heather Guy, heather.guy@ncas.ac.uk

Ryan Neely, ryan.neely@ncas.ac.uk

How Often:		Daily: Daily checks

Clean requirements: 	Clean frost/rime from hoods / windows

Supplies needed: 	Soft-bristled brush, alcohol (isopropanol) and lens cleaning cloth.

Access instrument log: https://docs.google.com/spreadsheets/d/130M7i6kIW5GgOFN35G3g53tDQgkNICgnL53GXTaL9gI/edit?usp=sharing

Access data quicklook plots: https://192.168.1.196/index_vis.html

Introduction:

The Biral is an SWS-100 series visibility and present weather sensor. There are two of these at Summit, “Beryl” is the primary installed on MSF roof, and “Bert” is the backup stored in MSF vestibule. The instrument consists of two arms facing each other, the right arm is an optical transmitter, and the left arm is a forward scatter receiver. Based on the measurement of forward scatter of the transmitted light, the instrument derives a measure of visibility (distance in m) and precipitation type. As such the path between the transmitted and receiver should not be obscured (although it is safe to do so, for example to clean the windows). There are two sets of heaters, the hood heaters (red heating pads under the hood of both arms) are powered by a separate 24V cable. The internal window heaters are powered by the main sensor cable. The 24V power supply for the Biral is under the desk in the MSF to the left of the MPL. The instrument itself is mounted on the small tower structure on the MSF roof.

24V power supply

Hood heater power

Instrument power

Instrument data

Port 1 light flashing means communication from instrument

Procedure:

Daily checks

- Visually check the condition of the sensor hoods and windows on the MSF roof. Remove any rime build up on the hoods and above the sensor on the tower with a soft bristled brush and log the time in the instrument log.

- If there is any noticeable obscuration of the instrument windows, clean the window with some isopropanol on a lens cleaning tissue or cloth. Log the cleaning time in the instrument log.

- Check the data on the quicklook plot, accessible at: 192.168.1.196/index_vis.html. Ensure that both plots have updated in the last two hours (check the time on the bottom right of the x-axis) and report if it has not. Monitor the data and report any extreme or unusual values.

- If ‘bad-self check’ values (shown by red lines on the quicklook plot), clean the instrument windows with isopropanol and a soft cloth as above.

- Complete the daily log and report all daily findings.

- If the local quicklooks page is not loading, you can also check the remote page at: Visiometer Quicklooks (leeds.ac.uk)

- Use WinSCP to verify an hourly file has been created and it is not empty (0 kb)

- The raspberry pi (ace-pi-2) can be accessed using ssh over the terraterm program. With the ip address: 192.168.1.196, username: ace-pi, password: ncasflux
To ensure the daily files are updating, you can use the following commands:

- cd /home/data/       (to change to the data directory)

- ls -latr                    (this will list all the data files in this directory with the most recent one last. Check that there is a *.biral file produced within the last hour and that the file is not empty, i.e. file size > 0 )

Every day the data are backed up to the external hard drive connected to ace-pi-2 in the MSF and transferred to dataman for back-up up with the rest of the ICECAPS data. Note that if the external hard drive becomes disconnected from the raspberry-pi for any reason, contact a PI to remount the drive – it will not mount automatically.

Power down:

To power down the instrument, disconnect the power supply box (grey box in photo above) from the extension lead that connects it to the UPS outlet. Inform PIs and note in log.

Power up:

Ensure rime and ice is cleared from the instrument hoods and windows on the MSF roof. Connect the 24V power supply to the extension lead going to the UPS outlet. Port 1 on the Moxa device (see photo above) should begin to flash indicating that the instrument is communicating. Inform PIs and note in log.

Data file access and locations:

Data are stored in hourly files at ace-pi@192.168.1.196:/home/data/

Consumable Resupply Items:

None

Retro:

None

# ICECAPS-ACE - Optical Particle Counter (OPC) – MSF

Contacts:		Ryan Neely, ryan.neely@ncas.ac.uk

Heather Guy, heather.guy@ncas.ac.uk

How Often: 		Daily: daily checks.

Clean requirements: 	None.

Supplies needed: 	Soft-bristled brush.

Coffee filters or equivalent.

Electrical tape.

Self-amalgamating tape (mastic).

Access to instrument logs:

MSF OPC: https://docs.google.com/spreadsheets/d/130M7i6kIW5GgOFN35G3g53tDQgkNICgnL53GXTaL9gI/edit?usp=sharing

AWO OPC:

https://docs.google.com/spreadsheets/d/110xWWWfh8if-Hx16oX-qpALVk6zBa4JkS_xvpjKA9e4/edit?usp=sharing

Access to data quicklook plots:  https://192.168.1.196/index_ace.html

Introduction:

The Optical Particle Counter (OPC) is an Alphasense N3, commercial grade aerosol optical particle counter that measures particle size distributions in real time. There are now two ICECAPS-ACE OPC’s; the MSF-OPC is located on top of the 2m tower on the MSF roof, the AWO-OPC is attached to the south railing on the AWO roof. On the lower side of the OPC is a metal inlet, on the top is an exhaust fan. There is a section of plastic piping mounted around the inlet that protects it from wind-blown riming (the wind shield). Air is pumped through the OPC and a laser is used to count particles and estimate particle size. The OPC can detect particles between 0.38 and 16 𝜇m.

The OPC’s are directly connected to the ICECAPS-ACE raspberry pi’s inside the building via USB.

Procedure:

Daily checks

- On the roof (AWO and MSF), check airflow inlet and exhaust fan on the OPC. Use a soft-bristled brush to remove any frost or rime. Ensure that neither the inlet nor the outflow are obstructed. Ensure the fan is working.

- In the winter months rime build up might start to fall into the exhaust fan and cause it to stop turning. To prevent this from happening, attach a coffee filter over the exhaust fan with electrical tape (defrost the OPC first so that the tape will stick). Do not pull the coffee filter tight over the fan – make it into a little ‘balloon’ so that it restricts flow as little as possible.

- In high winds, the cable connection at the instrument can become disconnected and cause the instrument to stop logging periodically. To prevent this, use self-fusing tape to secure the cable to the instrument.

- Check the data on the ICECAPS-ACE aerosol instrument quicklooks, which can be accessed at: http://192.168.1.196/index_ace.html. Ensure the relevant plot has updated in the last two hours (look at the latest time indicated on the right of the x-axis) and report if it has not. Monitor the data and report any extreme or unusual values. If the counts or flow rate drop to zero for an unusually long period, the instrument may need to be brought inside to defrost for 24 hours.

- If the local quicklooks page is not loading, you can also check the remote page at: Aerosol Quicklooks (leeds.ac.uk)

- Use WinSCP to verify an hourly file has been created and it is not empty (0 kb)

  - Data are stored in hourly files at ace-pi@192.168.1.196:/home/data/

  - The raspberry pi (ace-pi-2) can be accessed using ssh over the terraterm program. With the ip address: 192.168.1.196, username: ace-pi, password: ncasflux
To ensure the daily files are updating, you can use the following commands:

  - cd /home/data/       (to change to the data directory)

  - ls -latr                    (this will list all the data files in this directory with the most recent one last, check that there is a .biral file produced within the last hour and that the file is not empty, i.e. file size > 0 )

- Complete the daily log. Report all daily findings, and anything that might affect aerosol sampling (for example, heavy equipment operating around the MSF).

Power down:

The OPC is powered through the USB connection to the computer. To turn off the OPC, disconnect the USB.

Power up:

- Reconnect the USB to the computer (in the same port if possible).

- Check that quicklook plots start updating in the next hour.

Every day the data are backed up to the external hard drive connected to ace-pi-2 in the MSF and transferred to dataman for back-up up with the rest of the ICECAPS data. Note that if the external hard drive becomes disconnected from the raspberry-pi for any reason, contact a PI to remount the drive – it will not mount automatically.

Consumable Resupply Items:

None

Retro:

None

# ICECAPS – Albedo Stand and Solar Tracker – AWO

Primary contact: 	Michael Gallagher, michael.r.gallagher@noaa.gov

Secondary contacts:	Matt Shupe, matthew.shupe@noaa.gov

How often:  	Daily – instrument check and cleaning

Weekly – data download

As needed- raising fixed arm instrument

Clean requirements:  	None

Supplies needed: 	Log sheets, laptop computer for data download

Introduction:

The energy balance at the surface of the ice sheet is a key climate parameter, and the albedo stand and solar tracker measure the solar radiation, albedo, and longwave emission components of this energy budget. Major radiation components have been measured at Summit since the year 2000 (with a break in winter 2002/3 and spring 18). This long-term radiation monitoring allows analysis of both short-term variability and longer trends in short and long wave radiation.

Procedure:

- Daily check of the instruments:  The daily checks should be done before 10 am and after each rime event.

- General note: Disturb snow surface as little as possible and use same footprints every time you access the instruments  -> Figures 1a,b and 2a,b show the locations of all the instruments (‘I*’), the filters (‘F*’).

- Fixed arm (Fig 1a,b): Clean rime off black filters underneath the up-facing instrument (F1 and F2) and at the end of the two orange tubes (F3 and F4). If there is rime on glass domes or dark dome (I1-I3) remove it with Kimwipes and ethanol.

- After check/cleaning, record time and status in logsheet:

\\ftp\science\zurich\YY_FTP_ETH\YY_ETH_Logsheets.xls

Weekly data download:

Throughout this process, record the appropriate information in the Swiss Logsheet:

\ftp\science\zurich\YY_FTP_ETH\YY_ETH_Logsheets.xls

Solar Tracker

- On Dell ruggedized laptop in AWO (or other computer with RealPort and LoggerNet software) open LoggerNet

- Browse to ‘Main’ and click ‘Connect’ and then select ‘Solar Tracker – CR23X’

- If an error occurs when attempting to connect to the solar tracker, first click ‘disconnect’ on the software, then power cycle the Digi-One (plugged in behind the HATS sampler), then try to connect again.  Multiple power cycles of the Digi-One may be necessary in order to establish a connection.

- If the connection still does not work then click the Start/Windows icon and type ‘Device Manager’, open dropdown menu for COM and LPT ports, and check if COM5 – Digi One is there.  If not first check if the Digi One is visible by opening a web browser and typing 192.168.1.113.  This should take you to the login page for the Digi One serial to Ethernet converter device.  The username is ‘root’ and the password is ‘eth_wsl’.  Login to the Digi One with these settings.  Now open the RealPort software and make sure the Digi One with this IP address is located and assigned at COM5 port on the computer.  These steps should diagnose any software issues.  If it still doesn’t work there may be a lose connection or hardware issue.

- Click ‘Connect’ in the upper left.

- Next, determine how much the clock for the instrument’s data logger has drifted since the last data download by noting the difference between the datalogger time and computer system time.  This information can be found in the “Clocks” field in the upper right of the “Connect Screen”.  Datalogger time is displayed as “Station Date/Time” and the computer system time is displayed as “Adjusted Server Date/Time”.  Record the time difference in seconds in the ETH Logsheet under the tab ‘Weekly Download Solar Tracker’ (\ftp\science\zurich\YY_ETH_Logsheet), such that the value is a positive number if the “Station Date/Time” is ahead of the “Adjusted Server Date/Time” or a negative number if the “Station Date/Time” has fallen behind.

- Click ‘set’ to update time. Note difference between these clocks after setting time in ETH Logsheet.

- Click ‘Collect Now’ to download the data from the logger.  The default location for the data are on the network drive \ftp\science\zurich\DIF.dat (if it has been more than 1 week there may also be a DIF2.dat file, copy this over too).

- It may take ~40 min to download the data,

- Once transfer is complete, click ‘Disconnect’ in the upper left for the logger to resume logging data.

- Copy the file at \ftp\science\zurich\DIF.dat to the new weekly directory you created at \ftp\science\zurich\YY_FTP_ETH\YY_FTP_ETH\week##\.

- Record file parameters in Solar Tracker spreadsheet: YY_ETH_Logsheets

- Modify the file name of the new file in the \week##\ directory to match the incrementing DIF###.dat format in the ETH Logsheet.

Albedo Stand (Fixed Arm)

Process is the same as for the Solar Tracker, with the following adjustments.

- Browse to ‘Main’ and click ‘Connect’ and then select ‘Albedo Stand – CR1000’

- In case of connection issues, confirm operation of the Digi-One device on COM8.

- The appropriate tab in the ETH spreadsheet is ‘Weekly Download Albedo Stand’.

- The transferred data will appear at \ftp\science\zurich\ALB.dat

- Files should be named following the ALB###.dat convention outlined in the ETH Logsheet for Albedo Stand.

Instrument Shutdown:

To shut down the albedo stand unplug its power cord from the power box at the base of the AWO tower.

I3

I2

I1

Fig 1a: Fixed radiometer stand with instruments I 1-3.

Fig 1b: Fixed radiometer stand with filters F 1-4.

# ICECAPS-ACE – POPS – AWO

Contact: Ryan Neely, ryan.neely@ncas.ac.uk

Heather Guy, heather.guy@ncas.ac.uk

How Often:		Daily: Daily checks

Clean requirements: 	Clean frost/rime from inlet / exhaust fan.

Supplies needed: 	Soft-bristled brush

Access instrument log: https://docs.google.com/spreadsheets/d/110xWWWfh8if-Hx16oX-qpALVk6zBa4JkS_xvpjKA9e4/edit?usp=sharing

Access data quicklook plots: https://192.168.1.196/index_ace.html

Exhaust fan

Inlet

Power button

Introduction:

The Portable Optical Particle Spectrometer (POPS) is an optical particle counter by Handix Scientific that is similar in functionality to the SKYOPC and OPC-N3s. The instrument is mounted on the AWO inlet mast, it has its own computer and pump that are all located inside the main instrument box. The sample inlet is protected by a wind shield in the same way as the OPC-N3s. This instrument fills a data gap by measuring the size distribution of aerosol particles that are smaller than those that can be detected by the SKYOPC and OPC-N3 (between 140 nm and 3 um). The instrument is powered by a 12V power supply on the aerosol bench in the AWO. Data are stored locally and connected to Summit local network by an ethernet cable from the instrument to the network switch in the AWO. Data are transferred directly back to Leeds each day from the POPS internal computer.

Procedure:

Daily checks

- Check the condition of the POPS inlet on the AWO inlet mast. Use a soft-bristled brush to remove frost or rime from inlet and ensure it is not obstructed. Brush rime/frost from the exhaust fan on the underside of the instrument box.

  - During storms, snow can be forced into the POPS enclosure through the exhaust fan outlet. To mitigate this we have traditionally stuck a coffee filter (from the big house) loosely over the exhaust fans using electrical tape. This has worked well, but make sure the filter is loose (not stretched tight) so that it doesn’t impact air flow.

- Check the data on the quicklook plot, accessible at: https://192.168.1.196/index_ace.html, Ensure that the relevant plot has updated in the last two days (check the time on the bottom right of the x-axis) and report if it has not. Monitor the data and report any extreme or unusual values.

- Complete the daily log and report all daily findings, and anything that might affect aerosol sampling (for example, soldering in the AWO).

Power down:

- Never disconnect power from the instrument before first safely shutting down the system or the POPS computer can become corrupted.

- On the AWO roof, press and hold the power button (figure above) until the system powers down. It can take the system several seconds to recognize the shutdown command. When the system has shut down the pump will turn off and the system status LEDs inside the POPS enclosure will no longer be illuminated.

- It is now safe to disconnect the power cable from the instrument.

Power up:

- Connect the power cable to the instrument.

- Press and hold the instrument power button (see figure above) for about 1-2 seconds. The pump should begin operating within a minute and quickly settle on a steady operating speed to maintain a constant sample flow rate. After another minute, the POPS data acquisition software will automatically load the current configuration, begin acquiring and saving particle data to file.

- You can check for normal operational by looking at the LED lights inside the POPS enclosure. The first 3 LEDs indicate ‘power on and system configured’. During normal operation, the LED labelled 13 should be blinking (every second). If the LED labelled 14 is illuminated (red) this indicates a fault, leave instrument powered on and contact PI.

- Check that the quicklook plots update within the next hour.

Accessing the POPS internal computer (request permission from PI to do this)

The POPS internal computer can be accessed using terraterm with the following information:

IP address: 192.168.1.234

Username: root

Password: 1

Data is stored in the following directory:

/media/usb0/Data

Consumable Resupply Items:  None

Retro:  None

# ICECAPS-ACE – SKYOPC – AWO

Contact: Ryan Neely, ryan.neely@ncas.ac.uk

Heather Guy, heather.guy@ncas.ac.uk

How Often:		Daily: Daily checks

Clean requirements: 	None

Supplies needed: 	None

Access to ACE instrument log: https://docs.google.com/spreadsheets/d/1Y2gL0gLrTa2IRyox93SMCI8-y4Hi9l-6qAvazORQ9Bs/edit?usp=sharing

Access to ACE quicklook plots: https://192.168.1.196/index_ace.html

Introduction:

SKYOPC inlet

SKYOPC instrument

Outlet line to vacuum pump

Moxa Nport

ACE raspberry pi’s acting as data loggers

The SKYOPC is an optical particle counter by GRIMM and is similar in functionality to the Alphasense OPC’s. It consists of the spectrometer itself (on the AWO aerosol instrument shelf), sample inlet (shared with the NOAA aerosol instruments) and external vacuum pump (AWO vestibule labelled ‘SKYOPC pump’). This device was built for continuous measurement of concentration and particle size distribution of aerosol particles irrespective of the ambient air pressure. The data cable is connected to a Moxa Nport device (labelled in photo below), if data are being transferred correctly, the orange light on the SKYOPC port of the moxa should flash every 6 seconds.

Procedure:

Daily checks

- Inside the AWO vestibule, check that the pump is not too hot by placing the back of your hand against the pump box, it should be warm but not hot to touch. If you notice any changes in the temperature, sound or smell of the pump, immediately power down the pump and contact the POCs.

- Check the data on the ICECAPS-ACE aerosol instrument quicklooks.

Ensure that both plots have updated in the last two hours (check the time on the bottom right of the x-axis) and report if it has not. Monitor the data and report any extreme or unusual values.

Complete the daily log. Add any major events on the ‘master’ tab, and report daily data logging check and inlet icing check in the ‘daily log’ tabs. Report all daily findings, and anything that might affect aerosol sampling (for example, soldering in the AWO).

Power down:

- Enter power down time and reason in instrument log.

- First power down the pump by unplugging it.

- To power down the SKYOPC it can simply be unplugged.

Power up:

- Plug in the instrument, the SKYOPC should start automatically, if data are being communicated successfully, the port on the Moxa serial-port switch should have a flashing orange light.

- Plug in the external pump.

- Check the quicklook plots start updating in the next hour.

Data file access and locations:

Data are stored in hourly files at ace-pi@192.168.1.233:/home/data/

The raspberry pi (ace-pi-3) can be accessed using ssh over the terraterm program. With the ip address: 192.168.1.233, username: ace-pi, password: ncasflux
To ensure the daily files are updating, you can use the following commands:

cd /home/data/       (to change to the data directory)

ls -latr                    (this will list all the data files in this directory with the most recent one last, check that there is a .skyopc file produced within the last hour and that the file is not empty, i.e. file size > 0 )

Consumable Resupply Items:

Pump service kit, to travel to Summit in fall 2022

Retro:

None

# ICECAPS-ACE – Fluxtower – AWO

Contact:	Ryan Neely, ryan.neely@ncas.ac.uk
Heather Guy, heather.guy@ncas.ac.uk

Location of fluxtower log: https://docs.google.com/spreadsheets/d/110xWWWfh8if-Hx16oX-qpALVk6zBa4JkS_xvpjKA9e4/edit?usp=sharing

Address of fluxtower quicklook page: https://192.168.1.211/index_fluxtower.html

How Often: 	Daily: Daily checks
Annually: Raise instrumentation

Clean requirements:    Clean rime/frost from reachable aspirated shields and Licor window when necessary.

Supplies needed: 	Lens cleaning tissue / cloth. Soft bristled brush.

Introduction:

The fluxtower suite consists of 11 instruments mounted on the AWO Met tower, and the SIMBA thermistor string between the met tower and the albedo stand. The combined measurements from the Met tower instruments will be used to calculate turbulent heat and momentum fluxes required to estimate the surface energy budget (SEB) of the Ice Sheet.

The ICECAPS-ACE Enclosure, mounted at head height on the tower, controls the power distribution to each instrument, receives the data and passes it to the SRI wireless relay that transmits the data to the local network.

The fluxtower computer is a raspberry pi (labelled ace-pi-1) mounted on the side of aerosol instrument bench in the AWO. This computer logs all of the data from the ICECAPS-ACE fluxtower instrument suite. Note that if the wireless relay between the tower and the AWO is down, no data are logged, thus this infrastructure is a critical part of our experiment and the PI’s should be made aware of any issues or outages.

Each fluxtower instrument has a dedicated SOP. Use the individual instrument SOP’s if you have concerns or questions about one of the instruments, but you can just use this SOP for normal daily operations. There is an additional SOP that describes the layout and functionality of the ICECAPS-ACE Enclosure. This SOP describes the entire fluxtower setup in general, including locations of instruments on the tower and the data transfer on the raspberry pi.

All events related to the fluxtower or that may affect fluxtower data should be recorded in the fluxtower log, this includes: daily checks, work on or around the Met tower, internet outages ect.

Fluxtower Map:

This photo shows the fluxtower after the May 2019 install. Items in bold are ICEAPS-ACE instrumentation. In May 2019 the tower was approximately 14m tall. The Enclsoure, “2m” boom and Snow depth sensor (SnD) were installed approximately 1.5m from the snow surface. The booms are named after their ‘ideal’ height above snow surface but this does not have to be exact, due to the positioning of the other items on the tower, the “8m” boom was installed at around 9.3 m above snow surface and the “15m” boom is at 13 m.

Procedure:

Daily checks

- Check the fluxtower instruments quicklook page to ensure all figures have updated within the last hour. If the data looks odd or inconsistent to you, it probably is, make a note of it and contact the PI. A couple of things in particular to look out for:

  - The temperature and humidity from the four sensors should be similar, although large temperature differences between the top and the bottom of the tower can exist they are unlikely to persist in a systematic fashion. If one sensor is notably inconsistent with the others over multiple days, this is usually a sign of a failed aspirated shield fan, contact the PI.

  - The LiCOR does not log data when it is iced, or when it’s path is interrupted (by snow, blowing snow ect.). So it’s common to see big gaps in the LiCOR data (typically plotted as a straight line). The LiCOR should recover by itself as soon as the obstruction has cleared.

- Check the instruments on the tower for ice build up. Clean any ice off the LiCOR and the lowest aspirated shield with a soft-bristled brush. If there is any snow build up infront of the KT15 lens, remove very carefully with a soft-bristled brush (avoid touching the lens). Check that you can hear the lowest level aspirated shield running and the Metek buzzing. Make a note of the state of the instruments that are higher on the tower, it is not necessary to climb the tower to remove ice unless instructed by the PI.

- Complete fluxtower log: Complete the daily data and instrument checks, and log any other information relevant to the fluxtower including work on or around the tower, internet outages, ect, in the ‘master’ tab.

Data file access and locations:

Data are stored in hourly files at ace-pi@192.168.1.211:/home/data/

The raspberry pi (ace-pi-1) can be accessed using ssh over the terraterm program. With the ip address: 192.168.1.211, username: ace-pi, password: ncasflux

Terraterm settings:

To ensure the daily files are updating, you can use the following commands:

cd /home/data/       (to change to the data directory)

ls -latr                    (this will list all the data files in this directory with the most recent one last, check that there is a .simba file produced within the last hour and that the file is not empty, i.e. file size > 0 )

Hourly files have the following naming convention: YYMMDD_HHMMSS.ss.instrument

Everyday the data are backed up to the external hard drive connected to ace-pi-1 in the AWO and transferred to dataman for back-up up with the rest of the ICECAPS data. Note that if the external hard drive becomes disconnected from the raspberry-pi for any reason, contact a PI to remount the drive – it will not mount automatically.

Powering down individual instruments using the Tripplite software:

For more information / screenshots of the triplite software, or for instructions for a complete enclosure power down, see the fluxtower enclosure SOP.

- Open a new browser window and connect to the Tripplite power distribution unit in the fluxtower enclosure by navigating to https://192.168.1.205

- Log in using the username localadmin and password localadmin

- Navigate to the ‘Loads’ page. You should see a list of power ports, next to some of these there are descriptions of what the port powers. One port can power multiple instruments. Some of the ports are not connected. In some of the descriptions there is a note saying ‘DO NOT TURN OFF’ – pay heed to this.

- Identify the port you need to turn off to power down the instrument of interest. Power down the port by switching the green button next to that port in the ‘state’ column and selecting confirm.

- Use the same green button to power the instrument back on when necessary, note all times in the log.

Power down / reboot fluxtower raspberry pi:

Log onto ace-pi@192.168.1.211 using terra-term by following the instructions in the data access section above. 
Before running either of the commands below, double check that you are connected to the correct raspberry pi (192.168.1.211)

To restart the raspberry pi run the following command:

sudo shutdown -r now

To completely shut down the raspberry pi run the following command:

sudo shutdown -h now

Power up fluxtower raspberry pi:

The raspberry pi can be rebooted after a safe shutdown (described above) by disconnecting and re-connecting it’s mini-usb power cable. It can take up to 5 minutes to boot, so wait this long before checking data. To ensure it has started up correctly log onto terraterm per instructions in the data access section above and check all data are being collected.

Raising equipment on tower:

It will be necessary to raise the equipment on the fluxtower as snow accumulates. The SnD can not operate from less than 50 cm above the snow surface, however due to the locations of the cables under the ICECAPS-ACE Enclosure, it is likely that equipment will need to be raised before this limit is reached. This is a big job involving tower climbing therefore should be implemented before winter if necessary.

Mid-August: Evaluate the need to raise the tower equipment before the winter season. If it appears likely that the snow surface will approach 50 cm below the SnD over winter, a raise will be necessary - contact PI’s for instructions.

# ICECAPS-ACE – SIMBA – AWO

Contact: Ryan Neely, ryan.neely@ncas.ac.uk

Heather Guy, heather.guy@ncas.ac.uk

How Often:		Daily: Daily checks

Clean requirements: 	None

Supplies needed: 	None

Access instrument log: https://docs.google.com/spreadsheets/d/110xWWWfh8if-Hx16oX-qpALVk6zBa4JkS_xvpjKA9e4/edit?usp=sharing

Access data quicklook plots: https://192.168.1.211/index_simba.html

Introduction:

Data cable going to fluxtower enclosure

Temperature sensor

Thermistor string (extends 3 m into the ice)

Instrument computer, batteries, and local data storage

SIMBA stands for SRSL Ice Mass Balance Array- it is a chain of temperature sensors, spaced 2 cm apart. The chain extends vertically 3 m below the surface, and 2 m above the surface. The upper section is supported by a mast which is guyed down for stability. There is an additional air temperature sensor at the top of the mast, and the data logger box is a yellow pelican case on the snow surface, tied to a bamboo pole in case of strong winds. Snow can be allowed to accumulate on the pelican case, but it should be flagged well. The instrument is powered by batteries that are inside the yellow pelican case. Note: As of August 2023, a hardwired 12V power cable has been installed from the ICECAPS-ACE fluxtower enclosure on the Met tower to the SIMBA unit. This is now the main power source for SIMBA. The batteries and battery connector are still in the simba pelican case and have enough charge to power simba for a few months. Do not reconnect the batteries without consulting Heather first. At no point should the battery and fluxtower power be connected simultaneously. The data cable runs from the pelican case to the fluxtower enclosure, and the data are bundled with the rest of the fluxtower data. Data are collected every 15 minutes.

Other updates from August 2023:

- A second simba test unit was installed next to the primary simba unit (see below). This unit is completely self-sufficient with no need for daily checks. Please report if there are any obvious visual changes or damage to this system.

- A spare simba thermistor chain and PCB are now stored in AWO (see below).

- Note that there is a known issue with the simba thermistor chain which means IF THIS SYSTEM IS POWERED DOWN IT MAY NOT START BACK UP CORRECTLY. Avoid powering down the fluxtower enclosure where possible. If it does get powered down follow instructions in the ‘power up’ section for the best chance of getting it started back up.

Procedure:

Daily checks

- At the instrument, remove excess slack from the guy lines if necessary by sliding up the half-fisherman’s knot.

- Check the data on the quicklook plot, accessible at: https://192.168.1.211/index_simba.html. Ensure that the plot has updated in the last two days (check the time on the bottom right of the x-axis) and report if it has not. Monitor the data and report any extreme or unusual values.

- Complete the daily log and report any unusual findings.

Instrument reset (only do this if requested by a PI):

- Expose the yellow pelican case, it may be buried, if so take care not to damage the three cables (chain, data, and power) when you are uncovering it (see photo above).

- Ensure both the chain and data cable are well connected. Open the case (try to avoid getting snow inside the case).

- Press the ‘wake’ button to wake the system up (see photo below). Make a note of which LED’s light up when the system is woken up.

- Press the ‘reset’ button to reset the instrument. It will then go through its initialization phase that can last 5 minutes or so. This is described in the ‘power up’ section below.

Power (battery) connection

Data (serial) connection

GPS receiver

SD card (local storage and config)

Reset button (R) and wake button (W)

System health LEDS

Power down:

- Expose the yellow pelican case, it may be buried, if so, take care not to damage the data cable when you are uncovering it.

If it’s connected to power from the fluxtower enclosure:

- Disconnect the blue power cable from the outside of the pelican case (see photo above). It’s a push fit connector that only it’s in one way.

If it’s connected to batteries:

- Open the case and disconnect the power connection (identified in the photo above). Try to avoid getting snow inside the pelican case.

Power up:

- Make sure the chain is connected first (white cable), then reconnect the power connection.

- Sensor initialization takes around 5 minutes. Watch the LED light up sequence to ensure that initialization is successful:

  - The 5 LEDS on the board will flash from left to right to indicate that initialization has begun.

  - Then the first three lights in turn will go solid and remain solid.

  - After this an error code will flash 3 times (the same code, 3 repeats)

  - Then just the first system LED will remain red. System is ready to go.

- If point 2 doesn’t happen (i.e. any of the first 3 lights don’t light up and remain solid) check all of the connections, then press the reset button and try the initialization again.

- There is a known issue with this thermistor chain which means that the chain is not always detected, if this happens, the 2nd LED will go out instead of remaining solid. This means SIMBA has not started up correctly. After three reset attempts, if this problem persists try following the ‘EPROM power up sequence’ at the end of this document.

Data file access and locations:

Data are stored in hourly files at ace-pi@192.168.1.211:/home/data/

The raspberry pi (ace-pi-1) can be accessed using ssh over the terraterm program. With the ip address: 192.168.1.211, username: ace-pi, password: ncasflux

Terraterm settings:

To ensure the daily files are updating, you can use the following commands:

cd /home/data/       (to change to the data directory)

ls -latr                    (this will list all the data files in this directory with the most recent one last, check that there is a .simba file produced within the last hour and that the file is not empty, i.e. file size > 0 )

Everyday the data are backed up to the external hard drive connected to ace-pi-1 in the AWO and transferred to dataman for back-up up with the rest of the ICECAPS data. Note that if the external hard drive becomes disconnected from the raspberry-pi for any reason, contact a PI to remount the drive – it will not mount automatically.

Consumable Resupply Items: None

Retro: None

EPROM power up sequence

When SIMBA does not power up correctly (2nd LED does not remain solid), it is because it is not detecting one of the two EPROM chips at the head of the thermistor chain.

Below is a photo of the location of the two (black squares) EPROM chips at the head of thermistor chain relative to the cables going into the chain. They are covered by white heat shrink, but you can feel the location of the chips through the head shrink.

To get SIMBA to come back online correctly, you need to gently put pressure on (squeeze together) the two EPROM chips during the ‘reset’ sequence. I.e. One person squeezes the chips, whilst another person hits the reset button and watches the LED sequence. One the LED reset sequence is complete, the chips can be released (the unit only needs to read the data from these chips once).

# ICECAPS-ACE – Enclosure – AWO

Contact:	Ryan Neely, ryan.neely@ncas.ac.uk

Heather Guy, heather.guy@ncas.ac.uk

Location of fluxtower log: https://docs.google.com/spreadsheets/d/110xWWWfh8if-Hx16oX-qpALVk6zBa4JkS_xvpjKA9e4/edit?usp=sharing

IP Address of Enclosure power distribution unit: 192.168.1.205

IP Address of Enclosure Moxa serial device server: 192.168.1.204

How Often: 		Only for troubleshooting purposes or on direction from PI.

Clean requirements:    None

Supplies needed: 	None

Introduction:

The ICECAPS-ACE Enclosure, mounted at head height on the tower, controls the power distribution to each instrument, receives the data and passes it to the wireless relay that transmits the data to the local network. Under normal operating conditions it is not necessary to access the enclosure. The MOXA serial device server and the TrippLite power distribution unit are both connected to the network and can be accessed remotely from any browser on station at the IP addresses indicated above. The power consumption of the fluxtower suite can be monitored remotely if necessary. This SOP describes the contents of the ICECAPS-ACE enclosure for troubleshooting purposes only - again, the enclosure should not be accessed without permission from the PI.

Procedure:

*** ONLY CHECKING THE FLUXTOWER POWER CONSUMPTION CAN BE DONE WITHOUT APPROVAL OF PI ***

Checking the fluxtower power consumption.

Access the Tripplite PDU remotely at the following address: 192.168.1.205

Log in using the localadmin for both username and password.

You should see the following screen:

The input voltage should be ~120V (we are only using the secondary power input, input voltage 2). The ‘Output Current’ shows the total Amperage drawn by the flux tower suite. We expect this to be around 4-6A under normal operating conditions, it may increase in cold weather as several of the instruments are heated. It should not get above 11A.

Powering down the fluxtower enclosure:

The power to the entire flux tower suite and to individual power outlets can be controlled remotely using the TrippLite software. The screen below allows you to remotely turn off individual loads (1-16) or the entire system (Main Load). Note: Do not turn off the Main Load. If this is done accidently you will need to manually disconnect and reconnect the power from the ICECAPS-ACE Enclosure.

The individual loads are as follows:

| Load # | Description |
| --- | --- |
| 1 | KT15 instrument heater. Enclosure heater (disconnected). |
| 10 | Moxa. Do not turn off. |
| 12 | 24V supply: Metek2, Ventus2, HMP3, HMP4 and KT15 instruments. |
| 14 | 24V supply: Metek 1, Ventus1, HMP1, HMP2 and LiCOR instruments. |
| 15 | Network switch. Do not turn off. |
| 16 | 12V supply: All aspirated shields. Snow depth sensor. |

To completely power down the enclosure under consultation with PI:

- Power down the individual loads using the Tripplite software one at a time, apart from the moxa and the network switch (the ones that specify ‘Do not turn off’).

- At the fluxtower enclosure, physically disconnect the main power cable between the fluxtower enclosure and the main tower power distribution box.

To power back up:

- Reconnect the main power cable from the fluxtower enclosure to the main tower power distribution box.

- After a few seconds you should be able to access the Trippplite software again from the computer. Restart each of the instrument loads one at a time using the green state buttons on the ‘loads’ page. Wait 10-15s between powering on each load.

- Ensure data are logging per the main fluxtower protocol.

Inside the Enclosure:

SIMBA

Again the Enclosure should not be accessed without direction from the PI. The following illustrations are intended to be used as reference when following instructions from the PI. The photo below shows the underside of the Enclosure where all inlets/outlets are located. Each inlet/outlet is labelled. If a label is lost use this photo to determine the correct name and re-label. Photos of the inside of the enclosure follow.

Underside of enclosure - sockets to instruments

As of April 2020 there are now four power supplies in the lower part of the enclosure:

# ICECAPS-ACE - HMP (T/RH sensors in aspirated shields) – AWO

Contact:	Ryan Neely, ryan.neely@ncas.ac.uk

Heather Guy, heather.guy@ncas.ac.uk

Location of fluxtower log: https://docs.google.com/spreadsheets/d/110xWWWfh8if-Hx16oX-qpALVk6zBa4JkS_xvpjKA9e4/edit?usp=sharing

Address of fluxtower quicklook page: https://192.168.1.211/index_fluxtower.html

How Often: 		Daily: Daily checks

Clean requirements:    None

Supplies needed: 	Long-handled smooth bristled brush.

Introduction:

The ‘HMP’ is a Vaisala HMP155 humidity and temperature probe with heating. There are four of these instruments on the fluxtower. Instruments located at 2m, 4m, 8m and 15m are referred to as HMP1, HMP2, HMP3 and HMP4 respectively. The HMP’s are installed inside custom built aspirated shields. The aspirated shields suck ambient air up through the bottom and past the sensor head to prevent solar heating affecting the temperature measurements.

Cable to HMP instrument that sits inside the shield

Power for aspirated shield fan

Both the top and bottom sections of the aspirated shield can be removed whilst the instrument remains in place. The fan is housed in the top section.

Procedure:

Daily checks

- Check for icing: The aspirated shields can suck in fog and are prone to icing, particularly underneath where the air is pulled in. This ice should be removed where possible and reported on the icing page of the fluxtower log. Ice can be removed by  using a long handled soft-bristled brush. It may not be possible to remove the rime from the higher shields without climbing the tower. If it is not possible to remove the ice, record this in the icing log. It is not required to climb the tower to remove ice from the aspirated shields unless instructed by a PI.

- Check data: Weekly temperature and humidity measurements each HMP are plotted on the fluxtower quicklook page, accessible anywhere on station at: http://192.168.1.211/index_fluxtower.html

Complete the Data logging daily check in the fluxtower log. Issues with data should be logged, and reported to the PI’s if:

  - The temperature/ relative humidity plots are not up to date.

  - The plotted temperature/ relative humidity values appear unusual or unrealistic for any sensor (e.g. data gaps, sudden spikes, individual sensors differing greatly from others).

  - The technician has any reason to suspect bad data.

Power down:

In the event of a power outage, the HMP’s should automatically resume data collection as soon as power is restored (providing the raspberry pi is not powered down). If a manual reboot is deemed necessary, the HMP’s can be powered down simply by disconnecting the cable from the underside of the ICECAPS-ACE Enclosure. The ports are labelled HMP1-5. Once disconnected, wait at least 10 seconds before reconnecting.

Power up:

If the HMP is disconnected and then reconnected, data collection should automatically resume.

Data file access and locations:

- Raw data from the HMPs are saved in hourly files on the Fluxtower raspberry pi in the following directory:

  - ace-pir@192.168.1.211:/home/data/

- Hourly files have the following naming convention:

  - YYMMDD_HHMMSS.ss.HMP1

  - YYMMDD_HHMMSS.ss.HMP2

  - YYMMDD_HHMMSS.ss.HMP3

  - YYMMDD_HHMMSS.ss.HMP4

- Within each file, the data structure is as follows:

YYYY MM DD HH MM SS.sss RH Ta Tw ERR STAT

  - RH = relative humidity (%)

  - Ta = Temperature ©

  - Tw = Wet bulb temperature ©

  - ERR = error codes, STAT = heating status

Consumable Resupply Items:

N/A

Retro:

N/A

# ICECAPS-ACE - KT15 (Infrared radiation thermometer) – AWO

Contact:	Ryan Neely, ryan.neely@ncas.ac.uk

Heather Guy, heather.guy@ncas.ac.uk

Location of fluxtower log: https://docs.google.com/spreadsheets/d/110xWWWfh8if-Hx16oX-qpALVk6zBa4JkS_xvpjKA9e4/edit?usp=sharing

Address of fluxtower quicklook page: https://192.168.1.211/index_fluxtower.html

How Often: 		Daily: Daily checks

Clean requirements:     None

Supplies needed: 	None

Introduction:

The KT15 II is an infrared radiation thermometer by Heitronics. This instrument measures snow surface temperature and has been specially designed to detect temperatures as low as -70C. Snow surface temperature measurements can be used to estimate surface albedo and, combined with other flux tower and radiation measurements, will be used to make turbulent heat flux and surface energy budget calculations. The thermometer is protected by a ‘weather-proof’ casing with internal heating. Walking in front through the field of view of the infrared thermometer should be avoided where possible as it will impact the data, however it is perfectly safe to do so. If walking in front of the thermometer cannot be avoided, it should be noted in the flux tower log.

Procedure:

Daily checks

The KT15 is designed to be low maintenance. It has internal heating and should not ice. If ice is observed on the window of the KT15 enclosure, it should be brushed off and a note made in the logs; the PI’s should be contacted and can advise on troubleshooting the enclosure heating. Sometimes during blowing snow events ice builds up and obscures the instrument window. This should be noted in the log and removed as soon as possible. If it can’t be removed easily with a brush, it can be left to sublimate.

- Check data: Weekly snow surface temperatures from the KT are plotted as on the flux tower quicklook page.

- Complete the Data logging daily check in the flux tower log. Issues with data logging should be logged, and reported to the PI’s if:

  - The snow surface temperature plot is not up to date.

  - The plotted snow surface temperature values appear unusual or unrealistic, this might include sudden drastic changes from the usual daily temperature cycle or data gaps.

  - The technician has any reason to suspect bad data.

- During the annual instrument raise, keep the angle of this device in a 45 degree orientation, only unscrew the mounting plate that allow vertical relocation (never the angle adjusting screws).

Power down:

In the event of a power outage, the KT should automatically resume data collection as soon as power is restored. If a manual reboot is deemed necessary, the KT can simply be unplugged from the ICECAPS-ACE Enclosure. There are two ports for the KT on the ICECAPS-ACE enclosure, one labelled ‘KT15’ - this is the power and data transfer for the instrument itself. The other port is labelled ‘KT H’ - this is the power for the KT enclosure heater. For a manual reboot of the instrument, only the port/cable labelled ‘KT15’ should be disconnected. Once disconnected,wait at least 10 seconds before reconnecting.

Power up:

If the KT is disconnected from the ICECAPS-ACE Enclosure and then reconnected, data collection should automatically resume.

Data file access and locations:

- Raw data from the KT15 are saved in hourly files on the Fluxtower laptop in the following directory:

  - ace-pi@192.168.1.211:/home/data/

- Hourly files have the following naming convention:

  - YYMMDD_HHMMSS.ss.KT15

- Within each file, the data structure is as follows:

  - YYYY MM DD HH MM SS.sss TT.tt C

Where TT.tt is the snow surface temperature and C is the unit (celsius)

Consumable Resupply Items:

N/A

Retro:

N/A

# ICECAPS-ACE - LICOR (LI-7500 CO2/H2O Analyzer) – AWO

Contact:	Ryan Neely, ryan.neely@ncas.ac.uk

Heather Guy, heather.guy@ncas.ac.uk

Location of fluxtower log: https://docs.google.com/spreadsheets/d/110xWWWfh8if-Hx16oX-qpALVk6zBa4JkS_xvpjKA9e4/edit?usp=sharing

Address of fluxtower quicklook page: https://192.168.1.211/index_fluxtower.html

How Often: 		Daily: Daily checks

Weekly: Clean Licor lens

Clean requirements:   Replace Licor cleaning chemicals at request from PI.

Supplies needed: 	Lens cleaning wipes. Licor cleaning chemicals: Soda lime, magnesium perchlorate. All Licor cleaning supplies are stored in the small ICECAPS-ACE zarges in AWO.

Introduction:

The Licor LI-7500 is an open path infrared H2O/CO2 gas analyzer designed for use in eddy covariance flux measurement systems and measures H2O and CO2 in the free atmosphere at high resolution (10Hz). These measurements will be combined with other flux tower measurements to calculate latent heat flux, necessary to estimate surface energy balance. The sensor head is mounted on the ‘2m’ fluxtower boom, next to the 3D sonic. The Licor has it’s own separate electronics enclosure attached to the main body of the tower, the box is labelled LI-7500.

Procedure:

Daily checks

- Check for icing:

  - Rime ice will form on the Licor lens. There are two lens’ (top and bottom). Ice should be removed from the len’s daily. First try knocking or blowing off the ice, if this doesn’t work, ice can be removed using the lens cleaning wipes stored in the ICECAPS-ACE zarges in AWO.

  - Take care not to use anything that might scratch the lens.

  - Note cleaning times in the fluxtower logs - ‘blowing’ ice will effect H2O/CO2 measurements.

- Check data:

  - Daily CO2 and H2O measurements from the Licor are plotted as on the fluxtower quicklook page.

  - Complete the Data logging daily check in the fluxtower log.

  - The Licor will not collect good data in fog or if the lens is rimed, these will result in large spikes in the CO2 and H2O data and it is not necessary to contact the PI in this case. However suspected times of bad data due to fog/rime should be recorded in the notes section of the data-logging daily check in the fluxtower log.

  - Data issues should be reported to the PI’s if:

    - The the Licor plot is not up to date.

    - Unusual or unrealistic data is observed in the absence of fog or riming conditions (data spikes, data gaps)

    - The technician has any reason to suspect bad data that is not due to rime/fog.

Weekly lens clean:

- Even in the absence of rime ice, the Licor lens’ should be cleaned at least once per week using the lens cleaning wipes stored in the ICECAPS-ACE zarges in AWO. If the lens cleaning wipes have already been used to remove ice, it’s not necessary to repeat the procedure.

Power down:

In the event of a power outage, the Licor should automatically resume data collection as soon as power is restored. If a manual reboot is deemed necessary, the Licor can simply be unplugged from the ICECAPS-ACE Enclosure. There are two ports for the Licor on the ICECAPS-ACE enclosure, one labelled ‘LICOR’ - this is the data transfer the instrument itself. The other port is labelled ‘LI POWER’ - and this is the instrument power. For a manual reboot of the instrument, only the port/cable labelled ‘LI POWER’ should be disconnected. Once disconnected, wait at least 10 seconds before reconnecting.

Power up:

If the Licor is disconnected from the ICECAPS-ACE Enclosure and then reconnected, data collection should automatically resume.

Replace Licor cleaning chemicals:

The Licor uses Soda lime and magnesium perchlorate to keep the sensor housing free of H2O and CO2. The chemicals expire and the PI may request that they are changed, only change the chemicals at a request from the PI. Replacement Soda Lime and Magnesium perchlorate are stored in labelled containers in the small ICECAPS-ACE zarges in AWO. There is also a container for used Soda Lime and Magnesium perchlorate (that can safely be combined). Detailed instructions for replacing the chemicals are given on pages 96-100 of the LI-7500 manual. When disconnecting to the Licor sensor head, note that the cable pulls out straight - DO NOT TWIST.

Data file access and locations:

- Raw data from the LiCOR are saved in hourly files on the Fluxtower raspberry pi in the following directory:

  - ace-pi@192.168.1.211:/home/data/

- Hourly files have the following naming convention:

  - YYMMDD_HHMMSS.ss.licor

- Within each file, the data structure is as follows:

  - YYYY MM DD HH MM SS.sss Ndx Diag CO2R CO2D H2OR H2OD T P C

Ndx: Index value. Boolean controlled output. Increments every 6.5 ms (1/ 152). Resets to 0 whenever the configuration changes, or (with no changes) about every 163 days.

Diag:  Diagnostic value 0-255. Boolean controlled output.

CO2R: Raw  CO2 absorptance

CO2D:  CO2 density (mmol m-3)

H2OR: Raw H2O absorptance

H2OD: H2O density (mmol m-3)

T: temperature, degreesC

P: pressure kPa

Cooler:  Detector cooler (Volts)

Consumable Resupply Items:

Lens cleaning wipes.

Licor cleaning chemicals (Soda lime, magnesium perchlorate) - a resupply is not anticipated.

Retro:

Used Licor cleaning chemicals, to be stored in the dedicated labelled container in the small ICECAPS-ACE zarges in AWO.

# ICECAPS-ACE - Metek (3D sonic anemometer) – AWO

Contact:	Ryan Neely, ryan.neely@ncas.ac.uk

Heather Guy, heather.guy@ncas.ac.uk

Location of fluxtower log: https://docs.google.com/spreadsheets/d/110xWWWfh8if-Hx16oX-qpALVk6zBa4JkS_xvpjKA9e4/edit?usp=sharing

Address of fluxtower quicklook page: https://192.168.1.211/index_fluxtower.html

How Often: 		Daily: Daily checks

As needed: Recalibration

Clean requirements:    None

Supplies needed: 	Metek calibration kit

Introduction:

The ‘Metek’ is a uSonic-3 Scientific 3D Ultrasonic Anemometer made by Metek. There are two of these instruments on the fluxtower, the instrument on the ‘2m’ beam is referred to as ‘M1’ and the instrument on the ‘15m’ beam is referred to as ‘M2’. These instruments record 3D wind speed at high resolution (10Hz) for the calculation of turbulent momentum and heat flux required to calculate the surface energy budget. The sensor heads are heated and should not ice.

Procedure:

Daily checks

The Metek has internal heating and should not ice. If ice is observed on either Metek, the PI’s should be contacted and can advise on troubleshooting.

- Check data: Daily wind measurements from M1 and M2 are plotted on the fluxtower quicklook page.

- Complete the Data logging daily check in the fluxtower log. Issues with data should be logged, and reported to the PI’s if:

  - The 3D sonic winds plot is not up to date.

  - The plotted wind speeds appear unusual or unrealistic for either sensor, e.g. data gaps.

  - The technician has any reason to suspect bad data.

  - Note: Wind directions are aligned to Magnetic north as of 22 May 2019 and will need adjustment during data processing to align with true north.

Power down:

In the event of a power outage, the Metek should automatically resume data collection as soon as power is restored (providing the fluxtower raspberry pi is not powered down). If a manual reboot is deemed necessary, the Metek can be powered down by disconnecting the cable extending from the Metek to the junction box (see photo). For M2, this will require climbing the Met tower.

Once disconnected,wait at least 10 seconds before reconnecting.

Power up:

If the Metek is disconnected and then reconnected, data collection should automatically resume.

Recalibration:

In the event of continuous bad data, the PI might advise a recalibration of the instrument. In the event that a recalibration is required, the PI will provide a separate SOP.

Data file access and locations:

- Raw data from the Metek are saved in hourly files on the Fluxtower raspberry pi in the following directory:

  - ace-pi@192.168.1.211:/home/data/

- Hourly files have the following naming convention:

  - YYMMDD_HHMMSS.ss.m1

  - YYMMDD_HHMMSS.ss.m2

- Within each file, the data structure is as follows:

YYYY MM DD HH MM SS.sss X:x =    xx y =    yy z =    zz t =  tttt

  - X = Heater status (M=off, H=on, D=heater defect)

  - x,y,z components of wind in cm/s

  - t = acoustic temperature 2357 = 23.57C

Consumable Resupply Items:

N/A

Retro:

N/A

# ICECAPS-ACE - SnD (Sonic distance sensor) – AWO

Contact:	Ryan Neely, ryan.neely@ncas.ac.uk

Heather Guy, heather.guy@ncas.ac.uk

Location of fluxtower log: https://docs.google.com/spreadsheets/d/110xWWWfh8if-Hx16oX-qpALVk6zBa4JkS_xvpjKA9e4/edit?usp=sharing

Address of fluxtower quicklook page: https://192.168.1.211/index_fluxtower.html

How Often: 		Daily: Daily checks

Clean requirements:     None

Supplies needed: 	None

Introduction:

The SnD is a Campbell Scientific sonic distance sensor (model: SR50AH). The sensor is mounted looking vertically down at the snow and sends out a sonic pulse every 10 minutes, measuring the distance from the sensor to the snow surface - indicating snow depth change. The beam angle of view is 30 degrees from the sensor and snow within this range should be kept undisturbed. Changes in snow depth can be used to indicate heavy snowfall/ ablation and to calculate accurate distances between other fluxtower instruments and the snow surface. The sensor has a heated housing and is not expected to ice.

Procedure:

Daily checks

The SnD is expected to be low maintenance and it is not necessary to approach the instrument. The snow below and around this instrument should be kept undisturbed. It has internal heating and should not ice. If ice is observed, the PI’s should be contacted and can advise on troubleshooting the instrument heating.

- Check data: Weekly snow depth measurements from the SnD are plotted as on the fluxtower quicklook page. Measurements are the distance from the sensor to the snow surface. This instrument is notoriously finicky, check in with Heather every time you notice inconsistent data (it won’t update for several hours) and log it in the daily checks. Commonly a weather derived issue.

- Complete the Data logging daily check in the fluxtower log. Issues with data logging should be logged, and reported to the PI’s if:

  - The snow surface distance plot is not up to date.

  - The plotted snow surface distance values appear unusual or unrealistic, this might include sudden drastic changes or data gaps.

  - The technician has any reason to suspect bad data.

  - The snow surface is approaching 50 cm from the sensor. 50 cm is the measurement limit for the sensor and in this case the sensor requires raising.

Power down:

In the event of a power outage, the SnD should automatically resume data collection as soon as power is restored (providing the fluxtower laptop is not powered down). If a manual reboot is deemed necessary, the SnD can simply be unplugged from the ICECAPS-ACE Enclosure. There are two ports for the SnD on the ICECAPS-ACE enclosure, one labelled ‘SND’ - this is the power and data transfer for the instrument itself. The other port is labelled ‘SND H’ - this is the power for the SnD enclosure heater. For a manual reboot of the instrument, only the port/cable labelled ‘SND’ should be disconnected. Once disconnected,wait at least 10 seconds before reconnecting.

Power up:

If the SnD is disconnected from the ICECAPS-ACE Enclosure and then reconnected, data collection should automatically resume.

Data file access and locations:

- Raw data from the SnD are saved in hourly files on the Fluxtower laptop in the following directory:

  - ace-pir@192.168.1.211:/home/data/

- Hourly files have the following naming convention:

  - YYMMDD_HHMMSS.ss.SnD

- Within each file, the data structure is as follows:

YYYY MM DD HH MM SS.sss aa;D.DDD;QQQ; VVVVV;CC

  - aa = 33 (serial address of sensor)

  - D.DDD = Distance to target in m (will need temperature adjustment)

  - QQQ = Data quality, varies beteen 152-600, 600 is the poorest quality

  - VVVVV = diagnostic tests (only first two are actually something), 1 = pass.

  - CC = two-character checksum of data packet.

Consumable Resupply Items:

N/A

Retro:

N/A

# ICECAPS-ACE - Ventus (2D sonic anemometer) – AWO

Contact:	Ryan Neely, ryan.neely@ncas.ac.uk

Heather Guy, heather.guy@ncas.ac.uk

Location of fluxtower log: https://docs.google.com/spreadsheets/d/110xWWWfh8if-Hx16oX-qpALVk6zBa4JkS_xvpjKA9e4/edit?usp=sharing

Address of fluxtower quicklook page: https://192.168.1.211/index_fluxtower.html

How Often: 		Daily: Daily checks

Clean requirements:    None

Supplies needed: 	None

Introduction:

The Ventus-X is a 2D ultrasonic wind sensor by Lufft. There of two of these instruments on the fluxtower, the instrument on the 4m beam is referred to as ‘V1’ and the instrument on the 8m beam is referred to as ‘V2’. The Ventus-X measures wind speed and direction at a high resolution. These sensors are heated and should not ice.

Procedure:

Daily checks

The Ventus-X is designed to be low maintenance. It has internal heating and should not ice. If ice is observed on either Ventus, the PI’s should be contacted and can advise on troubleshooting.

- Check data: Daily wind measurements from V1 and V2 are plotted together on the fluxtower quicklook page.

- Complete the Data logging daily check in the fluxtower log. Issues with data logging should be logged, and reported to the PI’s if:

  - The 2D sonic winds plot is not up to date.

  - The plotted wind speed or direction values appear unusual or unrealistic for either sensor, e.g. data gaps.

  - The technician has any reason to suspect bad data.

  - Note: Wind directions are aligned to Magnetic north as of 22 May 2019 and will need adjustment in data processing to align with true north.

Power down:

In the event of a power outage, the Ventus should automatically resume data collection as soon as power is restored. If a manual reboot is deemed necessary, the Ventus can be powered down by disconnecting the cable extending from the Ventus to the junction box (see photo). This will require climbing the Met tower.

Once disconnected,wait at least 10 seconds before reconnecting.

Power up:

If the Ventus is disconnected and then reconnected, data collection should automatically resume.

Data file access and locations:

- Raw data from the Ventus are saved in hourly files on the Fluxtower laptop in the following directory:

  - ace-pi@192.168.1.211:/home/data/

- Hourly files have the following naming convention:

  - YYMMDD_HHMMSS.ss.v1

  - YYMMDD_HHMMSS.ss.v2

- Within each file, the data structure is as follows:

YYYY MM DD HH MM SS.sss SS.S DDD +TT.T xx*XX

  - SS.S = wind speed (m/s)

  - DDD = wind direction

  - +TT.T = signed virtual temperature

  - xx = status

  - XX = checksum

Consumable Resupply Items:

N/A

Retro:

N/A

# ICECAPS-ACE – Water-Based Condensation Particle Counter (CPC) – AWO

Primary contact:  	Betsy Andrews, NOAA GML, Boulder, CO

(303) 497-5171, Betsy.Andrews@noaa.gov

Secondary contact:  	Patrick Sheridan, NOAA GML, Boulder, CO

(303) 497-6672, Patrick.Sheridan@noaa.gov

Software contact:	Derek Hageman, NOAA GML, Boulder, CO

(303) 497-4652, Derek.Hageman@noaa.gov

How Often:		Daily: check that instrument is on, and talking to aerosol computer, check water level in bottle and fill if water is low

Weekly: fill water bottle. (Note: Please use the larger water bottle of the two that were shipped with the instrument, and be sure to keep it full, as air bubbles can get trapped in the instrument and interfere with measurements if the water bottle runs dry).

Clean Requirements:		None

Supplies Needed:		Distilled water from grocery store or milliQ water or Ultra-pure water (just not regular tap water!)

Science Introduction:

The concentration of particles in the atmosphere contributes to our knowledge about the spatial and temporal scales of aerosols.  Aerosol particles, along with other factors (relative humidity, aerosol chemistry, etc.), contribute to climate forcing.  The data acquired from the CNC are used as aerosol parameters for global models and ground-truth for satellite measurements.

The Moderated Aerosol Growth with Internal water Cycling (MAGIC) instrument draws sample air into a heated humid chamber where the air becomes saturated with water.  The air stream is then cooled and becomes super-saturated, causing condensation of the water on aerosol particles in the sample stream. This condensation increases the size of the particles, making them detectable through optical methods.

Procedure:

Daily:

- Make sure the numbers are changing on the LCD screen on the front of the instrument.

- Make sure that CPD3 is open on the laptop that is connected to the 8-port Edgeport box.  CPD3 will be a blue window with a series of numbers displayed in columns.

- After you have selected the CPD3 window so that it is the active window, hit <Enter>, and a menu will pop up.  Scroll down to “MAGIC CPC” using the arrow keys on the keyboard.  Once “MAGIC CPC” is highlighted, hit <Enter>.

- Make sure the numbers are changing in the MAGIC CPC window on the laptop blue screen and that the concentration is not 0.   Check MAGIC dewpoint temperature at its inlet by looking at the MAGIC CPC window on laptop - the dewpoint value should be greater than 10°C (for example see value of dewpoint in humidity section of picture below - in the photo the dewpoint is 16.9°C). Notify the primary contacts above if dew point is too low.

Weekly:

- Refill water bottle (it is very important that it does not go empty). Remove the cap on the bottle and top off with distilled, milliq or HPLC water.  Do not disconnect bottle from humidifier while refilling.  Replace bottle cap when done.

- If water bottle is ever found empty or measurements look supsicious, air bubbles will need to be purged from the instrument following the procedure outlined in the manual. Air bubbles can get trapped if the water bottle ever runs dry, or if dissolved gases in the water gradually accumulate in the humidifier housing.

Notes:

- Perform main inlet inspection as described in the protocol for the nephelometer.

Consumable Resupply Items:

Ultra-pure water or distilled water

Shutdown Procedures:

The CPC, nephelometer and CLAP power supply as well as the computer are powered by the UPS. The pump runs on a different non-UPS circuit. In the event of a planned extended power outage, notify the PIs beforehand. Add a note to the log and turn off power to the CPC, nephelometer, and CLAP. Shutdown for the CPC instrument is performed by using the on/off switch on the back of the instrument. Alternatively, the power supply connector can be detached from the back of the CPC or the instrument power supply can be unplugged from power strip/wall outlet). That way no invalid data are collected. On startup, the laptop will re-initialize communications with the instruments if they have power and are connected to the 8 port serial hub.

Emergency shutdown:

Turn off power to the CPC, nephelometer and CLAP at the power source. Shut down the laptop computer. Unplug pump. Turn off power to any other instrument components.

# ICECAPS – Significant Weather Observations

Primary contact: 	Von Walden, v.walden@wsu.edu

Secondary contact:	Matthew Shupe, matthew.shupe@noaa.gov

How often:		As appropriate: Observation and uploading of photos

Clean requirements:		None

Supplies needed:		Camera with UTC clock setting

Introduction:

The primary goal of significant weather observations is to document unusual atmospheric optics, including (but not limited to!) sun optics such as halos, sun dogs/parhelia, parhelic circles, columns or pillars, fogbows, fata morgana, along with any unusual clouds or ground events such as yukimarimo formation. It is unlikely that a comprehensive and statistically representative sample dataset of these phenomena at Summit can be collected. The focus, therefore, will be on documenting different events. Significant weather data can be a valuable tool in setting the context of remote sensing measurements made continuously by ICECAPS.

For example, these data can be used as an interpretation tool in the analysis of particle orientation as seen by ICECAPS lidars during precipitation events. This will help span the gap between the lidar measurement and a given set of meteorological conditions. Also, a longer term dataset of events considered to be of a similar type (such as diamond dust) can be analyzed with reference to radiance measurements by the ICECAPS Polar Atmospheric Emitted Radiance Interferometer (PAERI) to find the radiative effect of such events.

The following website provides excellent visuals and explanations of atmospheric optics: http://www.atoptics.co.uk/. It is bookmarked on the ICECAPS PC computer. There are also cloud classification books and posters around station.

Procedure:

- Whenever a significant weather observation is observed, note the time and length of the event. Try to record the time of the entire phenomena rather than just witness it once and forget about it.

- Take a photo of the event. If the sun is very bright and part of the optic, block it out with either your hand or something around station (flags work well).

- Record the event and time (in UTC) on the ICECAPS weekly report under “Significant Weather Observations.”

- Download the photos from the camera to a tech computer.

- Prepare files for uploading.

  - Name each photo in the following format: YYYYMMDDhhmmZ.NameOfOptic . If there is more than one photo for each event, continue with NameOfOptic1, NameOfOptic2, NameOfOptic3, ….

  - The file names CANNOT have spaces in them or Dataman will hang indefinitely when it tries to copy them to the ToBoulder folder.

  - If you feel extra descriptions or explanations are necessary, create a *.txt file the each sampling session in the folder with the photos. Include in the file:

    - The time of photograph and duration of phenomena

    - Any other comments or notes about the event

- Transfer the entire sampling session (individual photos and .txt file, not in a sub-folder) to the ICECAPS Linux computer (“Dataman”):

  - Open the WinSCP FTP client.

  - Select the existing profile for ‘dataman’ and click ‘Login’. If the profile dotes not exist, create a new profile:

    - File Protocol: SFTP

    - Host name: 192.168.0.73

    - Port: 22

    - Username: dataman

    - Password: see KeePass password manager

    - Save this profile to facilitate future access.

  - With the connection established, drag-and-drop the folder you’d prepared with IcePic data into this directory:

    - home/dataman/search/work/weeklyphotos

Retro:

None

Consumable Resupply Items:

None

# ICECAPS – Radiosonde Campaign

Primary contact: 	Matt Shupe, matthew.shupe@noaa.gov

Secondary contact:	Erik Olsen, eriko@ ssec.wisc.edu; Von Walden, v.walden@wsu.edu

How often:	Daily: Two launches per day corresponding to 00:00 and 12:00Z between April and September. One launch per day corresponding to 12:00Z starting September 2024.

Weekly: Check computer time, reboot computer

Each Phase: Change plastic on balloon table

Clean requirements:	None

Supplies needed:	Vaisala RS41 radiosonde

Helium (~0.2 cyl/launch)

Cotton string

Regulator and filler neck, with 500g weight

Scissors

Gloves to prevent powder chemical on balloon from irritating hands

Introduction:

Radiosondes are released twice per day and collect atmospheric profiles of temperature, humidity, wind speed, and pressure. These data are used in the ICECAPS measurement campaign and sent to the Global Telecommunications System (GTS) for incorporation into daily forecast models.

Data files from the radiosonde launches are transferred to the GTS by the cron scripts running on the ICECAPS Linux server; these scripts also archive the files locally and transfer them to servers in Boulder.

Radiosonde Launch Schedule in Local Times

| Launch | Local Standard Time (WGT) |  | Local Daylight Time (WGST) 3/31/2024 – 10/27/2024 |  |
| --- | --- | --- | --- | --- |
|  | Preparation Start | Launch Window | Preparation Start | Launch Window |
| 12Z (AM) | 08:45 | 09:15 – 09:30 | 09:45 | 10:15 – 10:30 |
| 24Z (PM) | 20:45 | 21:15 – 21:30 | 21:45 | 22:15 – 22:30 |

The timing of radiosonde launches is important; these radiosonde launches are being coordinated with others around the world that are launched near the “synoptic” times (12Z, 24Z).  The target launch time is 30-45 minutes prior to these times such that the sonde is mid-flight at 12Z or 24Z.  Messages are sent automatically from the Vaisala software upon the bursting of the balloon or loss of signal.

If a radiosonde is released outside of this window, the data will not be transferred to the GTS. However, even in this case the data is still incredibly valuable to the ICECAPS campaign and launches should be attempted even if they are outside of the specified window. This may be the case in the event that a hardware error delays the launch attempt or a hardware failure occurs during an initial launch attempt. Additional considerations for failures resulting from high wind conditions are detailed in the protocol.

Starting in Summer 2023, the ICECAPS radiosonde program transitioned from using 350g weather balloons to 150g ballons in an effort to reduce helium consumption. 150g balloons have proven to carry the radiosondes well through the troposphere where measurements are particularly important for diagnosing the effects of water vapor, with an average burst altitude calculated between XX and XX of ~20 km. However, 350g balloons are still launched periodically to collect measurements from higher in the stratosphere. For every 10 radiosonde releases, one 350g balloon is used.

Procedure:

- Preparation

While either the balloon or the sonde preparation can be done first; it is preferable that both the balloon and sonde have 10-15 minutes to rest and equilibrate prior to the launch.

  - Sonde prep

    - Check the computer time before starting a new sonde. It should be synced to the MSF time server. See “Weekly Checks” below for detailed instructions.

    - Open the MW41 software and log in to the admin account. Password credentials are on the password sheet.

    - Unpack the sonde being careful not to touch the sensors.

    - Place the sonde on the GC25 ground check station. There are no hard connections to make; the RS41 sondes are wireless.

    - Once open, the software will connect to the sonde through the ground check station and reconditioning will automatically start. During this time temperature, humidity, and pressure sensors are calibrated on the sonde. This process takes 5 minutes, during which time you may wish to start on balloon preparations, below.

    - When the reconditioning is complete, log the calibration humidity.

    - Remove the sonde from the ground check station. Insert the unwinder to the top of the sonde. Note that this pushes the sensors forward, ensuring the sensors will not come in contact with the sonde’s tether while in flight. Do not touch or breathe on the sensors.

    - Carry the sonde outside and set it on the red stool to equilibrate. Ideal placement is with the sensors, shaded from sun, aspirated by wind, away from heavy equipment danger, away from exhaust plumes, and in a location away from obstacles where it can be picked up while carrying the balloon.

    - After 5-10 minutes, when measurements have equilibrated to outside conditions, populate the log file with the sonde outdoor measurements and those from NOAA weather station:

      - Compare and record the sonde’s RH and temperature readouts with the AWO met tower’s RH and temperatures. This ensures accurate readings by the sonde. If the sonde RH/temperature is unusually erratic (changing more than 15% RH or 5 deg C) or differ from the AWO readings by more than 15% RH, start over with a new sonde.

Note:

        - When environmental temperatures fall below -60C, the NOAA aspirated humidity sensor hits its measurement limit and will often record -59.9 or -60 for dewpoint temperature value for as long as these temperature conditions persist. Notify PI when this occurs, note in the comments section of the log, and instead of using this incorrect dewpoint value, use the RH value from the equilibrated sonde itself for your surface observations.

        - When there is thick FZFG at the surface, rime might start to form on the sensor and result in high RH readings (> 100%) that may not agree with Met tower instruments. In this case, make a note in the log and continue with the launch. The RH should return to expected values soon after launch when the sonde leaves the surface fog layer.

      - Check the regularity of the measurements in the time series plots. If there are periods with missing data (i.e., sonde stops functioning or transmitting), and the problem continues, start over with a new sonde.

      - Check the strength of sonde’s signal on the signal window of the software; verify that there are no interfering signals in the same band.

      - Check the GPS output of the RS41 sondes via the GPS satellites window of the software.

    - If all values are acceptable, proceed with balloon preparations and release below. Do not bring the sonde back inside the SMG unless absolutely necessary due to high winds or safety issues. Note in the sonde log if so.

  - Balloon prep

    - During preparation and launch, make every effort to keep the body of the balloon as clean as possible:

      - Ensure that the balloon table is clean and covered with the blue tarp while not inflating. The white plastic should be replaced before each turnover.

      - Do not touch the balloon, especially with bare hands, except for the filler neck. If necessary, the balloon bag serves as a good “glove” to keep skin oils off the balloon.

      - Keep all sources of grease away from the balloon and table. Coordinate with other SMG users to ensure no equipment is running in the SMG while the balloon is being inflated or resting before launch.

      - Be careful not to bump the inflated balloon against anything in the SMG.

    - Fit the neck of the balloon completely to the base of filler nozzle; a weight should be suspended from the nozzle.

    - Turn on the cylinder valve, then dial the regulator up to 15 psi

    - Allow the balloon to fill until the weight just lifts from the table. Shut off the cylinder valve and open the regulator handle.

    - Tie off the balloon to both seal the balloon and to create an attachment point for the radiosonde.

    - Re-cover the balloon table, then allow the balloon to rest in the SMG for 10-15 minutes before launch.

*Be careful that the fill hose doesn’t get caught on the balloon table, if it gets hung up you will unintentionally fill the balloon more than necessary!

- Launch

  - Do not launch the balloon prior to the launch time in the table above.

  - Open the SMG garage door and walk with the balloon outside to the equilibration stand (a red stool).

  - Connect the payout reel hook completely into the loop created from the folded neck of the balloon.

  - Walk to a location 10 meters directly out from the equipment door of the SMG, aligned with the long axis of the building.

    - Try to keep the sonde equilibrated to the outside environment.

    - If the generator exhaust plume is low to the ground, avoid walking the sonde through it. Do not release the radiosonde through the generator plume; adjust the release location if necessary.

  - Check that the payout string is free to unspool and not tangled with yourself or any part of the radiosonde system. Lower the sonde close to the ground, and release the balloon allowing the payout reel to unspool. It is important to the PIs that the balloon be launched from approximately the same location (elevation included) each day: 10 meters out from the SMG bay door.

  - DECONFLICTING LAUNCHES WITH FLIGHT OPERATIONS – TO BE ADDED

  - HIGH WIND/EXTREME COLD LAUNCHES – TO BE UPDATED FOR BUDDY SYSTEM

    - If loading the sonde onto the balloon loop outdoors is impossible due to high winds or extreme cold:

      - At Step 2.c, bring the sonde back into the SMG and connect to the balloon at the table

      - Walk the balloon and sonde outside. As much as is reasonable given the conditions, try to allow the sonde re-equilibrate to the outside air before releasing the balloon

  - FINAL STEP: VERY IMPORTANT! Return to the radiosonde computer and enter the NOAA surface observations from the Met Tower. Click “Apply”. If this is not done, data will not be collected and sent to GTS!  See note above in section ix.1 if surface dewpoint temps at the time of launch are colder than -60c.

  - Check the sounding data to ensure that the sonde is ascending, that the measurements are reasonable, and that the data are being received without interruption. If the sonde fails in the first minutes of launch due to sensor issues, a second sonde should be launched (see Re-Launch below). It is better to have late data than no data at all. This should be rare—notify science PM of any failed soundings both to identify issues and to track ballooning material usage.

  - Enter the launch time into the electronic log, found at the end of the “Release” data line of the Sounding Status screen.

- Closeout

  - At about 50 to 90 minutes after release, the balloon will burst and the sonde’s measured pressure will start to increase. A data file message will automatically be generated. Once a minute the ICECAPS data management will look for this sort of file on the radiosonde computer and if found, it will copy it for backup and aggregation with ICECAPS data and send a copy to GTS. On the radiosonde computer, a window stating “Sounding Terminated” will fill the screen. Any time after termination, the closeout steps may be performed—this is usually when the next sonde is prepared about 12 hours later.

  - To close out. Click “OK” at the ‘Sounding Terminated’ alert.

  - Record the burst altitude and pressure (circled below in Figure 2) in the electronic log. Click “close sounding” to close the program.

Figure 2: Record burst values and closing out the sonde

  - Re-Launch

    - If the radiosonde or balloon fails early in a flight, the PIs would like a second launch to be performed.  Launches that fail in high-wind releases do not need to be repeated, due to the likelihood of further failures.

    - Initiate closing the current sounding using the “Terminate Sounding” button in the upper part of the window.

  - Click “Exit” to leave the current sounding; note the failure in the logs

  - Start from the beginning of the launch procedure, creating new log entries for the second launch

    - If the first, failed sonde is still alive and transmitting, you must adjust the transmit/receive frequency of the second sonde to reduce interference

    - Before starting the new sonde, in the “Administration” tab select “Sounding Processing Subsystem.” Do not adjust the frequency lower than 400.4 MHz or above 405.6 MHz. Click “Save.”

Accessing Past Launches

If for some reason a balloon flight needs to be ‘rerun,’ for example, surface observations were entered incorrectly or not at all, a simulation of a recent flight can be created. To do this, select “Archive” and “simulate.” You can then enter new surface observations and speed up the simulation so it is done in minutes rather than hours. This sends a new message to GTS and Dataman, so only run this simulation if need be, and notify both PIs.

If a launch was closed out before data was recorded by a tech, archives can easily be opened again by selecting “Archive,” checking the sonde you would like to analyze, and clicking “open.” Look though the messages to find data recorded, including burst height and surface observation values entered.

Weekly Checks

- Check the balloon table surface in the SMG. Clean the outer and inner plastic with diluted simple green and a clean rag. Replace plastic at least once a phase, or as needed.

- Restart the sonde desktop to keep systems updated and prevent crashes. Restart the SPS base station.

- Check the sonde desktop time

  - Open the date and time properties on the sonde desktop; then open the “Internet Time” tab.

  - Verify that desktop is setup to update using the MSF time server at 192.168.1.170.

  - Force the clock to update using “Update Now”

  - Verify that the time is reasonable

- Re-stock balloons, spoolers, and sondes as needed.

- Take out trash from balloon table can and mezzanine can.

- Confirm side fan is running. Upper fan is not plugged in to avoid negative pressure and ingress of grime.

Changing Out Air Filter:

- Every six months the air filter on the side of the cabinet should be changed.

- Prepare the filter and required tools at hand and ready for the swap well in advance of the swap.

- With the cabinet having been closed for at least an hour, record the time to start the work. Fine to leave the fan running.

- Perform the filter swap efficiently. You don’t need to rush the work, but we would like to minimize the time between the old filter being removed and the new filter being installed.

- Note the time at the end of the work.

- Leave the cabinet closed for an hour before opening it again.

Changing Out Cylinders

- When 25 empty cylinders have accumulated, arrange for a cylinder swap. This should be planned in advance to allow for deconfliction of the shop space and review of safety documents. Three people are ideal to do this task.  Review AHA first.

- Keep track of ICECAPS cylinder inventory at \PFS-ARC-Summit Techs - Documents\Inventories\Science_Inventory_Jan 2024.xlsx.

Troubleshooting – TO BE UPDATED

During the sonde preparation process, it could fail (see photo below). This error could be a bad sonde so first step is to take the current one off. Hold the power button down to turn off device. And open a new sonde and test it (test meaning put it on the conditioning device and wait to see if it fails). If it also fails, click file save in your launch log and exit out of everything, restart the computer and power cycling the SPS (press off button and let it sit for a minute then press button again). Log back into everything and check your second sonde again.

Recalculation Error: after launching, a “recalculation” error can occur at 4000-6000 meters. If this occurs, it will be necessary to relaunch a radiosonde. Often the recalibration error is associated with a “balloon burst” even if this is not actually the case. Because the sonde is still active, you will have to change the radio frequency of the new sonde (use the same procedure as noted above) and then launch a new sonde.

Planned Power Outage – TO BE UPDATED:

Sometimes station work requires a planned power outage; this will need to be coordinated with the science techs not only to send notice to the science community, but also to make sure that the radiosonde balloon launch has successfully terminated. Meaning balloon has burst and data has been transferred successfully. This can happen anytime between 60-90 minutes after the launch.

To confirm these steps have occurred, to give the okay for a power outage, see photos below:

Ascending icon (top right):

This indicates that the ballon is currently in flight; green arrow is the :

Generating Message icon:

This step is after the balloon has burst, but the data has NOT been sent yet, notice the balloon is shown as a red x (in the flight map):

Sounding Terminated icon:

Indicative of data sent off-site and terminated. At this stage, it’s ready to continue with other work: weekly computer reboot, planned power outage, next balloon launch, etc.

Retro:

He cylinders

Consumable Resupply Items:

Vaisala RS-41 radiosondes

Helium

150g and 350g balloons

Plastic sheeting for balloon table

# ICECAPS – Significant Weather Observations

Primary contact: 	Von Walden, v.walden@wsu.edu

Secondary contact:	Matthew Shupe, matthew.shupe@noaa.gov

How often:		As appropriate: Observation and uploading of photos

Clean requirements:		None

Supplies needed:		Camera with UTC clock setting

Introduction:

The primary goal of significant weather observations is to document unusual atmospheric optics, including (but not limited to!) sun optics such as halos, sun dogs/parhelia, parhelic circles, columns or pillars, fogbows, fata morgana, along with any unusual clouds or ground events such as yukimarimo formation. It is unlikely that a comprehensive and statistically representative sample dataset of these phenomena at Summit can be collected. The focus, therefore, will be on documenting different events. Significant weather data can be a valuable tool in setting the context of remote sensing measurements made continuously by ICECAPS.

For example, these data can be used as an interpretation tool in the analysis of particle orientation as seen by ICECAPS lidars during precipitation events. This will help span the gap between the lidar measurement and a given set of meteorological conditions. Also, a longer term dataset of events considered to be of a similar type (such as diamond dust) can be analyzed with reference to radiance measurements by the ICECAPS Polar Atmospheric Emitted Radiance Interferometer (PAERI) to find the radiative effect of such events.

The following website provides excellent visuals and explanations of atmospheric optics: http://www.atoptics.co.uk/. It is bookmarked on the ICECAPS PC computer. There are also cloud classification books and posters around station.

Procedure:

- Whenever a significant weather observation is observed, note the time and length of the event. Try to record the time of the entire phenomena rather than just witness it once and forget about it.

- Take a photo of the event. If the sun is very bright and part of the optic, block it out with either your hand or something around station (flags work well).

- Record the event and time (in UTC) on the ICECAPS weekly report under “Significant Weather Observations.”

- Download the photos from the camera to a tech computer.

- Prepare files for uploading.

  - Name each photo in the following format: YYYYMMDDhhmmZ.NameOfOptic . If there is more than one photo for each event, continue with NameOfOptic1, NameOfOptic2, NameOfOptic3, ….

  - The file names CANNOT have spaces in them or Dataman will hang indefinitely when it tries to copy them to the ToBoulder folder.

  - If you feel extra descriptions or explanations are necessary, create a *.txt file the each sampling session in the folder with the photos. Include in the file:

    - The time of photograph and duration of phenomena

    - Any other comments or notes about the event

- Transfer the entire sampling session (individual photos and .txt file, not in a sub-folder) to the ICECAPS Linux computer (“Dataman”):

  - Open the WinSCP FTP client.

  - Select the existing profile for ‘dataman’ and click ‘Login’. If the profile dotes not exist, create a new profile:

    - File Protocol: SFTP

    - Host name: 192.168.0.73

    - Port: 22

    - Username: dataman

    - Password: see KeePass password manager

    - Save this profile to facilitate future access.

  - With the connection established, drag-and-drop the folder you’d prepared with IcePic data into this directory:

    - home/dataman/search/work/weeklyphotos

Retro:

None

Consumable Resupply Items:

None

# NOAA – Carbon Cycle Greenhouse Gases (CCGG) Flasks – AWO

Primary contact:  	Don Neff, NOAA GML, Boulder, CO (sampler related questions)

303-497-4271 ccggflask@noaa.gov, don.neff@noaa.gov

NOAA Carbon Cycle Group, NOAA GML, Boulder, CO

Secondary contact:  	Christy Smith, NOAA GML, Boulder, CO

303-497-5397, gmd.met@noaa.gov

Shipping/Resupply:	Molly Crotwell, NOAA GML, Boulder CO

303-497-4728, ccggflask@noaa.gov,  molly.crotwell@noaa.gov

Eric Moglia, NOAA GML, Boulder CO

303-497-3988, ccggflask@noaa.gov

How Often:	Weekly

Clean Requirements:	Designated sampling site in the clean air sector upwind from AWO.

Sample ONLY when winds are from the clean air sector. If north winds are persistent and no alternative exists, attempt to make a weekly sample.

Supplies Needed:	Manual Portable Sampling Unit (PSU) flask case, 2 flasks, data sheet, pencil, watch

Introduction: The whole air samples in these glass flasks will be analyzed for trace gases that are important components of the global carbon cycle.  Summit is part of a network of sites around the world where samples of remote air are collected.  Your breath can ruin a sample, so imagine what a little heavy equipment or generator exhaust could do.

For more information: http://www.esrl.noaa.gov/gmd/ describes NOAA/GMD’s global measurements program and has past data available for download.  Full instructions for the Manual PSU are available at: http://www.esrl.noaa.gov/gmd/ccgg/psu/index.html.

Procedure:

Follow “Abbreviated PSU Air Sampling Instructions” below:

*When opening a new (unsampled) flask crate, it is helpful to remove some of the top packing foam and set it aside until all flasks are sampled and you are ready to repack the crate for shipment back to NOAA.

*Please keep flask storage crates closed at all times, except when actively loading and unloading flasks.

I.  Install the flasks in the PSU

1.   Ensure that the male flask connectors and female Teflon connectors are clean.

2.  Connect the “pump” connector to the inlet (center) of the back flask and the “return” connector to the outlet (side) of the front flask.  Connect the outlet of the back flask to the inlet of the front flask.

3.  Replace and tighten the flask holding plate.

II. Setting up the sampling site

- Walk to the designated sampling site, as indicated by a bamboo pole located approximately 75m southwest of AWO.

Make sure the sampling site is free of upwind structures or terrain features that would interfere with the flow of air to the PSU intake.

Set the PSU on a level spot, open the case, lift the mast to the upright position, and extend the mast.

NOTE: If winds are high mast does not need to be fully extended. If this is the case, use the lower and larger diameter rod sections for extension. Maintain a minimum of 2m of extension. In low winds use maximum extension.

III. Collecting samples

1.  Open all four stopcocks in any order.

2.  Lift the toggle valve to its upright (flushing) position.  Turn on the pump.  Record the flow rate and battery voltage.  Shut the PSU and walk 10 m downwind

3.  Flush the flasks for at least 5 minutes.

4.  While holding your breath, return to the sampler, open the case, and put the toggle valve in the down (pressurize) position. The toggle has both a correct and incorrect ‘down’ position. The correct down position has the finely textured grip surface toward the ground; when in this position, all spring tension is relieved from the toggle. Close the case, walk 10 m downwind, then resume normal breathing.

5. Wait one minute for the flasks to be pressurized to 6psi.

6.  Hold your breath, return to the sampler, open the case, and turn off the pump.  Resume normal breathing.

7.  Note the pressure and time and record them on the sample sheet.

8.  Close the stopcocks in the following order:

1)  Stopcock connected to “pump” connector

2)  Stopcock connected to “return” connector

3)  The remaining two stopcocks in any order

9.  Pack the mast down, fold it into the case, and latch the case shut.

NOTE: If temperature is below -40 deg C, wait to fold down the mast until indoors.  The small piece of connector tubing gets brittle at those cold temperatures and may break when folding the mast down.

IV. Unloading the samples

- Unload the samples indoors to avoid exposure to light.

- Check that the flask numbers, date, and time are recorded correctly on the sample sheet.

- Scan the sample sheet and save in \ftp\science\NOAA\CCGG\CCGG Scans\.

- Replace the red caps on the flasks and connectors.

- Pack the flasks along with original sample sheet back in the shipping box to indicate that the flasks have been sampled. Take care when closing the shipping box and ensure that the holes in the Styrofoam packing material line up with the flasks. If the lid is hard to close, then back off and recheck arrangement of the packing material. The glass flask necks can easily break if too much force is applied to them.

- Plug the battery charger in and connect it to the PSU.

- Please do not store any flasks in the portable sample unit.

- Sampled flasks are considered DNF (Do Not Freeze) and should always be stored inside a heated building.

- Unsampled flasks are NOT considered DNF but should also be stored inside a heated building as much as possible.

Managing sampled flasks:

Once all the flasks in a crate have been sampled, the crate should be repacked in its original shipping configuration and tagged to indicate that it has been fully sampled and is ready for retro. Outbound flask crates should be zip-tied shut.

Always repack the crates with all of the foam material that they arrived with. When properly packaged, the crate should have a continuous stack of foam all the way to the lid. If there is a gap, confirm that a layer of foam is not being left out.

Make sure to align the slits in the packing foam with each of the flask stems or they will experience too much pressure and can crack when the crate lid is closed.

Once the crate is repacked in its shipping configuration, attach a wire tag to the hasp on the front of the crate stating “Sampled and ready for retro”, as well as the date of repacking, and your initials.

Consumable Resupply Items:

CCGG flask crates (1 crate = 8 weeks)

In from pump

Return to Pump

Center connector on flask has dip tube

Jumper from outlet of 1st flask to inlet of 2nd flask

# NOAA – Halocarbons & Atmospheric Trace Species (HATS) Flasks – AWO

Primary contact:  	Stephen A. Montzka, NOAA GML, Boulder, CO

303-497-6657, Stephen.A.Montzka@noaa.gov

Secondary contact:  	Isaac Vimont, NOAA GML, Boulder, CO

isaac.vimont@noaa.gov

Shipping/Resupply:	Kyle Petersen

208-874-2658; oar.gml.hatsflask@noaa.gov

How Often:		Winter (Oct through June): Twice per month – approximately 8th and 24th

Summer (July, August, and September): Weekly.

Clean Requirements:	Sample ONLY when winds are from the clean air sector.

Supplies Needed:	Safety glasses, 2 un-sampled HATS flasks, data sheet, pencil, watch

Introduction:  The whole air samples in these glass flasks will be analyzed for trace gases that are important components of the global halocarbon chemistry.  Summit is part of a network of sites around the world where samples of remote air are collected.

For more information:  The HATS group within NOAA/ESRL/GMD was formed in the early 1970’s when it was found that chlorinated compounds released in the troposphere contribute to the depletion of stratospheric ozone.  The HATS group has a general mission of studying the sources, sinks, distributions and trends of halogenated compounds along with nitrous oxide and, more recently, carbonyl sulfide. The core activity of the HATS group is the flask network which includes sampling locations at South Pole, Tasmania, Samoa, Hawaii, California, Colorado, Barrow, and Alert. Summit Station is the latest addition to this list of locations for baseline measurements.

Flasks samples are analyzed in Boulder on at least four instruments and approximately 30 species are quantified.  The instruments are gas chromatographs with either electron capture or mass spectrometer detectors.  It is important to understand that most of these compounds are measured in the parts per trillion concentrations.  Contamination levels of many of these compounds exist inside buildings, or can be created by using products like halon fire extinguishers, refrigerant compounds, aerosol products such as dust off, or expanding foam sealants.  Some of these compounds are used for blowing foam so the use of “blueboard” or other foam insulation should be carefully considered.  Please keep in mind that levels of these compounds in seemingly harmless ppm or ppb range can be a serious problem if a sample is contaminated.

See http://www.esrl.noaa.gov/gmd/hats/ for more information.

Procedure:

Sampling from AWO inlet mast:

The flask sampler is designed for field use, but is currently plumbed for use inside the AWO building. The first two bullets apply only if the sampler has been returned from offsite firn air sampling:

- Ensure the 3/8” Dynaflex tubing from the AWO mast is connected to the inlet fitting inside the case.

- Check that all fittings are tight.

Before sampling:

Make sure hands are free of lotions, oils, dirt, ect. And dry

Make sure that winds are from the clean air sector throughout the entire sampling procedure (pump flush, flask flush, and flask filling). If winds blow from out of sector directions, restart the sampling procedures, but please ensure that winds have been within sector for at least ~1 hour before initiating sampling.

- Put on safety glasses!

- Plug in DC converter power.

- Flip toggle switch to Up/Flush position.

- Check that the red pump selection valve is closed to select Pump #2

- Remove the glass plug from the “In from Pump” flask connection point but leave the plug in the “Return to Pump” connection point during the pump flush. Turn on pump #2.  Run a minimum of 30 min. Make sure air is flowing out of the end of the “In from Pump” tube at the flask connection point.

- Turn off pump #2.

- Record length of pump flush.

- Find two un-sampled flasks and place them in the cut foam holder located on the bench next to the pump box.

  - Note:  Please keep flask crates closed at all times except when actively loading and unloading flasks.

- Make sure all sampled flasks are at room temperature before attempting to sample.

  - The viton o-rings in the Ultra-Torr fittings get hard in the cold and are not easy to slip onto the tapered glass connectors.

- Before connecting flasks, carefully bend the tubing so the fittings line up with the glass flask connections.  Take care to not put any sideways pressure on the glass connection.  THEY BREAK EASILY!  Remove the glass plug from the “Return to Pump” tubing.

- Connect the fittings to the flasks as shown in the drawing below.

  - The center inlet valves have dip tubes that go to the bottom of each flask.  Occasionally the valves will be reversed from what it is in the diagram, but make sure to plumb the center dip tube valve as the inlet and the other as the outlet.  Flasks are plumbed and flushed in series and always sampled as pairs.

  - It also may help to tighten the nut on the fitting only loosely (so that the o-ring just starts to become squeezed by the nut) before placing it on the glass.  Then, work the fitting onto the flask: some rotation can help.

  - Always hold the flask by the connector and glass valve when pushing on the fittings. Tighten the Ultra-Torr fittings finger tight and use the metal clips to keep the

  - fittings tight under pressurization.

In from pump

Return to Pump

Center connector on flask has dip tube

Jumper from outlet of 1st flask to inlet of 2nd flask

- Open valves on the flasks in any order.  You will hear a hiss as the flasks depressurize.

- Make sure pump valve is closed and toggle switch is up/open.

- Perform a preliminary leak check by doing the following:

1. Turn on pump #2 and flip toggle switch to Down/Fill.
2.  Allow the flasks to pressurize to 10 to 15 psi and then shut off pump.
3.  Monitor the pressure gauge for a change in pressure over at least 1 min.
If the pressure drops noticeably over this period, there must be a leak in the flask connections.  Flip the toggle switch up to release the pressure and try to reseal the fittings to the flasks.  Redo the preliminary leak check to be sure the leak is fixed.
4.  Once there is no noticeable change in pressure and the fittings appear to be well sealed, you are ready to collect a sample.

Flushing the flasks prior to sampling:

- Make sure toggle is UP for FLUSH

- Turn on pump #2.

- Record time and flow on the check sheet.

- If flow is <3.5 lpm on HATS flow meter, flush for 45 min.  If flow is >3.5 lpm, flush for 30 min. Record length of flask flush.

Taking the sample:

- Flip toggle switch down to pressurize flasks.

- Flasks should pressurize to ~15psi in 3min.  If the pressure goes over 15psi, (or stops at less than 15psi) adjust the backpressure regulator while flasks are filling (big black knob).

- Shut off pump and close the flask valves (as labeled in drawing) in the following order:

  - In from pump

  - Return to pump

  - The other two in any order

- Carefully remove the fitting from the flasks.

- Replace the glass plugs into the “In from Pump” and “Return to Pump” flask connection points.

- Double check the valves before storing the flasks in the crate.

- Record time of flask fill.

- Mark flasks somehow so you know they have been sampled!  Folding the original sample sheet in quarters and storing with its sampled flask pair is a good way to keep track of this.  If using tape to mark, don’t adhere the tape to the glass, just the red caps.

- Take care when closing the shipping box and ensure that the holes in the Styrofoam packing material line up with the flasks. If the lid is hard to close, then back off and recheck arrangement of the packing material. The glass flask necks can easily break if too much force is applied to them. We can’t control how they are treated after they leave Summit—and lots of effort goes into the sampling—so pack as best you can!

- Completely fill out sample sheet and scan a copy in p\HATS Scans\.  Store the original with flasks.

Managing sampled flasks:

Once all the flasks in a crate have been sampled, the crate should be repacked in its original shipping configuration and tagged to indicate that it has been fully sampled and is ready for retro.

Always repack the crates with all the foam material that they arrived with. When properly packaged, the crate should have a continuous stack of foam all the way to the lid. If there is a gap, confirm that a layer of foam is not being left out.

Make sure to align the slits in the packing foam with each of the flask stems or they will experience too much pressure and can crack when the crate lid is closed.

Once the crate is repacked in its shipping configuration, attach a wire tag to the hasp on the front of the crate stating, “Sampled and ready for retro”, as well as the date of repacking, and your initials.

NOTES:

- The original connections from the pump box to the flasks have been removed and saved to enable flasks to be filled straight from the white crate if samples are ever needed from the firn lines site again.  If this is done, Steve will provide detailed instructions.

Consumable Resupply Items:

HATS flask crates

Summer: 	1 crate = 2 months

Winter: 	1 crate = 4 months

# NOAA – Surface Ozone Monitoring – AWO

Document revision number: 1.1

Document revision date: 9/6/2022

Primary contact:  	Peter Effertz, NOAA GML, Boulder, CO

720-314-8380, peter.effertz@noaa.gov

Secondary contact:  	Irina Petropavlovskikh, NOAA GML, Boulder, CO

303-497-6279, irina.petro@noaa.gov

Christy Smith, NOAA GML, Boulder, CO

303-909-0375, Christine.smith@noaa.gov

How often:		Daily – Check

Weekly – calibration check

Six Months – replace particle filter (appx January 1st and July 1st)

Clean requirements:	 	None

Supplies needed: 		Ozone analyzer, timer, particle filters

Introduction: Ozone is a central player in tropospheric chemistry. It is not, however, preserved in snow and ice. In the Arctic there are three separate issues of concern. Stratospheric ozone is being depleted with the most severe depletion taking place in the polar vortices in the Antarctic and Arctic. Over the Arctic sea ice ozone is depleted to near zero levels in the spring near the surface, through complicated chemistry that is not fully understood. In between, in the troposphere, ozone precursors emitted in combustion and other industrial processes lead to enhancements in ozone. As a remote site, Summit provides information on the possible longer-term hemispheric scale alteration of ozone in the troposphere due to human activity. Ozone also provides a context for understanding other atmospheric measurements made at Summit.

Daily Checks:

- Check inlet on the AWO inlet tower (white and orange cylinder) for snow or frost.  Snow can be cleared with gentle tapping of the inlet.  Be careful not to damage the filter in the inlet housing.

- Check the ozone level on the front panel.  The ozone level will vary between approximately 30 ppb to 70 ppb.

- Check the engineering data per the AWO Dailies spreadsheet.  These values are displayed on the aerosol CED laptop within the TEI box.    The following are guidelines for normal ranges:

  - Bench Temperature: 20-40 C

  - Pressure: > 400 mb

  - Flow: 0.4-0.9 lpm

  - Intensity:  65,000-120,000Hz

Weekly Calibration Check: 	(Updated 2022-08-29)

- Enter a comment into CPD3 that weekly surface ozone calibration check is starting

- Push “ENTER” to operate analyzer in “LOCAL” mode.

- Turn on the external pump that forces air though the charcoal tower zero air source into the inlet marked “Zero Air” on the back of the ozone analyzer.

- Push the “RUN” button. This will activate the solenoid so that you will now be sampling from the zero-air source. This display will say “ZERO”

- Allow the analyzer to operate for 15 minutes in this mode. Record ozone level from the display. Display should indicate approximately 0 parts per billion.

- Push the “RUN” button. This will set the internal ozone generator to the “LEVEL 1” position. The display will say “LEVEL 1”.

- Allow the analyzer to operate for 15 minutes in this mode. Record the ozone level from the display. Display should be approximately 160 parts per billion.

- Push the “RUN” button. This will set the internal ozone generator to the “LEVEL 2” position. The display will say “LEVEL 2”.

- Allow the analyzer to operate for 15 minutes in this mode. Record to zone level form the display. Display should indicate approximately 40 parts per billion.

- Push the “RUN” button. This will return the analyzer to sampling outside air. The display will say “SAMPLE”.

- Turn external pump off and push “ENTER” to return analyzer to “REMOTE” mode.

- Enter a comment into CPD3 that the weekly calibration check is complete

- Enter the ppb data for Zero, Level 1, and Level 2 into AWO Dailies spreadsheet in the comment section of the Surface Ozone tab.

Bi-Annual:

- Change the particle filter in the inlet housing every six months. This should be done on approximately January 1 and July 1. Change the filter more frequently if it has been damaged, or after a prolonged period of North Wind equipment operations. Spare filters are stored in the “NOAA aerosol spare and repair” box in the AWO vestibule.

Other:

- Check that the inlet is free of snow.

- The analyzer’s clock should be set to GMT and kept within one minute of an accurate time source. To set the time:

  - Put the instrument in “LOCAL” mode by pressing “ENTER”

  - Press “MENU” and then select “Instrument Controls”

  - Select “Time” and use the cursor arrows to adjust the time

  - Press the “ENTER” button when UTC time is 0 seconds after the time entered.

Consumable Resupply Items:

Conditioned particulate filter: ~2 per year

Shutdown Procedures:

The surface ozone instrument is supported by the UPS and can stay running during power outages; however, if shutdown is required for some reason:

Turn off power using the on/off switch on the front of the TEI instrument.

If this will be an extended outage (more than a day), cap the inlet port into the instrument. Be sure to uncap after turning the instrument back on! Also after an extended outage, perform the weekly instrument calibration check.

TEI Ozone Calibration Check Quick Reference

- Press ENTER => display shows LOCAL

- Turn on external pump

- Turn valve on front panel to ZERO

- Press RUN => display shows ZERO

- Operate for 15 minutes

- Note ozone level (should be approx 0 ppb)

- Turn valve on front panel to CAL

- Press RUN => display shows LEVEL1

- Operate for 15 minutes

- Note ozone level (should be approx 160 ppb)

- Press RUN => display shows LEVEL2

- Operate for 15 minutes

- Note ozone level (should be approx 40 ppb)

- Press RUN => display shows SAMPLE

- Turn off external pump

- Press ENTER => display shows REMOTE

# NOAA – Meteorology Suite – AWO

Primary contact:  	Christy Smith, NOAA GML, Boulder, CO (she QCs data)

GMD Met – NOAA Service Account

303-497-5397, gmd.met@noaa.gov

After hours: (862) 266-2555

Secondary contact:  	Andy Clarke, NOAA GML, Boulder, CO (he does the technical work)

720-484-0754, andrew.clarke@noaa.gov

How Often:  	Daily checks

Annual raising of 2m and 10m instruments

Clean Requirements:  None.

Supplies Needed:  None.

Introduction:  NOAA/ESRL/GMD has a small suite of meteorological instruments at AWO to support Summit science and flight operations.  Temperature, wind speed, and wind direction can be useful in understanding the dynamics of the atmosphere and notify investigators of the meteorological conditions when samples are collected.  In July 2017, instruments were moved to the relocated AWO tower.  See Appendix 3 for instrumentation map.

AWO Instruments: The following instruments are installed on the AWO tower:

2m RTD (primary station temperature)

2m Vaisala HMP155 (dewpoint)

10m heated sonic anemometer (wind speed, wind direction)

10m (approx.) RM Young anemometer (wind speed, wind direction)

Setra pressure transducer (in enclosure at base of tower)

Backup NOAA-made pressure sensor (in enclosure at base of tower)

Software:

LoggerNet: Acquisition software for instruments connected to Campbell CR1000 data logger.

TeraTerm: Realtime UDP feed connected to the Campbell CR1000 data logger.

Near real-time data display website: Raw data available in near-real time – https://www.esrl.noaa.gov/gmd/obop/sum/met_all.html

On site access of the data is through either:

- The LoggerNet software located on the AWO Tech computer and the Tech Office computer. This software is to be used for trouble shooting purposes only.

-The summit weather page: https://conus.summitcamp.org/v1/weather

- The noaa_wx.py script output in the Windows Power Shell (Tech Office computer) console window showing realtime data. Files can be accessed at: \\\fileshare\FTP\science\NOAA\MetData\DailyData. Files are written every minute.

-The realtime TeraTerm UDP feed on the Tech Office computer and Balloon Lab computer.

Procedures:

Daily:

Check the temperature inside the enclosure by looking at the NOAA GML live data web page for Summit:  https://gml.noaa.gov/dv/site/sum/met_all.html. The voltage should be ~12 V.

Check the time using either the TeraTerm UDP output or the noaa_wx.py output in the Windows Power Shell (Tech Office computer) console window to make sure it is correct.

Visually inspect the mock inlet on the Green House tower to see if a visit to AWO is necessary.

If a visit is necessary, inspect all instruments and inlets, looking for excessive frost, listening for inoperable aspirator fans, frayed cables, etc. Note the condition of the instruments in the AWO Dailies for QC purposes. Use the following abbreviations:

clr= instruments clear

lf = light frost (not removed)

hf = heavy frost (not removed)

rm lf =  removed light frost

rm hf = removed heavy frost

N/A = not operational

After Big Storms:

Check AWO tower instruments, climb as necessary to check aspirators and clear frost.

In the Event of Total Power Outage to Meteorological Tower:

If the cause of the power outage has been determined and power can be safely restored to the tower, first complete the following:

Extended outages > 1 hour run the risk of cold soaking the NOAA instruments. Open the NOAA enclosure and disconnect all instrument power supply units except for the heater. Make sure there is no drifting or blowing snow that would contaminate the enclosure.

Restore power to the tower. Allow the heater to turn on and re-warm the internal hardware. Depending on ambient temperatures, this could take 30+ minutes. Following this, re-plug all other electrical connections.

Verify operation of instruments on NOAA webview.

Annually:

April: Check alignment of sonic anemometer. The arrow on top of the sonic should be pointing in the direction of the right (east) leg of AWO. This should be checked prior to the first flight period.

June/July/August: Measure height of instruments off snow and send to PI.  With their guidance, raise the instruments to a level off the snow that will help maintain their specified height throughout the next year.

Full System Shutdown:

Secure power to system:

Turn off the power strip in the enclosure at the base of the AWO Met Tower to shut off the CR1000, sonic power supply, heater, etc.

For an extended power outage:

Remove instruments from the tower (Lufft sonic anemometer, RM Young anemometer, RTD thermometer, Vaisala RH Sensor), store inside.

Remove instrument enclosure from the tower, store inside.

Cover the connectors to prevent them from being iced over.

Full System Startup – from Cold Soak:

Unplug all components from the enclosure power strip except heater.

Turn on power strip and allow enclosure to heat up.

Install any instruments on tower that had been previously removed.

Plug in aspirator fans.  If they don’t run, bring the instruments inside to warm up, reinstall on tower (RTD, HMP155).

Power up CR1000 and sonic power supply – plug in to power strip.

Verify all instruments are recording data.

# NOAA – Firn Air – AWO

Primary contact:  	Andy Clarke, NOAA GML, Boulder, CO

720-484-0754, andrew.clarke@noaa.gov

Secondary contact:  	Christy Smith, NOAA GML, Boulder, CO

303-497-5397, gmd.met@noaa.gov

How Often:  	Monthly checks

Annual raising of firn air lines

Clean Requirements:  None.

Supplies Needed:  Bare bamboo and wire for raises.

Introduction:  Please provide a few sentences on measurement context / science purpose

Spatial orientation and protecting the firn air lines:

The firn air lines exit the AWO building and extend to the ESE. The area where they disappear vertically into the snow surface is sensitive to local foot traffic. This area should be avoided except when required for firn air work tasking. Bamboo poles support the firn air lines in their vertical decent into the snow. Beyond the entry point of the firn air lines, two crossed stakes indicate that further travel to the SE is not allowed. An additional buried sample line extends 2 meters to the SE, and it is sensitive to foot traffic.

To the southwest (towards the Met Tower), a nearby pole has a red flag. This is the snow accumulation stake, used to estimate the depth of the buried inlets in the firn. Similar to the accumulation stakes on the ICESat Traverse and Bamboo Forest, this stake must not be disturbed unless performing well-planned work in consultation with the PI.

Routine raises of the bamboo supporting the horizontal firn air lines are important and can be done at any time. However, PI consultation is required before any raises or adjustments to the accumulation stake, or to last two bamboo supports where the firn air lines enter the snow.

Procedures:

Daily:

No formal daily checks. Investigate if a routine glance suggests anything out of the ordinary with tubing run between AWO and end of firn air lines.

Monthly:

Trace the path of the firn air lines from the AWO passthrough to the point where the firn air lines disappear into the snow. The firn air tubing should disappear into the snow surface with a vertical trajectory. All bamboo poles should be over .9 meters height above the surface? If these criteria are not being met, contact the PI.

Raise bamboo stakes supporting firn air lines but stop before the final two stakes at the entry point. Do not raise the flagged accumulation stake to the southwest of the entry point. If any firn air lines in the horizontal run are not well supported by the bamboo, adjust or add support as needed.

If any firn air lines are in danger of becoming buried raise the bamboo and adjust to prevent burial. If lines have become buried, gently remove them from the snow without the use of metallic tools. It is critical to the experiment not to damage the plastic tubing.

Yearly:

Extend firn air lines. Where the tubes go down into the snow there is a bamboo pole (shown in picture #2).  That one is wired to the firn air lines and can't be pulled out.  The idea is to preserve that spot vertically. When another pole is added, it should be done at that same location, so the bamboo poles and firn lines maintain their aligned vertical descent into the snow.  To raise, add a new stick of bare bamboo, and secure it to the adjacent pole with wire wraps. Additional firn air line tubing should be fed out of the AWO, past all bamboo supports, and to the descent point. With this extra slack, align the tubing vertically along the bamboo, and attach to the bamboo with wire. Estimate remaining length of firn air lines in building.

Measure snow accumulation. The accumulation pole is located to the southwest. Measure from the top of the bamboo stake to the snow surface. Looks like a stub when this was taken. If the pole is within 0.9 meters of the surface, add an additional bamboo pole adjacent, wire to the old pole, and measure the new accumulation stake from top to snow surface. Send measurements of both the old and new accumulation stake heights to PI.

Send before / after photos, measurements, and notes on your work to PI.

Firn air lines entering the snow surface at the correct vertical trajectory (left) and an incorrect angled trajectory (right). Accumulation stake is visible in the background in a normal configuration (left) and as an almost-lost stub (right).

# NOAA – Nephelometer – AWO

Primary contact:  	Erin Boedicker, NOAA GML, Boulder, CO

(720) 258-6017, Erin.Boedicker@noaa.gov

Secondary contact:  	Betsy Andrews, NOAA GML, Boulder, CO

(303) 497-5171, Betsy.Andrews@noaa.gov

Software contact:	Derek Hageman, NOAA GML, Boulder, CO

(303) 497-4652, Derek.Hageman@noaa.gov

How Often:		Daily: Check status to make sure it is logging data.

Monthly: Run a ‘span check’ to check K2 and K4 calibration constants.

Cleaning Requirements:		None

Supplies Needed:		CO2 cylinder in AWO

Introduction:  The nephelometer measures scattering and backscattering properties of aerosols. The data it collects is part of NOAA’s baseline atmospheric measurements.

For more information:

See the following websites for more information (NOAA Federated Aerosol Network & Aerosol Instrumentation – Nephelometer) or contact Betsy Andrews.

ESRL Global Monitoring Laboratory - Global Radiation and Aerosols (noaa.gov)

ESRL Global Monitoring Laboratory - Global Radiation and Aerosols (noaa.gov)

Procedure:

Daily:

- Make sure that CPD3 client window is open on the laptop that is connected to 8-port Edgeport box.  It will be a blue window with a series of numbers displayed in columns.

- After you have selected the CPD3 client window so that it is the active window, hit <Enter>, and a menu will pop up.  Scroll down to “TSI Neph” using the arrow keys on the keyboard.  Once “TSI Neph” is highlighted, hit <Enter>.

- The TSI Neph window will pop up within the CPD3 client window.  The neph window shows the scattering and backscattering values as well as various housekeeping information such as lamp voltage and current.

Example neph window (not the SUM neph!).

- Bsp=scattering, Bbsp=back scattering, BspBK and BbspBK are background values calculated from the hourly automatic neph zeros, BspBKShift and BbspBKShift indicate if there are shifts in the background values – want the BKShift numbers to be less than ~0.5.  Mode indicates whether the neph is in normal measuring mode (NBXX), blanking mode (BBXX) or hourly automated zero mode (ZBXX).  Lamp (A) is lamp current and Lamp (V) is lamp voltage.

- If the nephelometer is logging data, this will be reflected by changing numbers in the neph window, then everything is probably fine. Typically, the scattering numbers will be very small (i.e. less than 5).  If you see any very large numbers (>15) and it’s not a North winds time period you should email Erin and Betsy so that they can have a look at the data to make sure that it looks okay.

- Replace the lamp bulb when the current reaches >7.0 amps. This should occur once every few months. Also, keep an eye on the CPD3 automated daily emails.  When it is time for a lamp replacement a note will be included at the top of the daily email prompting you to “Replace lamp immediately”.  If possible, a lamp replacement should be completed within a day or two of receiving this email notification.  See separate document for TSI Neph bulb changing procedure.

  - Note: When replacing the bulb, there is a skinny wire in front of the bulb protection shield. Please be careful not to damage the wire, as this will damage the entire instrument. Additionally, the bulb has two prongs that slide into grooves housed on a ceramic piece that is quite fragile. Please be mindful and gentle while replacing the bulb so as not to break the ceramic.

- Summit Station experiences the 0100 flag frequently, indicating the “RH is out of range.” This is because cold, dry summit air when brought inside AWO via the sample lines 
(which are not well insulated), lowers the humidity inside the neph. The TSI neph further warms the air because of its 75W lamp so that the RH inside the neph is actually at or near 0%.  The TSI nephelometer interprets an internal RH of 0% as an error condition, even though this condition can occur at Summit. PI’s do not need to be notified of this error.

- Due to low bandwidth and slow internet speeds, Summit occasionally sees an error message in the Aerosol Daily email reporting that, “The latest file transferred has a time stamp that is __ minute(s) behind, relative to server time. Please check that the station computer is correctly synchronized to a time server.” Generally, this is not an issue at SUM but do check the CPD3 computer clock to confirm that it matches the time server.

NOTE: If you don't see what you expect on the screen: 1) observe/record the state of the system for possible troubleshooting, 2) double-click Start-CPD or, if that doesn't work, 3) reboot the computer.  4) Always make an online log entry and describe the symptoms if you encounter problems.

Monthly Nephelometer Span Check Quick Reference

Keep in mind at the end, you’ll want to note these values (start and end time, CO2 cylinder pressure, and average error) in the AWO Dailies spreadsheet under “Comments.”

Note: Span check cannot be started when the neph is in zero mode (ZBXX) - between hh:56 and hh:04.

- Note that a span check cannot be started when the neph is in zero mode (ZBXX) (between hh:56-hh:04).

- Add comment to CPD3 “Beginning monthly span check” + your initials

- In CPD3, press <enter><N><M><4> to get to the span check window.

- Use the <arrow keys> to choose begin and press <enter>.

- When prompted by the program, open the valve on the CO2 cylinder. Enter a comment with the CO2 tank pressure.  You can press enter at any time during the span check to access the main menu to enter a comment.

- A filtered air flush will run (~3 min.)

- When the filtered air flush ends and the CO2 flush begins, close the GREEN Swagelok neph valve on the manifold under the desk (it’s labeled ‘NEPH’ see image below).  CO2 flush will run (~10 min.). During this time the exhaust flow on the neph (measured by the rotameter) should go to zero.

- Enter a comment with the flow rate and CO2 delivery pressure.

- CO2 sampling will begin and run (~5 min.)

- A filtered air flush will run (~3 min.) As soon as this starts, close the CO2 cylinder and open the GREEN Swagelok neph valve on the manifold (see picture above). The neph exhaust flow should go back to what it was.

- Filtered air measurement will run (~10 min.)

- Span check results are calculated and are displayed on the screen.  The average error should be < 5%.  If the error exceeds 5% the span check should be repeated.

If the error is ~40% the exhaust flow was likely not completely cut off. During the second span check ensure that the correct valve is being used and the exhaust flow on the rotameter goes to zero.

If the error is ~60% CO2 was not making it to the instrument. Check the CO2 connections to the neph and the tank.

- Add a comment to CPD3 “Monthly span check ended” + your initials. Document the times for the HEPA filter and span check along with the CO2 tank pressure and average error on the daily log sheet in AWO.

Span check without CO2:

If you are instructed to run a span check without CO2, start the span check as you would normally but do not open the CO2 tank when prompted.  The instrument will run through the entire 30-40 min procedure, but there is no need to open any tanks or move any valves.  At the end you should get an error of ~60%.

Monthly Nephelometer Span Check: (quick reference at the end of this section)

- Keep in mind that at the end, you’ll want to note these values (start and end time, CO2 cylinder pressure, and the average result) collected during the check and log them in the AWO Dailies spreadsheet under “Comments”.

- Note that a span check cannot be started when the neph is in zero mode (ZBXX) (between hh:56-hh:04).

- The main CO2 cylinder valve is closed; this is used to open the flow of CO2. Do not touch the regulator or the black downward facing needle valve - they are set to the desired values for delivery pressure and flow. The delivery pressure should be set at 20-30psi and the flow ~5lpm.

- Note:  A span check cannot be started when the neph is in zero mode (ZBXX), between hh:56 and hh:04.

- Add a comment to CPD3 “Beginning monthly span check” + your initials.

- In CPD3, press <enter> <N><M><5> to get to the span check window.

- The program will notify you of a filtered air flush (~3 min).  It will then prompt you to open the valve on the CO2 tank.

- The program will notify you of the start of the filtered CO2 flush (~10 min).  At this time close the green Swagelok neph valve on the manifold on the bench below the instrument.  Enter a comment in the log with the CO2 tank pressure and flow rate. The flow rate should be approximately 5 lpm. You can press <enter> at any time during the span check to access the main menu to enter a comment.

- The program will notify you that it is sampling CO2 (~5 min).

- The program will notify you that it is flushing the neph with filtered air (~3 min).  Close the CO2 cylinder and open the green Swagelok neph valve on the manifold.

- The program notifies you of the measurement of filtered air (~10 min).  Span check results are calculated and are displayed on the screen.

- Enter a comment into CPD3 to indicate that the span check is complete.

- On the AWO Dailies log under “Comments”, be sure to record the times for the span check, the CO2 cylinder pressure, and the average results of the span check.

Notes:

- The plumbing for a span check is as follows: CO2 tankàRegulatoràFlow control Valve (used to open/control CO2)àRotameteràCopper tubingàPlastic tubingàfilteràneph CO2 inlet.  The plumbing should not change from week to week. Note: Be aware that the further away from the main tank valve you go to control the flow, the more likely leaks can happen.

- If the span check error is ~60%, this indicates that CO2 is not getting to the nephelometer. Check the tank (Is it full? Was the regulator all the way open?); the CO2 lines (cracked tubing? broken connections?); and flow through the nephelometer (did the main sample flow get turned off for the CO2 portion of the span check? Did the CO2 flow rate remain at ~5 lpm for the CO2 portion of the span check?)

- You may be asked to conduct a second span check if the average error is more than 5%. The point of the second span check is to make sure that the large error in the first span check wasn’t caused by operator error (e.g., forgetting to open the CO2 valve or not  having the CO2 valve open enough). If the second span check is also higher than 5%, then contact Erin/Betsy to discuss troubleshooting the problem.

- The trend of the span check results is just as important as any single result.  For example, if you see the average span check errors trending from 2.0% --> 3.0% --> 3.5% --> 4.0% --> 5.0%, then we have to start looking for a problem even though the average error never exceeded 5%.  One poor span check signifies the POSSIBILITY of a problem, but not necessarily a problem.  That is why we should repeat the span check when possible.  The next span check might turn out to be acceptable.

*NOTE: If NO COMMS (in red) starts flashing on the NEPH or CLAP screen, check all the connections on the laptop and USB converter and restart the computer.

Monthly Inlet Inspection:

This is a new inlet head design at SUM so we’re not sure what could go wrong. Two things to check:

- Take a good, close look at the inlet head to see if there appears to be snow blocking the screen under the louvre slats.

- Assess whether the vertical tube is still straight and firmly held.

Please let Erin/Betsy know if you see anything else that concerns you!

Consumable Resupply Items:

CO2 cylinder

Nephelometer lamps

Shutdown Procedures:

The neph, CLAP and MAGIC CPC as well as the computer are powered by the UPS. The pump runs on a different non-UPS circuit. In the event of a planned extended power outage, notify the PI’s beforehand. On startup, the laptop should re-initialize communications with the instruments.

Emergency shutdown:

Turn off power to the neph, CLAP and MAGIC at their individual power sources. Shut down the laptop computer. Unplug pump. Turn off power to any other instrument components.

Monthly Nephelometer Span Check Quick Reference

Span Check: Do not begin between hh:56 and hh:04 when Neph is in Zero mode.

- Add comment to CPD3 “Beginning monthly span check” + your initials

- In CPD3, press <enter><N><M><5> to get to the span check window.

- Use the <arrow keys> to choose begin and press <enter>.

- When prompted by the program, open the valve on the CO2 cylinder. Enter a comment with the CO2 tank pressure.  You can press enter at any time during the span check to access the main menu to enter a comment.

- A filtered air flush will run (~3 min.)

- When the filtered air flush ends and the CO2 flush begins, close the GREEN Swagelok neph valve on the manifold under the desk.  CO2 flush will run (~10 min.)

- Enter a comment with the flow rate and CO2 delivery pressure.

- CO2 sampling will begin and run (~5 min.)

- A filtered air flush will run (~3 min.) As soon as this starts, close the CO2 cylinder and open the GREEN Swagelok neph valve on the manifold.

- Filtered air measurement will run (~10 min.)

- Span check results are calculated and are displayed on the screen.  The average error should be < 5%.  If the error exceeds 5% the span check should be repeated.

- Add a comment to CPD3 “Monthly span check ended” + your initials.  Document the times for the span check along with the CO2 tank pressure and average error on the daily log sheet in AWO.

Span check without CO2:

Start the span check as you would normally.  The instrument will run through the entire 30-40 min procedure, but there is no need to open any tanks or move any valves.  At the end you should get an error of ~60%.

Navigating NOAA CPD3 Software Quick Reference

The CPD3 software should startup and begin logging data automatically when the NOAA aerosols laptop is powered on or restarted.

The software opens to the main base screen displaying current neph, CLAP and MAGIC CPC data (but not ozone data).  Individual windows for each instrument (neph, CLAP, MAGIC CPC, ozone) can be opened on top of the base screen.

To open individual instrument windows:

- Start from the base screen

- Hit <Enter> to bring up the main menu.

- Use up and down arrows to select an instrument from the menu and press <Enter>.

- A new window for the selected instrument will appear and will be highlighted with a brighter outline to indicate that this window is currently selected.

Note:

- When a window is highlighted you can use the arrow keys to change the window’s position on the screen.

- Use the <Tab> key to toggle through each open instrument window and the base screen.

To open sub-menus for individual instruments:

- Open the individual instrument window for the desired instrument making sure that it is highlighted.

- Press the letter <M> key to display the sub-menu for the selected instrument.

To close an individual instrument window:

- Open the sub-menu for the desired instrument window.

- Ensure that ‘Close Window’ is highlighted and press <Enter>

Note:  Opening and closing individual instrument windows in the software has no impact data collection.

# NOAA – Continuous Light Absorption Photometer (CLAP) – AWO

Primary contact:  	Betsy Andrews, NOAA GML, Boulder, CO

(303) 497-5171, Betsy.Andrews@noaa.gov

Secondary contact:  	Patrick Sheridan, NOAA GML, Boulder, CO

(303) 497-6672, Patrick.Sheridan@noaa.gov

Data contact:	Derek Hageman, NOAA GML, Boulder, CO

(303) 497-4652, Derek.Hageman@noaa.gov

How Often:		Daily: Check the transmittance and flow of the CLAP. Record values in the AWO Dailies spreadsheet.  If the transmittance is below 0.75 and the CLAP is on Spot 8, change the CLAP filter.

Clean Requirements:		Tweezers should be used to transfer the CLAP filters so that no residue from your skin mars the surface.

Supplies Needed:		For filter changes, use 47mm quartz fiber filters found in a small square plastic box with red tape on it labeled “CLAP filters.” Tweezers should be used to handle the filters. The red-handled torque driver in the aerosol supplies box should be used to remove and reinstall the nuts on the lid.  A Kimwipe moistened with ethanol or isopropyl alcohol may be used as a last resort to remove stuck on filter bits.

Introduction:  The CLAP measures absorption properties of aerosols. The aerosols are deposited onto a spot on the filter, and the transmittance of that spot is measured against two reference spots through which particle-free air passes. The data collected by the CLAP is a part of NOAA’s baseline atmospheric measurements.

For more information:  See http://www.esrl.noaa.gov/gmd/aero/ or contact Patrick Sheridan.

Also see the “Navigating NOAA CPD3 Software Quick Reference” at the end of the neph protocol for more information about using the CLAP data acquisition software.

Procedure:

Note: Maintenance should not be performed on CLAP during a Neph span check, as it could disrupt the span check.

Daily:

- Make sure that CPD3 is open on the laptop that is connected to the 8-port Edgeport box.  CPD3 will be a blue window with a series of numbers displayed in columns.

- After you have selected the CPD3 window so that it is the active window, hit <Enter>, and a menu will pop up.  Scroll down to “CLAP” using the arrow keys on the keyboard.  Once “CLAP” is highlighted, hit <Enter>.

- The CLAP window will pop up within the CPD3 window.

- Record the Transmittance and Spot Number from the CLAP window.  Once the Transmittance at a given spot drops below 0.7 in the green wavelength, the CLAP should automatically cycle to the next spot.  It should cycle in order from Spot 1 to Spot 8.  If it does not do so, contact Lauren Schmeisser and Betsy Andrews.  It may be necessary to manually advance the spot in this instance.  See next page.  Directions to do so are also included in the “CLAP_User_Manual” on the Summit FTP site.

- The flow of the CLAP should be 0.60 +/- 0.01 lpm as shown on the CPD3 display. If it is lower or higher than this, check the system for leaks, kinks in the tubing, or any other physical cause for the inconsistency and adjust the flow to 0.60 lpm (see below).

- If the CLAP is on spot 8 and the Transmittance < 0.75 at any of the three wavelengths, it is time to change the filter (see below). Record “Green” transmittance value in AWO dailies spreadsheet.

Adjusting CLAP flow:

- Make a comment in CPD3: “Adjusting CLAP flow from x.xx lpm to x.xx lpm” + your initials.

- Turn valve knob on front of CLAP to desired flow setting.  Make sure the flow shown on the CLAP screen corresponds approximately with the flow shown on the CLAP CPD3 screen. The two flow readings can vary by 0.01 lpm; in this case, go with the value shown on the laptop as it reflects the true calibrated flow of the CLAP.

Changing the filter:

- Add a note to CPD3: “Changing CLAP filter” + your initials.

- In the CPDclient (the blue screen), press <Enter><W><M><2> to select the CLAP menu and signal start of a filter change. The bottom of the CLAP window will indicate that a filter change is taking place. The red light on the side of the CLAP will be illuminated, providing an additional indication that a filter change is occurring.

- Using the torque driver unscrew the four nuts on the top of the CLAP, remove the lid, and lay it carefully to the side of the instrument or rest it on the back two bolts.  (Note: pictures below show prototype CLAP – actual nuts to unscrew are different.).

- Remove the filter. The flat surface of the top assembly should NOT be held over the bottom assembly when attempting to remove the filter.   That way, if the filter disintegrates while trying to remove it, no filter pieces will fall into the bottom holes and block the detectors.  Place the filter in a small zip bag and label the bag with the station name and the date the filter was pulled out of the CLAP. The label should have the format: YYMMDD-STN (so December 13, 2000 at Summit Station (SUM) would be labeled: 001213-SUM). Write your initials on the bag.  These filters will be stored on station until they are requested for retro by NOAA.

- Make sure that the top and bottom surface of the filter holder are clean.  You can clean the top section, rubbing off any stuck on debris with your finger while holding it so that the filter surface is facing down toward the benchtop (but not over the bottom section).  Doing the cleaning this way eliminates or at least minimizes the possibility of filter bits getting into either the bottom or top assembly holes.  If the filter crud on the top assembly is resistant to cleaning, a Kimwipe dampened with ethanol or isopropyl can be used to get it off.

- Place a new filter in the filter holder. The white, fuzzy side of the filter must be facing up. Secure the lid on the CLAP by replacing the lid and tightening the 4 nuts until the red-handled torque driver clicks indicating the proper torque has been reached. As you tighten, go in a star pattern, crossing over the middle of the CLAP, in order to ensure the best seal. Note: the lid only fits on one way.

This is an example of the type of stuck-on filter material you would need to remove with ethanol and a Kimwipe. Don't use a q-tip - the fibers can get stuck in the holes and change the light transmittance measurements.

- Let the system autodetect the end of the filter change. This requires you to step away from the CLAP and let it think for a minute or two. A CLAP filter change causes a disturbance in the CLAP measurement that can last several minutes. Letting CPD3 determine a filter change is over will result in a less noisy/spikey data set. Note: If red light on the CLAP is off or the CLAP window no longer says ‘Changing filter…’ then the system has used the autodetect function to determine the filter change is done. NOTE: If the software does not autodetect the end of the filter change after 5 minutes or so, you may have to force the end by pressing the red button on the CLAP instrument or by selecting “Stop filter change” in the CPD3 menu. If instructed by the PI to choose “white filter change”, the end of the filter change must be manually prompted by selecting “Stop filter change”.  A 'white filter change' is usually only necessary when an CLAP instrument is swapped out for a different CLAP and maybe after a really long power outage (many days).

- Check that the CLAP flow is approximately 0.60 lpm. The flow value you should look at is the one on the CLAP window in CPD3, NOT the flow on the instrument screen.

- Add a note to CPD3: “CLAP filter change complete” + your initials.

To Advance Spots Manually:

Note: you can only advance spots consecutively using this method.

- Add a note to CPD3: “Advancing CLAP to spot X” + your initials.

- Press <ENTER><W><M><4>.

Error Flags

The other thing to note about the CLAP is the value of the flags variable on the cpdclient screen. In the picture above the flags value is 0100. That is indicative of a lamp error. The table below lists the flags that might be shown on the cpdclient screen and their meanings. The flags are additive in the hexadecimal system so you could see a flag 07FE which would indicate temperature error, lamp brightness error, lamp error, all the transmittances are below 0.5, and there is a flow error. (i.e., lots of stuff is wrong!)

CLAP flags

| 0400 – temperature error (inlet or block) | 0010 – TrG<0.7 |
| --- | --- |
| 0200 – lamp brightness error | 0008 – TrB<0.5 |
| 0100 – lamp error | 0004 – TrB<0.7 |
| 0080 – TrR< 0.5 | 0002 – flow error |
| 0040 – TrR < 0.7 | 0001 – filter changing |
| 0020 – TrG < 0.5 | 0000 – all is good with instrument |

*NOTE: If NO COMMS (in red) starts flashing on the Neph or CLAP (or ozone monitor) screen, check all the connections on the laptop and USB converter and restart the computer.

Consumable Resupply Items:

47 mm quartz fiber filters

Small zip-lock bags for returning used filter

Shutdown Procedures:

The Neph and CLAP power supply as well as the computer are powered by the UPS. The pump runs on a non- UPS supported circuit. In the event of an extended power outage, notify the PI’s beforehand and shut down the laptop computer. That way no invalid data are collected and on bootup the laptop will re-initialize communications with the instruments. Power could also be removed from the Neph and CLAP in an emergency.

# GFZ Potsdam / Strollo – Seismometer – AWO

Primary contact:	Thomas Zieke, GFZ Potsdam, Germany

tzieke@gfz-potsdam.de

Secondary contact: 	Angelo Strollo (PI)

strollo@gfz-potsdam.de

How Often:  		Weekly:  Check seismometer data at web address listed below.

Upon request:  Level instrument

Clean Requirements:		None

Supplies Needed:  		During raises, plastic shovel

Introduction:

Most knowledge about the deeper interior of the earth is derived from seismological records. Seismic waves generated by earthquakes travel through the globe and sample its major structures on the way. Important information about seismic velocities and densities, structural boundaries, mineral composition, temperature and pressure regimes etc. are hidden in each recorded seismogram and can be retrieved by inverse methods. To obtain a complete picture, globally distributed high quality broadband seismological stations are required to record a full seismologically range in terms of frequency content (10**2 – 10**-6 Hz) and dynamic range (10**-9 – 10**-1 m/s). The technical equipment of the GEOFON network fulfills these requirements and is installed in 50 stations worldwide. (Near) real-time data transmission (via the Internet) from most stations makes the GEOFON data immediately available to the scientific community and provides a perfect tool for rapid determination of earthquake source parameters for scientific purposes but also for earthquake and tsunami early warnings and for use by disaster management. Both near real-time and archive data are openly available to the community from the GEOFON Data Center and are shared with other national and international data centers such as the European ORFEUS Data Center in De Bilt (Netherlands) and the global FDSN/IRIS Data Center (Seattle, USA).

For More Information:

Project website: http://www.gfz-potsdam.de/geofon.

GEOFON wiki page:

https://geofon.gfz-potsdam.de/redmine/login.

Navigate to Projects >>> SUMG; user/pass: techs / ge@summit

Batteries located in the seismometer box at AWO and also in the trench serve as a UPS. The system can run a maximum of 8 hours on this emergency power supply.

To access the seismometer, walk from AWO along the cables strung on bamboo poles in a ENE direction.  Where the cable disappear into the snow surface, poles mark the buried enclosure lid. And a GPS antenna is mounted on a bamboo pole above the surface.

Fig 1: Navigation routes and location of seismometer in relation to AWO and Clean Campaign and Undisturbed Snow Sector flag lines.

Procedure:

Weekly: Confirm that current data are being generated, and that these data show small variations and are not ‘flat-line’. To observe current data, visit station “SUMG” at the following page:

http://geofon.gfz-potsdam.de/waveform/gfn_liveseis.php. Contact the PI if data are missing or do not show variations.

Monthly (or after major storms): Walk the cable run to the seismometer. Raise bamboo monthly to ensure cables remain above the snow surface. At the seismometer, a GPS antenna is mounted on a bamboo pole. Raise the pole and antenna as needed to maintain them above the snow surface.

Emergency shutdown:

The PIs would like to shutdown the Meridian and seiscomp remotely via a web interface. If not possible, they would send instructions. If this is not possible due to urgent circumstances:

- In AWO, unplug PPU from mains power (orange cable).

- In AWO, unplug charger in the aluminum box from the PPU.

- In AWO, disconnect the battery from the power box/charger in aluminum box.

- In AWO Unplug the RJ11 and RJ45 cables from the VDSL modem

# NASA – Bamboo Forest Accumulation

Primary contacts:	Bob Hawley, Dartmouth, bob.hawley@dartmouth.edu

Thomas Neumann, NASA, thomas.neumann@nasa.gov

Secondary contact:	Joe McConnell, DRI, joe.mcconnell@dri.edu

How often:		Weekly – preferably on Wednesdays

Clean requirements:d		None

Supplies needed:		Skis, measuring stick, clipboard, log printout, logbook, pencil

Introduction: An important issue for all the investigations of the relationships between the composition of the atmosphere and snow here at Summit is the timing of snowfall, and how much it is redistributed by wind or lost by sublimation. The forest was established so that accumulation could be measured frequently and the survey is performed on skis to minimize artificial drifts and other surface disturbances in the forest.

When weather, work, and safety conditions allow, the survey is to be completed weekly on Wednesdays, or as close to Wednesday as possible. Use your best judgement and wait for a decent weather window, but if conditions align on Wednesdays, please capitalize on them! The collocated OGRE station makes similar accumulation measurements on Wednesdays and synching the Bamboo Forest survey with these measurements will allow for better comparisons.

More information: Data are available, including plots of previous year-round observations.

Procedure:

- Use skis to travel to the bamboo forest. Starting at the first stake, travel North down the first row and then back along row two traveling South. Continue this pattern for all eleven rows. Consult the Summit Travel Plan.

- Measure accumulation by recording length from bottom of tape to snow surface. This should be recorded on a log printout, best taped to a small clipboard. The log printout is available at: Science Tech OneDrive > Templates.

- Back in the Big House Office, place the log printout in the “Bamboo Forest” logbook, then transcribe data to the Excel data sheet found on the server: /ftp/science/bamboo_forest/Bamboo_Forest_Accumulation_Log.xls

- Scan the page in the logbook to: 
/ftp/science/bamboo_forest//BambooScans/Scans_YYYY/yymmdd.bamboo

Raising Bamboo forest:

- The Bamboo Forest accumulation stakes must be raised annually. This should be completed in August or September, when poles are still easy to remove from the snow. (in years past, this typically has occurred in July when average height reaches ~105cm).

- Go through and collect normal measurements of the bamboo poles, then repeat the process, but each time pulling the bamboo pole out, moving it 5 cm to the North, placing it back in the snow, and re-measuring it.

Consumable Resupply Items:

Logbook

# NASA – ICESat Accumulation

Primary contacts:	Bob Hawley, Dartmouth, 603-646-1425, robert.l.hawley@dartmouth.edu

Thomas Neumann, NASA, thomas.neumann@nasa.gov

How often:		monthly

Clean requirements:		none

Supplies needed:		meter stick, logbook, pencil

Introduction: An important issue for all the investigations of the relationships between the composition of the atmosphere and snow here at Summit is the timing of snowfall, and how much it is redistributed by wind or lost by sublimation. The ICESat transect is used to track accumulation in a similar manner to the Bamboo Forest, but at a larger scale.

More information: Data are available, including plots of previous year-round observations.

Procedure:

- Conducted along the ICESat Transect route.

- Measure accumulation by recording length from bottom of black tape to snow surface, in the “Accumulation” worksheet at each pole in the ICESat Transect. Note: some poles also have white reflective tape; be sure to always measure to the black tape. Every other pole in the middle long section lacks black tape and does not require measurement. The ICESat-2 extension does not need to be measured for accumulation.

- Back in the office, transfer data from the logbook to the Excel data sheet at: 
/ftp/science/ICESat/ICESat_accumulation_log.xls

- Scan the page in the logbook to: 
/ftp/science/ICESat/ICESat_Scans/yyyy/yymmdd.icesat.jpeg

# NASA – ICESat Transect

Primary contacts: 	Bob Hawley, Dartmouth, 603-646-1425, robert.l.hawley@dartmouth.edu

Secondary contacts: 	Thomas Neumann, NASA, thomas.neumann@nasa.gov

Joe Pettit, Earthscope, joe.pettit@earthscope.org; 303-358-1097

How often: 	Standard surveys performed monthly, with a two-traverse sequence coordinated with ICESat-2 overpasses every three months.

Clean requirements:	none

Supplies needed:

- Trimble Zephyr GPS antenna, and antenna cable cap for PolyPod

- Trimble R7 roving receiver

- Earthscope lead-acid battery for receiver

- Yellow SICK Surface Roughness Laser

- Siglin sled with aluminum attachment points

- PolyPod survival sled

- Snow machines (2)

- Equipment for the accumulation survey (see NASA – ICESat Accumulation Protocol)

- Surveyor’s leveling staff

- Safety gear for off-station travel – see seasonally appropriate travel policy

Location: The ICESat transect is located across the skiway from the main station. A flag line extends from the southern end of the skiway to the start of the transect.

Introduction: The launch of ICESat-2 was accompanied by changes in these protocols—please review closely. The purpose of the ICESat Traverse survey is to maintain an accurate ground-truth measurement for spaceborne and airborne laser and radar altimetry systems. It was established to run directly along and across the ground track of NASA’s Ice, Cloud and land Elevation Satellite (ICESat), which was operational between 2003 and 2009. The main instrument onboard ICESat was the Geoscience Laser Altimeter System (GLAS), which was designed to measure ice-sheet surface elevation and elevation change. ICESat-2, successfully launched in September 2018, has a similar mission requirement over the Earth’s ice sheets. In the intervening years between ICESat and ICESat-2 (2009-2018), NASA undertook an airborne campaign known as Operation IceBridge, which aims to “bridge the gap” between the 2 satellites, using airborne laser altimetry in both Greenland and the Antarctic. Additionally, the European Space Agency has launched CryoSat-2, which carries as its primary payload a radar altimetry system with the parallel goal of measuring changes in the Earth's ice cover.

With these disparate past and present systems contributing to the altimetry record, it is more important than ever to maintain our ground measurements. Inter-comparison between different measurement platforms (i.e., what changes took place on the ice sheet between ICESat and ICESat-2) can only be meaningful if they are both cross-calibrated to the same “truth”, and we have established Summit's ICESat transect as that truth for operations in Greenland. Having a continuous set of measurements will allow meaningful comparisons between ICESat, the instruments on IceBridge aircraft, CryoSat-2, and ICESat-2. In addition, the time-series of elevation change provided by these measurements is invaluable to researchers in its own right, as it can provide completely independent rate-of-change estimates.

In 2016, the original ICESat traverse route was augmented with the ‘ICESat-2 extension’. The value of the original ICESat section of the traverse (the southern end of the traverse line) lies in the 10-year time-series. Ten years of data provide an unprecedented time series of surface change and help to fill in the information that can’t be obtained from quarterly satellite assessments or semi-annual airborne campaigns. The value of the ICESat-2 extension (the northern end of the traverse line) is that it consists of a series of waypoints intended to optimize the ground sampling specific to ICESat-2 track placement; the extension will hit a region where 2 of the ICESat-2 tracks cross one another, providing better statistics for ICESat-2 validation.

As of 2022, work is underway to transition the ICESat GNSS survey to a new GNSS logging instrument (the “SLIVER box”).  This new instrument will hopefully make the GNSS portion of the traverse simpler.  In a long time series continuity is critical, so it is important for us to evaluate the two instruments side-by-side for a period of overlap, before discontinuing use of the Earthscope R7 for the ICESat traverse.  As this is a new instrument, we welcome any and all feedback on how it works and are open to suggestions for modifications to make it easier to work with, especially in the cold and dark.

Given that the value of the ICESat section of the traverse is in the time series, do not make changes to this protocol without contacting all of the Primary and Secondary contacts.

Procedures:

NOTE: It is necessary that participants review and follow ‘CP423 Summit ICESat Traverse, Operational’, which covers safety and communication preparations in advance of the traverse, and pre-planned responses in the event of issues. Given the time spent away from station and the demanding ice sheet environment, these preparations are a crucial part of all ICESat Traverses.

With ICESat-2 now active in orbit, the schedule and to protocol have changed. Most months, a coordinated overpass is not scheduled, and the standard traverses are performed as has been done in the past, targeting the dates shown in the calendar at the following link:

https://docs.google.com/spreadsheets/d/1D09MtjVyRGainFbOn1u6WPxd6_R2yPSuukBeDrMyebc

Every three months, the ICESat-2 satellite performs a sequence of two overpasses of the Summit traverse route, with the two overpasses separated by ~8.6 days. They’d like for the technicians to perform an overpass-coordinated traverse, followed by a second traverse (reverse traverse) ~8.6 days later, in which only GPS measurements are made. In the GPS-only traverse, no manual accumulation measurements are made at the bamboo stakes; however, the track depth measurements must still be made.

Reverse Traverse

These two overpass-coordinated traverses should be timed to coincide as closely with the overpasses as safe conditions allow. The best data quality is achieved when the traverse date is closest to the overpass date (before or after), so if a good weather window appears in the days leading up to the overpass date, take advantage! PIs would like us to target traverses for +/- three days from scheduled overpass date, but that may not always be possible. Crew safety absolutely prevails. The desire for close timing increases the importance of advance coordination between the technicians, manager and mechanic, as well as off-site staff when needed. Technicians should initiate a discussion between these parties about one week before each overpass date.

If the survey cannot be performed on the day of the overpass, technicians should record conditions on the intervening days, and communicate these conditions when the post-survey notice is sent to the PIs. The PIs are particularly interested in snowfall, drifting, wind scouring or rime. This information is valuable as it suggests how the snow surface may have changed between the time the satellite measures the surface height and the time that the technicians with the PolyPod measure the surface height.

Preparing the Trimble R7 survey system:

- Ensure all batteries (particularly the large battery to power the R7) are fully charged 1 day before the survey. A discharged battery could take 12 hours to fully charge.

- Verify AWO base station operation (Trimble R9 designated SMM4 accessed at 192.168.1.235). Ensure hourly data files are being collected as expected.

- Verify the R7 receiver configuration file:

  - Connect the R7 receiver to the science tech laptop in the MSF using the power and communications cable plugged into (a) the R7 port 2 and (b) a COM port on the PC or serial-to-USB adapter. The R7 should power on automatically when connected. If not, power on the receiver by pressing the green button.

  - From the start menu, launch “Configuration toolbox” (Start > All Programs > Trimble > Configuration Toolbox > Configuration Toolbox) NOTE: The configuration Toolbox software and related information is available on-line: https://kb.unavco.org/article/trimble-configuration-toolbox-all-versions-160.html Beware of the “save before transmit” bug.

  - Select Communications -> Get File...

  - From the dialog box, select “power_up” and click “Get File”. Note: if “power_up file isn’t on the list, this means the proper configuration file isn’t on the R7. Proceed to Step 5 below, uploading the correct configuration.

  - You will now be presented with a dialog box in which you can verify the correct settings of the receiver. The settings should be as follows:

i.	File

Applied immediately:			box checked

Stored in receiver:			box checked

As auto power up file:			circle checked

Reset to defaults before applying:	box checked

ii.	General

Elevation mask (deg):			10

PDOP mask:				20

Measurement rate:			1 Hz

RTK mode:				Low latency

iii.	Logging

Measurement logging rate:		1 Hz

Position logging rate:			5 mins

Fast static timers:			4 SV: 20; 5 SV: 15; 6+ SV: 10

2nd Measurement logging rate:	Off

2nd Position logging rate:		Off

Enable automatic data logging:	box checked

Enable automatic deletion:		box not checked

iv. 	Antenna

Group:					All

Type:					Zephyr

Method:				Bottom of antenna mount

Measured height (m):			Leave at 0.000 (We will apply the height
							antenna correction in post-processing).

True vertical height (m):		This cannot be changed and will be the 
							measured height (0.000) plus 0.046m 
							for the Zephyr.

v.	Static

Static mode:				Kinematic

vi.	Time Activation

Time Activation Enabled:		Yes

Date and Time:			1/1/2000 00:00:00 (UTC)

Repeat Interval:			24 Hours

- If the parameters are shown correctly, as above, click “Close”

- If the parameters are incorrect, upload the correct configuration to the R7 using the following procedure:

  - Verify that any useful GPS data on the CompactFlash memory card in the R7 has been transferred, as this process will erase the card. Perform a hard reset of the receiver with the memory card installed. A hard reset is performed by pressing and holding the power button for 30 seconds.  The indicator light will flash at the end of the interval.  This will reformat the card and reset to factory defaults. To reset to factory settings without reformatting flash card, press and hold the power button for 15 seconds.  Note that when the flash card is reformatted during a hard reset, all data are lost on the card.  The receiver will be reset, the card will be re-initialized, and any data on the card will be deleted. In Configuration Toolbox, Select File -> Open...In the dialog box choose the file “ICESat_1sec_ppk”. If this file is not shown, navigate to the directory: “C:\Users\ARO-Summit-Technicia\Polar Field Services\PFS-ARC-Summit Techs - Documents\Instrument Manuals and Supporting Documents\Earthscope\Receiver Configuration Files”.

  - Click “Open”.

  - Verify the file settings are as indicated above.

  - Click “Transmit”.

  - If the dialog comes up asking if you want to replace the file “power up” say “Yes”.

  - You should see the message “the application successfully transmitted to the receiver- click “OK”.

- Power down the receiver and unplug the cable from the receiver.

- Take the insulated Styrofoam case with receiver, battery, hot water bottles, and the antenna out to the PolyPod sled.

- Mount the antenna on the stud at the rear of the PolyPod then attach the antenna cable. Verify that the antenna is screwed down completely so that the antenna height is consistent survey to survey.

Connect the antenna cable to the R7 receiver inside.

Connect the battery cable to the R7 receiver. The receiver should turn on automatically. If it does not, press the power button to turn on. Once the R7 receiver is running at the MSF, leave it running continuously until you return with the receiver to the MSF, after the full completion of the ICESat transect.

Important: Always double check your power connection!  Data has been lost in the past due to the R7 having an improperly seated power cable.  When connecting the cable be sure to align the red dots and push in FIRMLY.  Give the cable a tug to make sure that it is locked into place.

Verify that the R7 receiver is powered, receiving satellite signals, and is logging data by verifying the following: The green power LED is on solidly, the satellite LED is flashing red once per second and the yellow “log data” LED is on solidly. All of these lights should come on with 1 minute of the R7 being powered on.

Solid green power light and solid red Satellite light, but NO yellow “log data” light indicates the proper ‘power_up’ configuration file either isn’t present or failed to load. Take the receiver back inside and perform the verify Config File procedure above.

After ~15 minutes the yellow LED will begin to blink once per second to indicate that the R7 has initialized (a position solution has been calculated). It is not necessary to wait for this indication before starting the survey.

Ensure that all safety and support equipment is present, as described in ‘CP423 Summit ICESat Traverse, Operational’. This includes, but is not limited to, adequate ECW gear and spares, a survival bag, food, a thermos of hot water, and appropriate communications equipment. A typical survey requires about 4 hours but bring sufficient supplies for flexibility and safety in case of issues.

SICK Surface Roughness Laser preparation for survey:

See document titled “SLIVERProtocols2024.pdf” at location: C:\Users\ARO-SummitScienceTec\Polar Field Services\PFS-ARC-Summit Techs - Documents\Instrument Protocols\ICESat

Performing the ICESat Traverse

- Drive to the start of the ICESat flag line.

  - At the start of the ICESat flag line, the halfway point, and the turnoff to the ICESat extension route, and more frequently if any issues are suspected, stop to inspect the status lights on the R7 and SLIVER receivers and confirm that the lights indicate 1) good power connection 2) good satellite signals and 3) writing to file as normal. This is to avoid extended data outages if cable-related or other issues develop during the traverse travel.

- Note the start time in UTC.

- At the start of the ICESat flag line, measure the depth to which the PolyPod's deepest runner sinks in the snow (which will vary with snow conditions) using the following procedure (see photo):

  - Take the ICESat field book or clipboard and lay it on the snow at the edge of the PolyPod's track, approximating the actual surface

  - Measure and record the distance between the bottom edge of the book and the deepest part of the PolyPod's track.

  - This distance will be used by the PIs to subtract from the fixed distance between the bottom of the antenna mount and the bottom runner of the PolyPod (indicated inside the PolyPod).

- Drive the accumulation line, stopping to measure the stakes (see ICESat Accumulation Protocol). The PolyPod should be towed on the right side of the bamboo stakes.

  - The survey sled must be towed slowly. Maximum speed during the survey is 20 km/h or 12 mph. Confirm whether the snowmobile in use has km/h or mph speedometer units. This speed will result in position coordinates with 5-meter spacing, which is the science team requirement.

  - There is no need to align the GPS antenna (at the back of the PolyPod) with the bamboo flags when stopping. Measuring accumulation with the snow machines stopped adjacent the flag (allowing one person to measure and the other to record) is the preferred procedure.

- Drive the entire distance of each line of bamboo and then make an outside, donut shaped turn to line up on the next line (as shown below in green; do not make sweeping turns as shown in red).

- Each corner will be marked with a directional flag indicating you’ve reached a turn:

Left turn							Right turn

- The survey track was extended in 2016 for ICESat-2; this additional ‘ICESat-2 Extension’ segment is shown below. At the cluster of flags marking the end of the original accumulation survey, turn nearly around and follow the ICESat-2 Extension flag line in a clockwise direction, with the receivers still collecting data. Drive with the PolyPod to the right of the red flags. Several things are done differently for the extension:

  - There is no need to turn the PolyPod in a loop at these corner flags, as these waypoints are outside of the intersection of the ICESat-2 tracks

  - There is no need to loiter at the corners or route markers.

  - Accumulation is not measured at the ICESat-2 Extension flags: there is no need to stop and measure bamboo height along the extension.

At the end of the ICESat-2 extension, measure and record the depth to which the PolyPod's deepest runner sinks in the snow as above.

Note the finish time in UTC.

Return to the MSF.

At the MSF, bring the receivers inside. The data record should start and stop at the MSF.

R7 receiver: note the appearance of the indicator lights on the R7 receiver, then power off the receiver by holding the power button down for two seconds, or by unplugging the power cable.

    - Holding down the power button on the R7 for longer than 2 seconds

    - format CF card

  - SLIVER box: It is not necessary to bring the logging box inside immediately. When ready, open the box and disconnect the battery from the GPS.  If convenient, you can plug the battery in to charge while you download the data (It will likely finish charging by the time you are done downloading and ready to put the box away for the next traverse)

Take a break, relax, get a cup of hot chocolate. Come back and resume procedure in an hour when you are ready.

After the ICESat Traverse:

Create a directory at \ftp\science\ICESat\ICESat_Data\YYYY named: ICESATYYMMDD where YYYY is the year and YYMMDD is the year, month, and day of the survey.

Download the data from the R7 receiver and convert file type:

Remove the Compact Flash (CF) card from the R7 via the hinged yellow door on the bottom.

Connect it to a Science Tech laptop computer running Windows 10 with a USB card reader and locate the CF card drive.

Drag and drop the .T01 file(s) for your recent survey from the flash card into the desktop folder called “runpkr00”.

The files are all named with the following file naming scheme: 9876xxxy.T01 where xxx is the Julian Day and y is a sequence number of files within the day. Select the “Details” button in the upper right corner to view the creation date of the files, to help you figure out which one is the one you want. Be sure to select the current day's files.

  - Julian dates are no longer accurate. Instead, determine which files belong to the day using the size of the files. Once located and all data processing is done, move this data file to the Archive folder and correct year folder so that on the next traverse it remains clear which data file corresponds to each. Note all other files in the main folder and their sizes to ensure that in the future those aren’t mistaken for your traverse’s data.

  - Update the READ ME.txt file with which data file corresponds to the traverse dates.

Note that the receiver logs data whenever it is powered on. If the R7 has lost and then regained power over the course of your survey (this has historically been an issue) then multiple files will have been created.  Look at the file size for guidance: Larger files indicate longer data recording times. Select all files from the time you left the MSF to the time you returned to the MSF. You can select multiple files by hitting shift-click.

Eject and remove the CF card from the reader and return to the R7.

Open a PowerShell command prompt window.

Navigate to runpkr00 directory: in the OS (C:) drive

Folder name: runpkr00.exe

Drag the ICESat file into the runprk00.exe folder

Open PowerShell and run the command: cd C:\runpkr00.exe [[Enter]

Run the command ls [Enter]

Convert file: ‘.\runpkr00 -d [filename].t01’ [Enter]

Confirm presence of new [filename].dat file: ‘ls’ [Enter]

For example : if your file name is 98762310 the command you run is:

.\runpkr00 -d 98762310.t01

Move your new *.dat and original *.t01 roving data file(s) to the survey directory created on the FTP. All data files should be removed from the runpkr00 directory.

Recharge the battery after each use, to prolong battery life.

- Download data from the SMM4 base station and place in the folder you created on the FTP server.

Data can be downloaded from the base station via any computer on the Summit local network. First, open a web browser and connect to: http://192.168.1.235. When/if prompted, enter the login credentials. The username and password for SMM4 Summit users are: Username: summit, Password: gps. The receiver records high-rate (1 hz) data compiled into hourly data files, however these files are not archived at Earthscope and will not remain on the receiver indefinitely. Therefore, after each ICESat survey download the simultaneous high-rate data files and place them on the FTP server for safekeeping. Download the 1hr1Hz data for SMM4.

For the Trimble (SMM4) base station, one high-rate data file is written per hour. These files are stored in per-day subdirectories, and are named SMM4yyyymmddhhhhB.T02. For example, if the ICESat survey was performed between 1315 and 1615 UTC, download the high-rate base data files for the hours 1300,1400,1500, and 1600 of that day. Note that you cannot download a file to which data are currently being written; you must wait until after the top of the hour.

Download the .T02 files for the corresponding hours of the survey by clicking on the icon to the left of the filename. The SMM4 files are also converted to the RINEX format. Click “Convert” and select Rinex 2.11 and Observables before selecting “Convert and Download”. The extension of the converted file should be .24O.

Once downloaded, save both the .T02 and .24O files into a folder titled “SMM4” in the FTP directory created for the survey: \ftp\Science\ICESat\ICESat_Data\YYYY\ICESATYYMMDD

Create a .txt file with background metadata for your survey. Copy the “polypod_depth_template.txt” file from the \ftp\science\ICESat\resources\ directory and save a copy titled “polypod_depth_YYMMDD.txt” in the survey directory titled.

Edit the file to reflect the survey date, start and end times, and the PolyPod track depth measurements made at the start and end of the survey. Note, that the template also includes a “Distance between Antenna Base Plane and Bottom of the Runners”. This distance should remain consistent as long as there are no changes to the antenna or polypod mounting configuration. With the current configuration this distance is 1.797m.

  - If this was an overpass-coordinated survey, note relevant weather conditions (e.g., snowfall, drifting, scouring, rime) between the date of the overpass and the date the survey was performed. Save the file as “polypod_depth_YYMMDD.txt” in the same directory as the GPS files.

- All the files should now be saved to the FTP directory: \ftp\science\ICESat\ICESat_Data\YYYY. The new folder should be named: ICESATYYMMDD where YYMMDD is the year, month, day.

- Please notify Tom (thomas.neumann@nasa.gov) and Bob (robert.l.hawley@dartmouth.edu) when the base station and rover segment files have been uploaded. And you're done! Thanks for collecting these data for us!

Raising ICESat:

- Should be completed over the course of one day in July or August, when poles are still easy to remove from the snow.

- Complete the ICESat transect once as usual, then go through and raise all bamboo poles. Contact the PI regarding which way to move the poles (e.g. 15cm forward). Then repeat ICESat.

Caution: Remember that safety is the over-riding issue for measuring the ICESat Array. Venture out when conditions are favorable and minimize risks due to weather and other factors. Ride snowmobiles conservatively. Postpone measurements if conditions are questionable.

R-7 with cable connections to battery and antenna.

Ready for ICESat.

Before you leave for ICESat…

- Mount antenna to Polypod

- Connect cable: Antenna → Trimble R7 (GPS Port)

- Connect cable: Battery → Trimble R7 (Port 2)

(Press green power button to turn on, if necessary)

- Verify:

1. Green power LED solid

2. Red satellite LED blinking 1x/second

3. Yellow data log LED solid

When you depart the MSF and arrive at the ICESat traverse…

- Press green power button on Trimble R7

- Verify LEDs as above

- Measure depth of Polypod track at start of the ICESat traverse and record in notebook

- Drive transect at <20km/hr. Stop ~10 seconds at each pole.

When you finish the ICESat-2 extension and return to the MSF…

- Measure depth of Polypod track at end of the ICESat-2 extension and record in notebook

- Verify yellow data log LED is blinking

** If LED is solid, wait until it starts blinking **

- Turn off Trimble R7

# DTU – Magnetometer – AWO

Primary contact: 	Anna Naemi Willer, anna@space.dtu.dk

How often:	As requested by PI

Clean requirements:		None

Supplies needed: 		None

Introduction: This work supports two complementary scientific objectives.

(1) Project IceBase is a high altitude geomagnetic survey to be proposed by a consortium around Goddard Space Flight Center to NASA to investigate the geothermal heat flux below the Greenland ice cap. The project aims at producing a Greenland-wide map of magnetic crust depth (Curie-depth), indicative for geothermal heat flux. The derived heat flux map is a boundary condition for ice sheet models to improve, among other things, estimates for global sea level rise due to melting of the Greenland ice sheet. Ground magnetometers are critical when correcting the survey data for natural geomagnetic time variations. Data from Summit Station, due to its location in Central Greenland, in combination with the below mentioned array, is crucial here.

(2) The Greenland Magnetometer Array operated by DTU Space is a permanent array of some 15 magnetometer stations located on the Greenland East and West Coasts. The array is ideal for investigating the polar ionospheric current systems and processes related to the coupling of energy and momentum from the solar wind to the magnetosphere and ionosphere.  Data are interpreted in combination with satellite data (e.g. NASA's Themis mission, ESA's Cluster mission), or with conjugate stations from Antarctica.  The proposed Summit magnetometer experiment will, apart from improved geographical coverage, provide data from the electrically insulating ice cap.  This data will be less affected by induced electric currents in surrounding oceans and underlying bedrock than the coastal stations, thus improving the scientific value of the array data as a whole.

Installation details:

This project was installed by the PI in June 2014. The magnetometer was placed in a pit (approximately one meter, at time of installation) approximately 100 meters northwest of AWO. A small data logger, electronics box, and transformer is housed inside AWO. A cable is buried under the snow surface between AWO and the magnetometer.

Procedure:

The instrument runs autonomously. Any support needs will be communicated by the PI.

Shutdown: If shutdown is required (e.g., UPS depleted), please shutdown the computer via the normal Windows menu and start the machine once power returns. The magnetometer should start up automatically after restarting. Record the times of the outage and inform the PIs. Advance notification is otherwise preferred.

# Vieregg - RNO-G

Primary Contact (technical):      Cosmin Deaconu, UChicago
				      775-846-9105(cell), cozzyd@kicp.uchicago.edu
Secondary Contact (technical):  Eric Oberla, UChicago
				      614-546-7733 (cell), ejo@uchicago.edu
Administrative Contact:             Abigail Vieregg 
	                                          avieregg@kicp.uchicago.edu

How often: Only as directed or emergency shut down.

Clean requirements:  None (RFI requirements if visiting site).

Introduction:

This document describes operations and procedures for science techs for the RNO-G detector array at Summit Station. This document is separated into three parts, describing different locations of our equipment.

1) The Detector Array, which requires no planned regular intervention from PFS at this time, but procedures for approaching it are documented in case someone wants to approach the array.

2) The server rack inside MSF. This may require intervention in the case of a long unplanned power outage (longer than 20 minutes). Instructions for turning off all the equipment (in case of a serious unplanned emergency like a fire) are also provided.

3) The communications equipment inside the Big House office. This does not require any intervention in case of a power outage (it will safely turn off and back on by itself), but may require to be turned off if someone is working on the roof near the rooftop antennas. Additionally, notification of any power events would be helpful.

Detector Array

Checks required: None

Emergency shutdown: None

Special requirements: RFI mitigation if approaching, see below.

Description:

Our detector array consists of independent stations, each autonomously powered via solar power and communicating to the station via wireless systems with no intervention from PFS. We do not require any regular checks, and such checks would be difficult in any case. A map of the layout is below, also available at https://rno-g.github.io/station-map/

Our detector array consists of a set of radio antennas (some near the surface, some deep underground) for detecting neutrinos (our band is approximately 100 MHz – 700 MHz), associated readout electronics (including a GPS), a power system (solar panels + batteries), and a communications system (LTE Band 8 + LoRaWAN, both approximately 900 MHz).

While we do not require any regular checks, if for some reason someone needs to approach one of our stations, special requirements are needed due to the fact that any radio frequency interference (RFI) is problematic for our stations.

Procedures for Approaching Station:

If approaching a station is required, we ask that we are notified. If prior notification is not feasible, at least log when you get there and leave. The radio emitted from snowmobiles and portable radios can potentially damage the sensitive radio equipment at our stations. For this reason, please park a snowmobile at least 50 m away from our stations and do not transmit with a radio within 50 m (receiving is fine, since that does not produce RF emission, but it might be safer to just turn off the radio).

MSF Server Rack

Checks required: None unless directed.

Emergency shutdown: Turn off server in case of prolonged power outage.

Special requirements: None

(Photo: Jennie Mowatt)

Our server rack in MSF houses the “backend” for our detector system. Included within our rack are:

- Our server, which stores data, and is the backend for our comms systems. The server is a DELL R7515, with a manual available here: https://dl.dell.com/topicspdf/poweredge-r7515_owners-manual_en-us.pdf

- A network switch

- A raspberry pi that serves as an airplane detector and network time server.

- A GPS receiver (connected via USB to the raspberry pi), used to get GPS timing (and can also be used as an RTK base).

- A software defined radio (connected via USB to raspberry pi), used for airplane tracker

- A UPS + extra battery.  The UPS model is this: https://www.tripplite.com/smartonline-120v-1kva-900w-double-conversion-ups-2u-rack-tower-extended-run-snmpwebcard-option-lcd-display-usb-db9-serial~SU1000RTXLCD2U

- A LoRaWAN gateway, the radio access part of our LoRaWAN network.

The rack is connected to several antennas on the roof

- Our LoRaWAN communications antenna (900 MHz), used for our backup network

- A GPS antenna, used for the NTP timekeeping

- A plane tracker antenna, to keep track of flights flying overhead.

The UPS + extra battery allow running the server for at least two hours without power. The other devices on the UPS require very little power by comparison. The server will automatically turn off when the UPS battery is below 20% capacity, however for a prolonged power outage it is preferable to turn off the server before the battery is critical.

We will be maintaining the server remotely. Note however if there is a network issue between the Big House and MSF, we will notice quickly and need it to be resolved ASAP, particularly when we are taking data.

Emergency Server Shutdown procedures:

There are two ways to turn off the server in case of an unexpected prolonged power outage.

1) press the power button on the server, near where “3” points to in the figure below.

2) Use the management interface (note that this requires the network to be working, which might not be the case if the Big House office is down).

- From Summit, this involves navigating a web browser to https://192.168.99.51

- Log into the management system with username sciencetech and password !RNO-G!10565

- Press the Graceful Shutdown button (below)

In either case, please notify us that this was done.

After power is restored, the server may be turned on by either method (press button, or use management interface). We can also do it via the management interface, but if it’s a weird hour, it may be helpful for you to do it.

Emergency Hard Power Off Procedure:

If for some reason, all equipment in the rack must be turned off, the procedure is to turn off the UPS by holding the off button until it turns off. You can unplug the UPS to turn it completely off.

Big House Communications Equipment

Checks required: None unless directed

Emergency shutdown: None

Special requirements: Turn off if anybody will be near (within 1 m) of rooftop antennas.

RNO-G LTE + GPS Antenna

In the Big House office, we have our LTE radio, a UPS (https://www.tripplite.com/smartonline-120v-750va-600w-double-conversion-ups-tower-extended-run-snmpwebcard-option-usb-db9-serial~su750xl) , and associated network cables. On the roof of the big house, we have our LTE antennas (2 big panel antennas) and a GPS antenna (for time synchronization). The LTE radio provides the radio access part of our LTE network (though the backend is running on the server).

The UPS will keep the LTE up for short power outages, but if it runs out of power nothing bad will happen (the only way to turn off the LTE radio is to remove power anyway).

Note that if there is a disruption to the network between the MSF and Big House, or between the Big House switch and the LTE radio, our cell network will fail to work.

Antenna Approach Requirements:

Due to relatively high output power and directional antennas, it is against FCC safety regulations for anybody to be right next to the antennas on the roof while the radio is on. If someone must get within 1 m of the antennas for some reason, it is advisable to turn off the LTE radio temporarily. Please let us know in advance if possible so we are aware.

To turn off the LTE radio, simply hold the off button on the UPS until you hear a beep. Alternatively, you may unplug the LTE radio from the UPS. When the LTE radio is off, the green light will shut off.

To turn the LTE radio back on, press hold power on  button on the UPS until you hear a beep (or, if you unplugged it, plug it back in). A light should turn on on the LTE radio.

# Dartmouth / SCO – GIS data management

Primary contact: 	Bob Hawley, Dartmouth, 603-646-1425, robert.l.hawley@dartmouth.edu

Secondary contact:    Ian Geraghty, (970) 946-6771, ian.g@polarfield.com

How often:	Once per phase

Clean requirements:	None

Supplies needed		Handheld GPS unit

Computer with Garmin BaseCamp software

Introduction: For any disturbance or location, we wish to record the coordinates for input into the Summit GIS for future use. This will ultimately serve as a resource for the science and operations community where each year of activity at Summit can be viewed as a GIS layer. For some applications, such as a snow pit or firn core, a single point is appropriate.  For other applications, such as a flagged route, a track is more appropriate.  For large structures, you may wish to mark each corner with a separate waypoint, or record a track around the structure.

Points are collected using handheld GPS units (see “Instruction Manual for Garmin GPSMAP 76CSx” for more detail on unit specifics) and are processed into a cohesive set using the Garmin BaseCamp (formerly MapSource) software at Summit. This set will then be exported in a text file for further processing and integration into the larger GIS database at Dartmouth.

Determining disturbance and location sites:

Data should be collected for any site that could be of importance to future operational, construction, or science activities at Summit Station or in the vicinity of Summit Station. Below is a list of on-station and off-station examples of sites that require data collection.

| On-station | Off-station |
| --- | --- |
| Buildings | Buildings |
| Utility vaults | Power lines |
| Utility lines | Boreholes |
| Power lines | Science installations |
| Outhouses | Science vaults / subsurface work areas |
| Tent city | Towers |
| Summer temporary structures | Science activity areas |
| Boreholes | Sampling / pit sites |
| Science installations | Tracks for access into Clean Air Zone |
| Towers | Tracks for regular travel routes |
| HF Antenna arrays |  |
| Freezer trenches |  |
| Berms |  |

Procedure:

Collecting New Points:

- Make sure your hand-held GPS is wiped of existing saved data. This will simplify processing later as you won’t have extra points/tracks to sort.

- Near the end of each month, or throughout the month, travel to all new disturbances on station as well as field camps, and record waypoints and/or tracks to capture them for later processing.

- Make sure to write down the point and/or track names and a brief description for later.

Downloading Points to Garmin BaseCamp:

Connect the GPS to the computer using the Garmin connector cable and a serial to USB converter dongle and turn on the GPS unit.

Open the BaseCamp program.

Select “File” -> “New” -> List and name the new list in the following format:  YYYYMMDD

Under the “Device” menu select “Receive From Device”

A window will appear. Select “Waypoints”, “Routes”, and/or “Tracks” depending on what you have collected, then click “Receive”.

Another window will appear, click “Ok”

Your data will then populate a list on the left, and a map on the right.

At this point, you can go through and rename your collected points to appropriately represent the disturbance or location marked if you did not do this during data collection. To rename a point or route, simply double click the point in the list which you wish to rename to open its properties and then edit the name field in that properties window.

Once you are finished naming all points, save the file to the FTP server at:

\ftp\science\Hawley_SCO_GIS\gis\GIS Data 20YY

First, make sure the list you just created is highlighted in the menu on the left hand side of the screen, then Select “File” -> “Export” –> “Export YYYYMMDD”, where YYYYMMDD is the name of the list you want to save.

In the window that pops up, name the file with the following nomenclature: YYMMDD.GISdata.gpx, making sure to select .gpx for the file type

Entering Metadata:

For each point/route you collected, metadata should be entered into BaseCamp.

Double click on each new point or route you created to access the properties window for that point (same as renaming, above), then click on the “Notes” tab and enter information that provides context for each point/route collected. This must include the season and year, and a brief description. For example, “Bamboo forest snow accumulation study, site visited weekly, summer 2014.” You should have Metadata entered for every point, and every route.

Exporting Points to a Text Document:

Next you will need to save the same set of points as a text file for processing at Dartmouth. Save it in the same location as the .GISdata file.
\ftp\science\Hawley_SCO_GIS\gis\GIS Data 20YY

To do this, click “File” -> “Export” -> “Export YYYYMMDD” then click the drop down arrow on “Save as Type” and select “Text (tab delimited)(*.txt)”

Name the file with the following nomenclature: YYMMDD.GIStext.txt

Uploading safety points to hand-held units for station use:

Connect the GPS to the computer using the Garmin connector cable and a serial to USB converter dongle and turn on the GPS unit.

Open the BaseCamp program.

In the left hand sidebar select the most recent safety points file “SummitSafetyPoints20YY” from the list, where YY is the current year.

From the menu bar at the top of the screen select “Device” -> “Send to Device” -> “Send SummitSafetyPoints20YY to Device”

Please email Bob (robert.l.hawley@dartmouth.edu ) when the files have been uploaded, making sure to attach the text file. And you're done! Thanks for collecting these data for us!

# SCO – Clean Air / Science Sector Reporting

Primary contact:	Sam Dorsi, 406-461-5298, sam@polarfield.com

How Often:		Weekly

Clean Requirements:		None

Supplies Needed:		Computer and email

Introduction:

The pristine environment of the ice sheet surrounding Summit Station, and the historical context of the GISP 2 paleoclimate record, has made Summit Station an ongoing research site for investigating high latitude physical processes, atmospheric and snow science, ice core interpretation, and climate change. As the station has grown over the last 30 years, so too have the local impacts of human activities. Many of the research projects at Summit Station can be significantly impacted by local pollution sources. Documentation of certain human traffic and emissions impacts on science activities is important for preserving the scientific record in the context of human activities at Summit.

Procedure:

Certain activities must be documented in the GEOSummit North Winds Operations Log. This document is located at:

\ftp\science\clean_air_traffic\geosummit_north_winds_ops_logYYYY.xls

The activities requiring documentation are:

- any operation of equipment during north winds

- any equipment operation on or across the skiway, including ICESat and all flight ops

- any pedestrian, heavy equipment, scientific, or other human traffic into the science sectors aside from Bamboo Forest, Seismometer, Met Tower, or AWO.

Technicians can enter events directly into the digital clean air log above. Other station personnel will make entries a the physical clipboard located in the Big House office, so technicians should check the clipboard and transfer to the digital log once each week (typical prior to completing the weekly report).

A full explanation of North Winds and Clean Air Zone protocols is available in the Clean Air Management Plan. The most current version of this document is posted on the Big House safety board. Any questions should be directed to Sam Dorsi.

Weekly Weather Report:

- In PowerShell:

cd  "C:\Users\ARO-Summit-Technicia\Polar Field Services\PFS-ARC-Summit Techs - Documents\Reports\weekly_wx\"

python ./weekly_wx.py

- The resulting report and downloaded NOAA minute weather data files can be located in the Tech One Drive/Reports/weekly_wx/

- For ICECAPS 1200z weather data use

python weekly_wx_icecaps.py

# NOAA / Battelle-ARO – Weather Station Displays

Primary contact:	NOAA Data:  Christy Smith - GML Met (update to Andy Clarke)

303-497-5397, gmd.met@noaa.gov

Summit Wx Displays: PFS ITC, ITC-Support@polarfield.com

General Questions:	Sam Dorsi, 406-461-5298, sam@polarfield.com

How often:		Daily

Clean requirements:		None

Supplies needed:		None

Introduction: This protocol describes the various station weather equipment, displays, and their points of contact.  The protocol for the altimeter settings and its related data and displays are in a separate document located in the /ftp/science/protocols directory.

There are two sources of weather data at Summit Station: the NOAA weather instruments on the AWO tower and the Battelle-ARO instruments on the Berthing Module roof.  Details about the NOAA weather station and its maintenance are provided in the section of this protocol “NOAA – Coastal Environmental Meteorology Suite.”  The Berthing Module weather station instruments, located on the south railing of the Berthing Module roof, consist of an RM Young anemometer and an aspirated temperature sensor.  Spares for these instruments are located in the MSF vestibule.  The data from the Berthing Module instruments are visible on an LCD display located with the communications equipment in the main hallway of the Berthing Module.  The Berthing Module met instruments are considered the backup, non-official instruments.

Summit Station weather is displayed in four places.  Locations displaying NOAA data are:  NOAA GML webpage, Summit Weather html display, and the ARSLS InSite webpage.  The LCD screen in the Berthing Module main hallway display data for the secondary met instruments mounted on the Berthing Module roof.  More details about these displays are provided below.

#1 – the NOAA GML website for near real-time Summit weather data is available both inside and outside of camp at the URL: https://gml.noaa.gov/dv/site/sum/met_all.html

#2 - the Summit Weather html display is an html file generated by a python script at /usr/lib/python2.4/site-packages/sri/weatherstation/NOAAMet.py. The data it uses is the raw data provided by the NOAA met instruments (not drawn from Virtual Weather Station); however, the display is maintained by Battelle-ARO. The display is available both inside and outside of camp at the URL: https://conus.summitcamp.org/v1/weather. The display is shown in Figure 1.

#3 - plots on the ARSLS InSite website use data from the NOAA met instruments.  The display is maintained by Battelle-ARO.  The display is available both inside and outside of camp at the URL:  https://insite.arsls.org/insite2-devel/summitcamp/environment/weather/combined.  The web page display is shown in Figure 2.

#4 - the LCD display in the Berthing Module is located near the communications equipment in the building’s main hallway.  The LCD display is maintained by the science techs.

Figure 1: Summit weather display.

Figure 2: ARSLS InSite weather web page.

Procedure:

The displays should be checked daily for rough data accuracy and updated time.

The Berthing Module met instruments are considered the secondary, non-official instruments; however, it is often a convenience to the community to keep the instruments clear of rime so that the Berthing Module LCD read-out indicates accurate data.  It is also important to keep them clear of rime so that they do not malfunction.  The anemometer can read lower than accurate wind speeds if it is not clean, which can lead to confusion about when north wind conditions exist.  The met instruments on the Berthing Module roof should be cleaned periodically as conditions require.

# NOAA – Surface Ozone Monitoring – AWO

Document revision number: 1.1

Document revision date: 9/6/2022

Primary contact:  	Peter Effertz, NOAA GML, Boulder, CO

720-314-8380, peter.effertz@noaa.gov

Secondary contact:  	Irina Petropavlovskikh, NOAA GML, Boulder, CO

303-497-6279, irina.petro@noaa.gov

Christy Smith, NOAA GML, Boulder, CO

303-909-0375, Christine.smith@noaa.gov

How often:		Daily – Check

Weekly – calibration check

Six Months – replace particle filter (appx January 1st and July 1st)

Clean requirements:	 	None

Supplies needed: 		Ozone analyzer, timer, particle filters

Introduction: Ozone is a central player in tropospheric chemistry. It is not, however, preserved in snow and ice. In the Arctic there are three separate issues of concern. Stratospheric ozone is being depleted with the most severe depletion taking place in the polar vortices in the Antarctic and Arctic. Over the Arctic sea ice ozone is depleted to near zero levels in the spring near the surface, through complicated chemistry that is not fully understood. In between, in the troposphere, ozone precursors emitted in combustion and other industrial processes lead to enhancements in ozone. As a remote site, Summit provides information on the possible longer-term hemispheric scale alteration of ozone in the troposphere due to human activity. Ozone also provides a context for understanding other atmospheric measurements made at Summit.

Daily Checks:

- Check inlet on the AWO inlet tower (white and orange cylinder) for snow or frost.  Snow can be cleared with gentle tapping of the inlet.  Be careful not to damage the filter in the inlet housing.

- Check the ozone level on the front panel.  The ozone level will vary between approximately 30 ppb to 70 ppb.

- Check the engineering data per the AWO Dailies spreadsheet.  These values are displayed on the aerosol CED laptop within the TEI box.    The following are guidelines for normal ranges:

  - Bench Temperature: 20-40 C

  - Pressure: > 400 mb

  - Flow: 0.4-0.9 lpm

  - Intensity:  65,000-120,000Hz

Weekly Calibration Check: 	(Updated 2022-08-29)

- Enter a comment into CPD3 that weekly surface ozone calibration check is starting

- Push “ENTER” to operate analyzer in “LOCAL” mode.

- Turn on the external pump that forces air though the charcoal tower zero air source into the inlet marked “Zero Air” on the back of the ozone analyzer.

- Push the “RUN” button. This will activate the solenoid so that you will now be sampling from the zero-air source. This display will say “ZERO”

- Allow the analyzer to operate for 15 minutes in this mode. Record ozone level from the display. Display should indicate approximately 0 parts per billion.

- Push the “RUN” button. This will set the internal ozone generator to the “LEVEL 1” position. The display will say “LEVEL 1”.

- Allow the analyzer to operate for 15 minutes in this mode. Record the ozone level from the display. Display should be approximately 160 parts per billion.

- Push the “RUN” button. This will set the internal ozone generator to the “LEVEL 2” position. The display will say “LEVEL 2”.

- Allow the analyzer to operate for 15 minutes in this mode. Record to zone level form the display. Display should indicate approximately 40 parts per billion.

- Push the “RUN” button. This will return the analyzer to sampling outside air. The display will say “SAMPLE”.

- Turn external pump off and push “ENTER” to return analyzer to “REMOTE” mode.

- Enter a comment into CPD3 that the weekly calibration check is complete

- Enter the ppb data for Zero, Level 1, and Level 2 into AWO Dailies spreadsheet in the comment section of the Surface Ozone tab.

Bi-Annual:

- Change the particle filter in the inlet housing every six months. This should be done on approximately January 1 and July 1. Change the filter more frequently if it has been damaged, or after a prolonged period of North Wind equipment operations. Spare filters are stored in the “NOAA aerosol spare and repair” box in the AWO vestibule.

Other:

- Check that the inlet is free of snow.

- The analyzer’s clock should be set to GMT and kept within one minute of an accurate time source. To set the time:

  - Put the instrument in “LOCAL” mode by pressing “ENTER”

  - Press “MENU” and then select “Instrument Controls”

  - Select “Time” and use the cursor arrows to adjust the time

  - Press the “ENTER” button when UTC time is 0 seconds after the time entered.

Consumable Resupply Items:

Conditioned particulate filter: ~2 per year

Shutdown Procedures:

The surface ozone instrument is supported by the UPS and can stay running during power outages; however, if shutdown is required for some reason:

Turn off power using the on/off switch on the front of the TEI instrument.

If this will be an extended outage (more than a day), cap the inlet port into the instrument. Be sure to uncap after turning the instrument back on! Also after an extended outage, perform the weekly instrument calibration check.

TEI Ozone Calibration Check Quick Reference

- Press ENTER => display shows LOCAL

- Turn on external pump

- Turn valve on front panel to ZERO

- Press RUN => display shows ZERO

- Operate for 15 minutes

- Note ozone level (should be approx 0 ppb)

- Turn valve on front panel to CAL

- Press RUN => display shows LEVEL1

- Operate for 15 minutes

- Note ozone level (should be approx 160 ppb)

- Press RUN => display shows LEVEL2

- Operate for 15 minutes

- Note ozone level (should be approx 40 ppb)

- Press RUN => display shows SAMPLE

- Turn off external pump

- Press ENTER => display shows REMOTE

# Battelle-ARO – Elevation and Altimeter Reporting

Primary contact:	Sam Dorsi, 406-461-5298, sam@polarfield.com

Introduction:

For communication to flight crews, an altimeter setting parameter is provided on the Battelle ARO weather display. This is a calculated parameter. An atmospheric pressure reading is made by the NOAA GML meteorology equipment at the base of the Met Tower. This atmospheric pressure value, along with the elevation at which the pressure was measured, are used to calculate an altimeter setting using a script in the Battelle ARO weather display software.

If the altimeter setting is unavailable on the weather display, it can be calculated using an atmospheric pressure measurement and accurate information about the elevation at which the pressure is measured or ‘station height’. To calculate altimeter setting, retrieve a pressure measurement and a corresponding orthometric station height, and utilize this recommended US National Weather Service website:

https://www.weather.gov/epz/wxcalc_altimetersetting

The current value for the elevation of the NOAA GML sensor is typically posted near the altimeter in the Big House. The value is approximately 3209.5 m AMSL or 10,530 ft AMSL, however the posted up-to-date value must be used.

Note that atmospheric pressure depends strongly on elevation. So, if pressure measurements for this purpose are made at locations different than the NOAA GML sensor at the base of the Met Tower, the posted ‘station height’ elevation should not be used, and a high-accuracy elevation measurement at the location of the pressure measurement should be made. Note too that GPS elevations are typically reported as height with respect to a reference ellipsoid, while flight applications require height with respect to a reference geoid (orthometric height = GPS ellipsoidal height - geoid height). The geoid height at Summit Station is 44.129 meters (EGM96).
Battelle-ARO – Weather Observations

Primary contact: 	Sam Dorsi, 406-461-5298, sam@polarfield.com

How often:		Three times daily

Clean requirements:		None

Supplies needed:		Possibly binoculars, headlamp/flashlight, dark fabric

Introduction: Weather observations are taken at Summit three times daily at 0000UTC, 1200UTC, and 1800UTC.  Observations are based on METAR protocols.  The local time of the observation changes during the year when Summit Station uses Daylight Savings Time.  Data from these observations are used by science PIs to understand the context for their own observations.

Procedure:

At the time of the weather observation, three things must be recorded by the observer:  horizontal visibility, present weather, and sky condition.  METAR protocol should be consulted for specifics on the reporting of these conditions, but this procedure will highlight some of the Summit-specific tips and procedures.  The other information in the observation is available from the NOAA weather station data:  wind speed and direction, pressure, and temperature.

Horizontal Visibility

Prevailing visibility (visibility for over half the horizon) is reported in meters.  There are station landmarks as well as mile markers that can be utilized to determine visibility.  If a marker can be observed, even using binoculars, then it is within the horizontal visibility.  To the north, markers are at 1 mile, 2 miles, and 3 miles.  To the east, mile markers are at ½ mile, 1 mile, and 3 miles.  Visual aids are listed in the table below:

| Station Landmark | Distance from Berthing Mod |
| --- | --- |
| Big House or Shop | 100m |
| MSF | 300m |
| AWO | 750m |
| 1st East Visibility Marker | 800m |
| 1st North & 2nd East Visibility Marker | 1600m |
| 2nd North Visibility Marker | 3200m |
| 3rd North & East Visibility Marker | 4800m |

In dark conditions, the light at AWO can also help determine visibility.  In an internet browser, access the following IP address: 192.168.1.89 and log in using the default settings.  To turn on the light, click ‘All outlets on’ under the Master Control.  The light can be turned off by clicking ‘All outlets off.’  The strength and clarity of light can be used to help estimate visibility.  If the AWO light does not help and it is too dark to make a true estimate, either code ‘Too dark’ or make the best guess possible using + or – to help define your confidence (800+m or 400-m).

If visibility is unlimited (to the horizon), code ‘9999’ as the visibility.

Present Weather

The present weather section of the observation refers to any precipitation on station or in the vicinity of the station.  Using a dark surface such as a black glove, dark fabric on a sleeve, or 3-ring binder will help to determine if there is anything in the air.

- Precipitation:

Precipitation is coded with an indicator of intensity.  The intensity indicator is chosen based on the visibility created by the falling precipitation:

| Precipitation Intensity | Visibility |
| --- | --- |
| “ - “ (minus sign) | > ½ mile |
| “ “ (no symbol) | ¼ - ½ mile |
| “ + “ (plus sign) | < ¼ mile |
| “SH” | Showers – precipitation starts and stops rapidly, often associated with rapid changes in sky condition |

After the precipitation intensity, precipitation type is coded.  Below are the precipitation types commonly reported at Summit.  Other types can be found in the METAR binder in the tech office:

  - SN – Snow crystals: branched, six-sided dendrites

  - SG – Snow grains: very small, white grains of ice

  - IC – Ice crystals: unbranched crystals in form of needles, columns, plates

- Obscurations:

Obscurations are not coded with intensity indicators. The first two obscurations are caused by wind:

  - DRSN – Drifting snow: winds raise snow into the air less than 6ft

  - BLSN – Blowing snow: winds raise snow into the air greater than 6ft

The other common obscurations at Summit are variations of suspended water and ice droplets in the air.  They differ in their ability to reduce visibility:

  - FG – Fog: suspended water droplets when temperatures are above 0ºC that reduce visibility to less than 5/8 mile

  - tsi – Freezing Fog: suspended water or ice droplets when temperatures are below 0ºC that reduce visibility to less than 5/8 mile

  - BR – Mist: suspended water or ice droplets at any temperature that reduce visibility  from 5/8 to 7 miles

Fog and mist can also be reported as ‘patchy’, ‘partial’, or ‘shallow’.  See the METAR binder in the tech office for more information.

- Sky Condition

Two observations are coded for the sky condition: amount of clouds and height of clouds.  The amount of cloud cover is reported in ‘oktas’, or eights of the sky.  The okta for each cloud height observed is reported, in order of height.  Each layer of increasing height includes the obscuration of lower height.  For example, if you have scattered clouds at 1000 ft, you cannot code few clouds at 2000 ft; you can only code scattered or greater oktas.

| Oktas of sky | Code | Plain English |
| --- | --- | --- |
| 0/8 | SKC | Sky clear |
| 1/8-2/8 | FEW | Few clouds |
| 3/8-4/8 | SCT | Scattered clouds |
| 5/8-7/8 | BKN | Broken sky |
| 8/8 | OVC | Overcast sky |
| 8/8 | VV | Vertical Visibility |

Accompanying the okta of sky obscured by each cloud layer is the height of the clouds.  Cloud height is reported in 100’s of feet after the okta designation.  A helpful resource is the data taken by various ICECAPS instruments.  This data are available to be viewed in real time using VNC to the computers running the various instruments.  Use any VNC viewer to access the following instruments:

192.168.1.166	 	Micro Pulse Lidar (MPL)

192.168.1.167 		Ceilometer

Each instrument shows the cloud heights in a different format and with different units and each has its drawbacks and benefits depending on the fog or cloud conditions.

The data are also available as a daily time series after a day or two at:

http://www.esrl.noaa.gov/psd/arctic/observatories/summit/browser/

Putting it together, sky condition is coded as XXXyyy, where XXX is the code for the oktas of sky obscured by a cloud layer and yyy is the height of the cloud in hundreds of feet (dropping the zeros after the hundreds).  For example:

FEW005  BKN010  OVC050

translates to: Few clouds at 500 ft,  broken skies at 1000ft, and overcast sky at 5000 ft.

‘000’ can be used for the height for obscurations at the surface.  It must also be reported in the remarks.  See  in the Remarks section below.

‘VV’, or vertical visibility is used when the cloud height cannot be determined, such as when there is fog, blowing snow, and other obscurations at the surface.  In this case, METAR defines vertical visibility as the distance one can observe into the indefinite ceiling (look at the towers, check out the distance you can see the LIDAR beam, or how long you can see the radiosondes after they are launched).

- Cloud Type

The reporting of the cloud type is up to the observer.  At a minimum, the basic types of clouds for each coded cloud height (stratus, cumulus, cirrus, etc.).  If the observer wants to be more specific, the Weather Identification Handbook located in the tech office is a fantastic resource for determining specific cloud types and species.

- Remarks

The Remarks section gives the observer the opportunity to explain an event they coded in the main observation, or add other observations they find interesting or useful.  For example, if FZFG and FEW000 were coded in the main observation, it would useful for the observer to write “FZFG FEW000” to indicate that freezing fog was the reason part of the sky was obscured at the surface.  Other common remarks are as follow:

- CIG THN – Thin ceiling, indicated the clouds are thin and you can clearly see the   sun/moon through the clouds

- SUN DMLY VSBL – Sun is dimly visible through the clouds/obscuration

- MOON DMLY VSBL – Moon is dimply visible through the clouds/obscuration

- AURBO – Aurora borealis visible!

(See the METAR binder in the tech office for more possible comments)

- Wind, Temperature, Pressure

The wind, temperature, and pressure data are gathered from the NOAA weather displays or the Met station data.  For convenience, a macro in Excel has been developed to automatically pull out a day’s data.  The macro is saved in an excel workbook called personal.xls and located in the Tech folder on the FTP drive.  This workbook is opened every time any excel workbook is opened so that the macro will always be available.  To use the macro, open the AR* file in excel and run the macro by pressing CTRL-W.  The macro will automatically create a new sheet that only contains the data needed for the three weather observation times.  If this is done after 1800UTC, all the data will be there and can be cut and pasted into the Weather Observations Excel workbook.

For the weekly report:

- In PowerShell:

cd  "C:\Users\ARO-Summit-Technicia\Polar Field Services\PFS-ARC-Summit Techs - Documents\Reports\weekly_wx\"

python ./weekly_wx.py

- The resulting report and downloaded NOAA minute weather data files can be located in the Tech One Drive/Reports/weekly_wx/

- For ICECAPS 1200z weather data use

python weekly_wx_icecaps.py

# Battelle ARO - Mobile Science Facility Protocol

Primary contact: 	Sam Dorsi, science project manager, sam@polarfield.com

Secondary contacts:	Chris Crane, Greenland maintenance lead, chris.c@polarfield.com

How often:		Daily: Daily checks

Weekly: None

Clean requirements:		None

Supplies needed:		Various

Introduction:

The Mobile Science Facility (MSF) provides infrastructure for the ICECAPS instrument suite and several other projects. The on-site science techs are responsible for monitoring the facility, addressing issues that arise, and communicating with station manager and facilities personnel when appropriate. This includes:

-- Identifying and communicating any hazards or changes to the facility.

-- Monitoring the power supply and Uninterrupted Power Supply (UPS) and notifying facilities personnel in the event of a system failure.

-- Maintaining a stable temperature inside the main room of the building.

-- Keeping ventilation ducts and exhaust/intake vents clear of snow.

-- Shoveling snow to preserve access to the main doors and emergency exit door at all times.

-- Clearing snow from the MSF roof deck.

-- Recording daily temperatures and HVAC parameters, and reporting on weekly basis.

-- Performing regular housekeeping (wiping surface, vacuuming, organizing tools and storage, etc.) to ensure MSF is a safe, functional, and presentable workspace.

In addition, several intermittent tasks are performed on an as-needed basis. These tasks are outlined in the protocol documents at: \PFS-ARC-Summit Techs - Documents\Facilities

-- Leveling the MSF as dictated by PI science needs.

-- Assisting in periodic relocation of MSF facility

-- Maintaining MSF network connection, in consultation with station manager and SRI.

The MSF is heated primarily from instrumentation waste heat. Two electric heaters offer backup heat. Cooling is provided by a fresh air handling system and by a heat recovery ventilator.

The temperature in the MSF must be maintained between 55° and 63° F. This temperature represents a compromise between PI desires for warmer (MMCR) or cooler (PAERI) conditions for their instruments.

UPS-sustained power is supplied to the instrumentation by a facility-integrated UPS; there is a mixture of clearly labeled UPS outlets and unsupported outlets in the MSF.

Procedure:

Daily Checks

- Verify that no power outages have occurred in the MSF, note any outages and their duration in the ICECAPS weekly report, and the science weekly report.

  - Event logs for the UPS, including outages, can be accessed at 192.168.1.32

- Verify that the temperature in the MSF is between 55° and 62° F (56°-66° per Sam Dorsi 10/29/2022); the temperature is displayed on the thermostat located on the wall between the MPL and CAPABL, or by using the X-300 (see below). Address any discrepancies:

  - Read the MSF Ventilation SOP F-011 for adjustment instructions.

  - Consult with the MSF facilities contacts before any non-emergency changes.

- During warmer periods, verify that the fresh air cooling system is operating properly by observing the following:

  - Periodically, the system will switch on, the fan will begin running, the exhaust damper (north wall) will open, and the mix air damper will close down somewhat.

  - During operation, the exhaust damper and the outside air damper (not visible) will open and close in unison, and the mix air damper will move in the opposite sense of the other two.

  - When finished cooling, the fan will shut down, the mix air damper will open completely, and the other two dampers will close completely.

- Verify that the two west wall electric heaters are set to “Heat” and that the thermostats are set correctly (typically 55°).

  - The east wall electric heater must always be OFF; or it will overheat the MMCR.

  - If either heater is running, this is an indication of out-of-balance HVAC systems. Consult the MSF facilities contacts.

  - Verify that the heat recovery ventilator is operating correctly by confirming airflow out of the central air duct in the main room and return airflow into the vent just above the electrical panels in the main room.

- If there has been recent snow accumulation, or a blowing snow event:

  - Clear the porches and doors of snow

  - Shovel snow from stairs to allow safe access to the building.

  - Verify that the emergency exit door operates and is not frozen shut

  - Check the four intake and exhaust vents for blockage

  - Sweep the roof clear of snow

  - Check for and clear the HVAC system filters of any snow that may have accumulated.

- Complete the daily entry in the MSF Facility Temperatures log at ‘\PFS-ARC-Summit Techs - Documents\Facilities\Daily Logs\MSF_HVAC_log.xlsx’.

Monitoring the Temperature (X-300)

The MSF is equipped with a Control-by-Web X-300 thermostat/temperature monitor and logger and three thermocouples at various locations inside the main room of the MSF. The X-300 can be accessed at 192.168.1.181. Some notes on the X-300:

- Username and password are in the password document

- Access the setup for the unit at 192.168.1.181/setup.html

- “Relay 1” is connected to a signal-over-AC relay located at an outlet below the air handler; it controls two unused outlets on the south wall of the vestibule

- Access logged data for the last several weeks as comma-separated-values at 192.168.1.181/log.txt

Adjusting the MSF Main Room Temperature

The MSF Facility Ventilation SOP (F-011) is the primary resource for this system and should be consulted before changes are made. The following outline is a general familiarization tool.

Maintaining a steady temperature in the MSF main room demands careful consideration of the many heat sources (instruments and electric heaters) and cooling mechanisms (fresh air handler, ventilator, and the ambient weather). All changes should be recorded in the MSF Facility Temperature log, and the effect on the facility should be monitored closely. Adherence to protocols and consultation with facilities personnel will minimize disruption to science equipment.

All changes to thermostats or ventilation settings must be made in consultation with Sam Dorsi.

- Heating sources

  - Instruments

    - Waste heat from instruments is the primary heat source for the MSF.

    - Be aware of possible room temperature changes if an instrument is switched on or off for a long period

  - Electric heaters

    - The west wall heaters are typically set to 55°.

    - These are considered a backup; if found to be running, consult the facilities contacts for recommended HVAC adjustments.

    - Never switch on the east wall heater.

- Cooling sources

  - Fresh air handler

    - The air handler controls the mixing dampers that admit outside air in warmer weather, or recirculate indoor air in cooler weather.

    - The SOLO controller, below the air handler, acts at thermostat.

  - Heat recovery ventilator

    - The heat recovery ventilator provides fresh air circulation in the main room with minimal heat loss through use of a heat exchanger. Due to inefficiencies, this unit can “overcool” the MSF in cold conditions.

    - The ventilator fan speed and duty cycle are adjustable on the west wall controller, between the door and the electrical panels.

    - In summer, set the ventilator fan speed to 5. In winter, the fan speed may need to be decreased to 3 or 4.

  - Doors and Windows

    - Except in the most extreme emergency, neither doors nor windows should be used to admit cool air, due to the high risk of damage to science equipment.

    - When entering or exiting the MSF, use the west vestibule double-door system as an airlock to prevent cold air drafts.

    - Do not use the east wall exit except during emergencies.

Monitoring and Controlling the UPS

The MSF is equipped with an Eaton Powerware 9355 10 kVA UPS. This system provides backup power to a subset of the instruments.

- The outlets around the MSF are clearly labeled as being “UPS” outlets, or normal unsupported outlets

- Certain instrument components should not be connected to UPS outlets, as they draw too much power and would drastically reduce UPS runtime:

  - MWR blowers/heaters, webcam

  - Any appliances in the room

  - Lights

- Monitor all data and events pertaining to the UPS using its front control panel, or the web interface at 192.168.1.32; see password sheet for login information. The event log and status menu on the UPS front panel LCD contain more information than the web interface; this front panel should be consulted when there are concerns regarding the UPS.  Periodically verify that:

  - the status is “System Normal” on the web interface and on UPS

  - the event log on the web interface and the UPS don’t contain strange messages. “Charger on” and “Charger off” messages are normal.

  - the runtime is between 45 – 60 minutes

- The MMCR is plugged into the Eaton UPS, but also has its own rack-mounted UPS.

- Consult the power outage information in the OneNote Tech Guide for info on how to handle power outages outage and protect the sensitive PAERI Stirling cooler.

# CRREL - Station Accumulation/Drift Survey

Primary Contact: Ian Geraghty, (970) 946-6771, ian.g@polarfield.com

Secondary Contacts: 	Zoe Courville, zcourville@gmail.com

Deeb, Elias, Elias.J.Deeb@erdc.dren.mil

Deeb, Elias, Elias.J.Deeb@usace.army.mil

How Often:	Twice a year:
Spring (typically late March) before major drift mitigation occurs
Fall (early September) before major winter storms cause drifting

Introduction: Science technicians at Summit Station have conducted GPS-based surface-elevation surveys bi-annually in the vicinity of the main buildings since 2007.  CRREL personnel compile the data and generate maps of snow accumulation and drift mitigation to support station drift management activities.

Survey Objectives

- Document drift patterns around the main buildings

- Estimate annual winter snowdrift accumulation volumes

- Provide guidance on preferred locations of buildings to minimize maintenance owing to snowdrifts

- Provide guidance on building layout at proposed Isi station to minimize maintenance owing to snowdrifts

- Acquire datasets to validate snowdrift simulation models for building design and station layout at Summit and Isi

Supplies needed:

- Trimble equipment for the accumulation survey:

  - Trimble backpack

  - 12’’ extension pole for antenna

  - Trimble Zephyr GPS antenna

  - 12V battery for receiver

  - Trimble R7 roving receiver

- Handheld Garmin GPS to mark survey corners

- Skis or snowshoes

Survey Methods

The survey is performed using a Trimble R7 GPS receiver and a base station.  Personnel walk the length of the survey using the roving antenna with the backpack setup.  The data are high quality but their usefulness depends on the area surveyed and the data spacing.

A serpentine pattern offers a good approach to execute the survey on foot.  Based on past surveys, we recommend

- 5-m data spacing along the direction of travel

- 5-m spacing between adjacent survey lines (perpendicular to the direction of travel)

This provides a compromise between resolution needed to meet the objectives and the duration (cold exposure time) to conduct the survey.

Because the objectives focus on the drift patterns near Summit’s main buildings, we recommend limiting the extent of the survey to that area.  Figure 1 shows the recommended survey area and approximate serpentine pattern.  Tables 1 and 2 provide the corner coordinates.  Survey lines across the direction of the main snowdrifts provide good definition of those drifts, hence we recommend generally east-west line orientation.

This survey requires you to walk over snowdrifts and uneven terrain, so please wait for a day with good visibility! Avoid days with flat light. Avoid windy days so you can see your tracks during the entire survey.

Procedure

Put the 12-volt battery used for ICESAT on the charger the night before you intend to run the survey.

Measure the distance (in centimeters) from the bottom of the antenna mount to the ground while the person who will be conducting the survey has the backpack on (see below for instructions). It is important that the measurement take place while the tech is wearing the gear (boots especially) they will have on during the survey, and the backpack is fit to comfort.

Verify the proper Survey Configuration file is on to the R7 receiver. Note that the configuration file is the same as used for ICESat.

Connect the R7 receiver into the science tech laptop in the MSF using the power and comms cable plugged into the R7 Port 2 and a comm port on the PC.

Power on the receiver by pressing the green button.

From the start menu, launch “Configuration toolbox” (Start > All Programs >Trimble >Configuration Toolbox > Configuration Toolbox)

Select Communications -> Get File…

From the dialog box, select “power_up” and click “Get File”

Select File -> Open…

In the dialog box choose the file “ICESat_1sec_ppk”. If the file is not present, it can be copied from “C:\Users\ARO-Summit-Technicia\Polar Field Services\PFS-ARC-Summit Techs - Documents\Instrument Manuals and Supporting Documents\Earthscope\Receiver Configuration Files”.

Click “Open”

Verify the file settings are as indicated in the ICESat protocol

Click “Transmit”

If the dialog comes up asking if you want to replace the file “power_up,” select “Yes”

You should see the message “the application successfully transmitted to the receiver”- click “OK”

Click “Save”

Click “Close”

Note: If two different people perform the survey, be sure to record both height measurements and which surveys these heights correspond to.

Depending on battery capacity and state of charge, you may wish to place the battery back on the charger between legs of the survey.

Set up the backpack as pictured below:

Connect the battery to Port 2 of the R7, being careful to line up the red marks from the cable to the R7 port. This connection can come loose quite easily, so placing a piece of tape to hold it may be helpful.

Attach the extension pole to the outside of the backpack.

Attach the Trimble Zephyr GPS antenna to the top of the extension pole.

Connect the antenna to the R7 using the yellow TNC antenna cable, running the cable through the top of the backpack. The 90 degree end attaches to the antenna. Make sure there is plenty of slack so neither end is stressed.

When the connections have been made, place the system upright outdoors with a good sky view to initialize. Before starting work, confirm that that LED status lights on the receiver show the presence of satellite signals and that data are being saved to the compact flash card—both the red satellite LED and the orange data LEDs should be blinking. See the ICESat Traverse protocol for a quick reference sheet for these status lights.

If it is very cold outside, hot water bottles can be added to the pack to extend the life of the battery.

GPS Antenna connection

Port 2 battery connection

It is advisable to make a layout of the survey prior to starting to make navigating easier. You may do this by marking the four corners with flags and driving the perimeter with two parallel snowmobile tracks or adding intermediary flags for visual reference of the perimeter of the survey.  It may be helpful to use two additional flags along each N-S perimeter line to keep track of your progress and to provide a visual reference for walking in relatively straight lines.

After you have a good idea of the route and the survey gear is set-up, you can begin the survey. Although the area is rather contained, expect the survey to take 6-8 hours.

When you are finished with the survey, process the data in the same way you would for an ICESat survey:

Convert all .t01 files to .dat using .\runpkr. Verify the .dat file size(s) are as expected, totaling 5000-10000+ KB, depending on how fast the survey was conducted.

Convert the appropriate SMM4 base station data files to RINEX.

Save the survey in a new folder named “YYMMDD_camp_survey” under \ftp\operations\GPS\Accumulation_Surveys

Transfer the base station GPS files for the survey period into a folder in this directory.

In the same folder, create a metadata file with technician names, antenna heights and basic information on survey sessions, for example:

Summit Station Accumulation Survey

Antenna height information

2018-09-29

Performed by Mara Menahan and Hannah James

First session: Mara Menahan

File 98762680.T01

Start Time: 13:11 UTC

End time: 13:59 UTC

Antenna height (measured from snow surface to base of Zephyr antenna, when carrying survey backpack): 172 cm

Email the science project manager to let them know that the data are processed and available on the FTP.

Science Project Manager will send the FTP link to the survey data to the appreciative stakeholders.

Figure 1. Recommended survey area and serpentine pattern.  Data points should be 5 m apart along the direction of travel and lines should be 5 m apart perpendicular to travel.  The box is ~ 330 m x 230 m.

| Survey corner coordinates in Degrees and Decimal Minutes |  |  |
| --- | --- | --- |
| Corner | Lat | Long |
| NW | 72° 34.904 | -38° 27.677 |
| NE | 72° 34.864 | -38° 27.285 |
| SE | 72° 34.696 | -38° 27.474 |
| SW | 72° 34.735 | -38° 27.864 |

| Survey corner coordinates in Decimal Degrees |  |
| --- | --- |
| Corner | Lat/Long |
| NW | N72.58173 W38.46128 |
| NE | N72.58107 W38.45475 |
| SE | N72.57827 W38.45790 |
| SW | N72.57892 W38.46440 |