ICECAPS Weekly Report

Oct. 7 – Oct. 13, 2012

Lance Roth – Phase II Tech

General Items

• Turnover was completed and Phase I has departed.

Dataman Account:

• Normal

PAERI:

• 11/9, Partial blue flags

• 11/10, Partial blue flags

• 11/13, Fan quit working.  Will replace ASAP.

MPL:

• 11/7, 12:43:05, "AMCS Timeout", "Can't Open AMCS Card: 2". Reboot. Measurements resumed at

14:17:37.

• 11/10, 14:53:00, Found Computer Logged Outed. Measurements resumed at 14:53:00.

• 11/12, 17:45:00, Found Computer Logged Outed. Measurements resumed at 17:45:00

CAPABL:

• Normal

MMCR:

• Normal

MWR:

• 11/11, 17:30:00, Receiver 2 Temp Stability was Red.

ASIA-A:

• 11/7, break in communication w/radiometer at 1740 UTC. Power cycle radiometer, restart software. Back up at 1832 UTC.

• 11/8, software frozen at 1123 UTC. No computer control, ctl-alt-del will not work. Manually turned off computer w/power button. Computer will not reboot. Contacted Pierre.

• 11/9, Replaced old computer with spare computer. Reconnected to radiometer and currently collecting data. Elevated new computer to increase ventilation. 1445 UTC.

SODAR:

• Normal

POSS:

• Normal

VCEIL:

• Normal

IceCAM:

• Normal

IcePIC:

• No snowflake pictures taken this week.

Radiosonde:

• Normal