ICECAPS Weekly Report

Nov. 21 – Nov. 27, 2012

Lance Roth – Phase II Tech

Interesting weather events:

• 11/21 Little to No Winds, Cloud, and clear.

• 11/23 High Winds exceeding 25 knots

• 11/26-27 Warm temps (-8 F)

General Items

• 11/27 MSF extremely warm (80.0 F near PAERI).  Turned down heaters and increased air flow.

Dataman Account:

• Normal

PAERI:

• 11/21 At 1220 noted. LW HBB NEN, SW HBB NEN, LW Responsivity, and SW Resoponsivity.

Scene mirror  w/ Light Snow. Brushed snow off mirror 1240 UTC.  Flagged returned to Green.

• 11/25 Partial Blue Flags

• 11/27 Partial Blue Flags

MPL:

• 11/23, 1641 UTC, “Can’t find AMCS Card.  Restarted MPL at 1644 UTC.

CAPABL:

• Normal

MMCR:

• Normal

MWR:

• Normal

SODAR:

• Normal

POSS:

• Normal

VCEIL:

• Normal

IceCAM:

• Normal

IcePIC:

• No pictures taken this week.

Radiosonde:

• Normal.

ASIA-A:

• 11/23 Lost connection at 1939 UTC.  Re-established connection at 0124 UTC

• 11/26 Lost connection at 1322 UTC.  Re-established connection at 1327 UTC

Lost connection at 1451 UTC.  Re-established connection at 1453 UTC

• 11/26 Trouble shooting Arduino camera.