# ICECAPS Weekly Report

November 18 – November 24, 2013

John Lyons

TAWO floodlight, 11/22 12Z (roughly)

General:

- NOAA/TAWO Met 2m temperature sensor down from 11/21 – 11/24 due to programming of Zeno DAS. NOAA webpage also down 11/23 – 11/25. Sonde ground check data from GH tower, supplemented with Swiss Tower RH data.

- Electric wall heaters not switched on yet, as warmer weather has required cooling rather than heating

- MSF facility operating normally.

Significant Weather Observations:

- 11/18: Winds 174 at 6 kn, -48C, 61%, SCT, SN from very thin cloud in late afternoon and evening, large halo around moon

- 11/19: Winds 310 at 9 kn, -49C, 60%, SCT-OVC

- 11/20: Winds 237 at 21 kn, -19C, 90%, SN/BLSN, moderating at midday, but picking back up at night

- 11/21: Winds 281 at 17 kn, -35C, 67%, DRSN, overnight and early AM was higher winds (gusts to 30+ kn) but moon visible

- 11/22: Winds 243 at 14kn, -46C, 63%, FEW, diamond dust overnight, plus vertical cross haloes around moon and TAWO floodlight in AM

- 11/23: Winds 194 at 14kn, -43C, 65%, SCT, some diamond dust overnight, above average riming this morning

- 11/24: Winds 227 at 25kn, -20C, 80%, SN-BLSN, occasional views upwards to moon

Dataman Account:

- Operating normally

MWR:

- 11/19: Calibration body fan on HATPRO frozen up and will not restart, even after cleaning

- 11/19: R2CH software froze up; Erik discovered and cleared

- 11/20: Recvr 2 Temp Instability on 90/150; not coincident with climbing roof (but high winds), cleared itself

- 11/24: Restarted HATPRO calibration body fan with help of wind

SODAR:

- Operating normally

- 11/19 Sodar cleaned of snow (mostly transmitter dish)

- 11/22 Sodar cleaned of snow

POSS:

- Operating normally

MMCR:

- 11/23: Contour display data found to be junk in AM (all vertical lines), all other checks OK. Three restarts attempted afternoon of 11/23; on third restart MMCR showed cyclical helix overcurrent fault; shut instrument down. Cleared to restart again 11/25 PM; successful.

- Operating normally

CAPABL:

- 11/20: Failed to startup again upon reboot.

- 11/22: Laser and Labview turned off

MPL:

- Operating normally

- 11/19: Afterpulse calibration

- 11/20: MPL mount to Dataman went down; remounted at 1300

VCEIL:

- Operating normally

IceCAM:

- Operating normally

PAERI:

- 11/20: LW and SW Responsivity alarms at 06Z and 13Z (at least). Self-clearing, probably due to SN-BLSN

- 11/21: Continuous LW and SW responsivity alarms from 0830Z (BLSN) until mirror manually cleared at 1645Z

- Operating normally

ASIA-A:

- 11/19: Communications error sometime between 04Z and 06Z, Pierre discovered and rebooted radiometer remotely.

- 11/20: 24hr stability testing initiated in early AM by Pierre

- Operating normally

TSI:

- Offline for winter

IcePIC:

- 11/18 2200Z: cruddy bullet rosettes, needles, hollow column, one triangular prism

- 11/20 1345Z: small prisms and rime, most poorly shaped

- 11/22 1400Z: mostly bullet rosettes, some columns, few triangular and stellar plates

- 11/24 1300Z: rime, prisms, poorly formed stellar plates

Radiosonde:

- Twice daily sondes