# ICECAPS Weekly Report

December 23 – December 29, 2013

John Lyons

New Ceilometer Jacket

General:

- MSF facility operating normally.

Significant Weather Observations:

- 12/23: Winds 104 at 8 kn, -40C, 69%, FEW-SN (see IcePIC – rimed crystals) with a 0.5 – 1km cloud layer between 15Z and 18Z, layer lowering to ground FZFG, then clearing

- 12/24: Winds 51 at 11 kn, -42C, 67%, FEW, new dense hoar/rime layers on NE side of most structures/poles/etc.

- 12/25: Winds 122 at 14kn, -41C, 68%, SCT

- 12/26: Winds 155 at 10kn, -44C, 66%, SCT, SN later in day

- 12/27: Winds 150 at 10kn, -45C, 63%, FEW

- 12/28: Winds 170 at 6kn, -52C, 55%, SKC, qualitatively the “clearest atmosphere of the phase so far”

- 12/29: Winds vrb, 0 – 3kn, -48C, 63%, SKC, new layer of hoar frost, 2-3m thick patchy ground FZFG moving around summit

Dataman Account:

- 12/23: ToBoulder clear after several days of backlogs

- 12/28: ToBoulder has 54 Icecam images remaining

- 12/29: ToBoulder clear once again

MWR:

- Operating normally

SODAR:

- Operating normally

- 12/29 1403 - 1410: Sodar cleaned

POSS:

- Operating normally

MMCR:

- Operating normally

CAPABL:

- No longer operating

MPL:

- Operating normally

VCEIL:

- 12/23 1420: Insulation and mylar jacket installed around instrument body

| Time | NOAA Ambient Temp | VCeil Outside Temp | Tdiff |
| --- | --- | --- | --- |
| Before Insul. | -40C | -35C | +5 |
| 12/23 1425 | -40C | -34C | +6 |
| 12/24 1300 | -42C | -33C | +9 |
| 12/25 1436 | -41C | -29C | +12 |
| 12/26 1400 | -44C | -20C | +24 |
| 12/27 1305 | -45C | -16C* | +29 |
| 12/28 1308 | -52C | -22C | +30 |
| 12/29 1340 | -48C | -21C* | +27 |

* Inside heater off, blower heater still on

IceCAM:

- Operating normally

PAERI:

- Operating normally

- 12/27 1600 - 1700: Motor end insulation installed and secured

| Time | Outside Air Temp | Mirror Motor Temp | Tdiff |
| --- | --- | --- | --- |
| Before Insul | 225.5 | 244.1 | +18.6 |
| 12/28 1315 | 220.9 | 238.4 | +18.4 |
| 12/29 1345 | 230.0 | 251.0 | +21.0 |

ASIA-A:

- 12/26 2145: Communications error; reboot and re-time-sync required

TSI:

- Offline for winter

IcePIC:

- 12/23 1725Z: Rimed rosettes and columns, rime clumps, ill-formed crystals, columns, prisms, maybe associated with a low, possibly mixed-phase(?) cloud

- 12/25 1515Z: All very small rime pellets, ill-formed clumps, columns, needles, few prisms

- 12/26

  - 1700Z: Small rime pellets, columns, needles, very fine needles, prisms

  - 1845Z: Rime pellets, larger rime clumps, bullet rosettes, columns, prisms. Fewer needles than 1700Z, more, larger particles. Coincident with SN and change to higher reflectivity region on radar

Radiosonde:

- 12/26 12Z K-sonde likely erroneous RH reading (RH declines monotonically to 50%, then steady for rest of sounding

- 12/27 12Z – Bad K-sonde RH measurement; same symptoms as 12/26 12Z

- Twice daily sondes