# ICECAPS Weekly Report

October 31, 2016 – November 6, 2016

Sam Dorsi

In anticipation of fan replacement, the MWR blower unit warms on the MSF bench. SWD.

General:

- 11/2: Network outage due to stateside issues with ISP, roughly 18:00-22:00z.

Significant Weather Observations:

- 10/31: Overcast stratus, freezing fog -33C, 10 kts.

- 11/1: Overcast stratus, blowing snow, -20C, 18 kts.

- 11/2: Overcast stratus, falling ice crystals (12:40-14:00z), -37C, 8 kts.

- 11/3: Few altostratus, -42C, 12 kts.

- 11/4: Few altostratus, freezing fog, -40C, 6 kts.

- 11/5: Overcast stratus, freezing fog, blowing snow, -18C, 18 kts.

- 11/6: Overcast stratus, blowing snow, -17C, 18 kts.

ICECAPS Data Management:

- 10/31: Data transfer complete at 15:35z.

- 11/1: Data transfer complete at 17:16z.

- 11/2: Data transfer complete at 11:51z.

- 11/3: Data did not transfer. This resulted from an issue with the NOAA server.

- 11/4: Data transfer complete at 23:48z. The backlog from yesterday’s issue was successfully transferred.

- 11/5: Data transfer complete at 13:33z.

- 11/6: Data transfer complete at 13:53z.

MWR:

- Operating normally.

- 11/1: The HATPRO heater fan was found inoperative at 12:57z. To reduce the chance of the unventilated heater causing damage, the heater was unplugged at 15:53z. Replacement fans have been located and will be swapped out.

- 11/2: The MWR blower was removed at 12:52z, and moved into the MSF for warming. The blower unit was disassembled, a new fan installed, connections soldered and heat-shrunk, and the unit reassembled. The fan was restored to service at 18:25z and the heater power restored at 18:27z. It sounds normal and appears to be drawing error well.

- 11/3: Using photos from previous day’s job, a step-by-step protocol for the fan replacement was written and placed in the /ICECAPS/MWR/resources/ directory.

SODAR:

- Operating normally.

POSS:

- Operating normally.

MMCR:

- Operating normally.

- Pitch 0.20, roll -0.15 as of 11/6.

CAPABL:

- Operating normally.

MPL:

- Operating normally.

VCEIL:

- Operating normally.

Hotplate:

- Offsite for repairs since 6/9/2016.

IceCAM:

- Operating normally.

PAERI:

- Operating normally.

- 11/3-4: Yellow alert flags related to lower indoor temperatures in the MSF.

TSI:

- Stored in MSF for winter.

IcePIC:

- 11/2: Slides placed for collection 12:48 – 14:18z. Particles photographed include sheaths, hexagonal plates and bullet rosettes.

Radiosonde:

- Twice daily soundings.

MASC:

- Offsite for repairs since 5/3/2016.