# ICECAPS Weekly Report

September 26, 2016 – October 2, 2016

Elissa Barris, Travis Guy

Reflected halo. –Elissa, 9/29

General:

- 9/26: Internet set to winter bandwidth

Significant Weather Observations:

- 9/26: Few Stratus, BR, -30C, 8kn, Visibility 1600m. Fog bow 1817z.

- 9/27: Ovc Stratus, BR, -18C, 8kn, Visibility 1600m. Parhelia 0954z. Halo 1815z.

- 9/28: Few Stratus, FZFG, -28C, 5kn, Visibility 700m.

- 9/29: Bkn Stratus, Sct Altocumulus, -24C, 7kn, Visibility 1600m. Halo 1200z and 1557z

- 9/30: Few Stratus (BR), Sct Stratus, Bkn altostratus, -28C, 12kn, Visibility 3200m. Halo 1151z and 1548z.

- 10/01: Bkn Stratus, -33C, 14kn, Visibility 1600m.  Halo and upper tangent arc 1806z.

- 10/02: Few Stratus (BR), Bkn Stratus, -30C, 9kn, Visibility 3200 in the morning becoming unlimited by midday. Increased clouds and winds in the evening with BLSN

Dataman Account:

- 9/26: Data transfer complete at 0556z.

- 9/27: Data transfer complete at 1116z.

- 9/28: Data transfer complete at 1317z.

- 9/29: Data transfer complete at 1141z.

- 9/30: Data transfer complete at 1456z.

- 10/01: Data transfer complete at 1207z.

- 10/02: Data transfer complete at 1153z.

MWR:

- Operating normally.

SODAR:

- Operating normally.

- Heavy equipment operating daily for berm construction.

POSS:

- Operating normally.

- 10/2: Heater current error in MESPOS and MAINLOG error log text files. First time I have seen this, but protocol says this is normal and acceptable.

MMCR:

- Operating normally.

- TWTA Cathode Voltage out of range throughout week.

- Pitch 0.11, roll -0.11 as of 10/02.

CAPABL:

- Operating normally.

MPL:

- Operating normally.

- VantageVue communication error (15:57) remains on log.

- “Missing calibration files: Overlap, Dead Time Correction” messages also present throughout week.

VCEIL:

- Operating normally.

Hotplate:

- Off site for repairs since 6/9/2016.

IceCAM:

- Operating normally.

- 9/26: Pictures through most of the day were white until frost was removed from the camera

PAERI:

- Operating normally.

TSI:

- Operating normally.

IcePIC:

- 8/XX Slides placed for collection XXXXXz. Images photographed include XXXXXXXX.

Radiosonde:

- Twice daily soundings.

- Weekly (or more) computer restarts seem to solve most software glitches.

MASC:

- Offsite for repairs since 5/3/2016.