# ICECAPS Weekly Report

January 9 – January 15, 2017

Sam Dorsi

With the sun returning, the sky grows ever brighter behind the MSF. SWD.

General:

- 1/9: Unplanned power outage from 23:07z to 23:08z.

- 1/10: Planned outages for generator swap from 12:09z to 12:10z, and again from 19:16z to 19:17z.

- 1/11-12: Network outage occurred ~22:00z to ~13:00z.

- 1/13: Planned outages from 12:47z to 12:48z, and from 14:43z to 14:45z.

- 1/13: MSF structure was re-leveled 17:15z – 18:00z. Pitch was changed from 0.33° to 0.01°, and roll changed from -0.27° to 0.00° (MMCR 1-hour averages).

Significant Weather Observations:

- 1/9: Clear, -50C, 14 kts.

- 1/10: Few stratus, -47C, 6 kts.

- 1/11: Scattered stratus, -52C, 6 kts.

- 1/12: Few stratus, -50C, 12 kts.

- 1/13: Broken stratus, -40C, 14 kts.

- 1/14: Scattered altostratus, -42C, 15 kts.

- 1/15: Blowing snow, -23C, 29 kts.

ICECAPS Data Management:

- 1/9: Data transfer complete at 12:23z.

- 1/10: Data transfer complete at 11:43z.

- 1/11: Data transfer complete at 11:28z.

- 1/12: Data transfer complete at 23:43z. Transfer was manually restarted after partial completion due to outage.

- 1/13: Data transfer complete at 15:57z.

- 1/14: Data transfer complete at 12:42z.

- 1/15: Data transfer complete at 11:37z.

MWR:

- Operating normally.

SODAR:

- Operating normally.

POSS:

- Operating normally.

MMCR:

- Operating normally.

- Pitch 0.33, roll -0.25 as of 1/9.

- Pitch 0.00, roll 0.00 as of 1/13.

CAPABL:

- Operating normally.

- 1/10: Laser 1 was fully powered down and its head unit was opened and inspected at 17:45z. Inspection revealed condensed water on many interior surfaces. An attempt was made to identify leaks by operating the cooling water pump with the head open; however, no leaks were apparent. Head was left open to dry.

- 1/11: Confirmed visually that Laser 1 head was dry, then powered water pump and examined for leaks inside laser head over the course of 2 hours. No leaks were detected.

- 1/12: Instrument was discovered in an error state (19:30z), with the .vi block diagram shown instead of the real-time display. Laser 2 was firing as normal, but the Licel was not triggering and the shot counter was not incrementing. LabView was shutdown, the Licel power-cycled, and the vi reopened. The system resumed normal operations at 22:23z.

MPL:

- Operating normally.

VCEIL:

- Operating normally.

Hotplate:

- Offsite for repairs since 6/9/2016.

IceCAM:

- Operating normally.

PAERI:

- Operating normally.

- 1/12: Throughout the day, PAERI exhibit cold-related status flags.

TSI:

- Stored in MSF for winter.

IcePIC:

- 1/13: Slides placed for collection 13:10 to 13:26z. Particles photographed include columns, hexagonal plates, bullet rosettes and needles, possible triangular ice, possible rimed particles.

Radiosonde:

- Twice daily soundings.

MASC:

- Offsite for repairs since 5/3/2016.