# ICECAPS-ACE Weekly Report

August 30th – September 5, 2022

Molly Peek

Gull-winged upper tangent arc, partial 22-degree halo, and sun pillar on the evening of 2022-09-04.

Photo taken by: Molly Peek

General Notes:

This week at the Mobile Science Facility (MSF), work was focused on fixing bugs and instrument errors. The Millimeter Cloud Radar (MMCR) experienced issues with the MMCR-DM computer, and there was no MMCR data transfer after 2022-09-01. The Raspberry Pi also had issues processing data and troubleshooting continues. After some small operating issues on 2022-09-03, the Dataman computer is successfully working and processing data. Summit has experienced anomalously warm temperatures since 2022-09-03, with temperatures intermittently above 0 degrees Celsius throughout the day on 2022-09-03, and hovering closely below freezing since then. All regular instrument checks and the twice-daily radiosonde balloon launches continued successfully.

General Weather Observations:

| Date | Time (Z) | Wind (kt) | Weather | Visibility | Sky | Temp (C) | Notes |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 2022-08-30 | 12:00 | 25 | BLSN, DRSN, +SN | 750 | OVC | -5 |  |
| 2022-08-31 | 12:00 | 12 |  | 1200 | FEW | -12 |  |
| 2022-09-01 | 12:00 | 5 | FZFG | 400 | OVC | -18 |  |
| 2022-09-02 | 12:00 | 11 |  | 9999 | SCT | -14 |  |
| 2022-09-03 | 12:00 | 17 | FZFG | 3000 | BKN | -1 |  |
| 2022-09-04 | 12:00 | 14 |  | 9999 | BKN | -2 |  |
| 2022-09-05 | 12:00 | 9 |  | 9999 | BKN | -3 | Unverified freezing rain in the night leading into 2022-09-06 |

Significant Weather Observations:

- Partial 44-degree halo; diamond dust 20220831:1106z

- Diamond dust 20220831:1200z

- Iridescence 20220831:1300z

- Fogbow 20220901:1900z

- Iridescence 20220902:0845z

- Partial rainbow 20220903:1127z

- Lower tangent arc 20220903:1240z

- Gull-winged upper tangent arc; partial 22-degree halo; sun pillar 20220904:2015z

MSF

- Operating normally

ICECAPS-ACE Data Management:

- Data transfer completed 2022-08-30:0400z

- Data transfer did not complete 2022-08-31

- Data transfer completed 2022-09-01:0447z

- Data transfer completed 2022-09-02:0419z

- Data transfer completed 2022-09-03:0558z

- Data transfer did not complete 2022-09-04

- Data transfer completed 2022-09-05:0427z

- Issue opening desktop folders and completing commands in the Terminal Emulator 20220903

TSI:

- Was partially iced during warm and cold weather fluctuations

- Operating normally

Radiosonde:

- Several Java script errors continue to pop up, does not appear to affect launch or data collection.

- Operating normally

MWR:

- Operating normally

SODAR:

- Amplifier turned off for cleaning 2022-09-05:1138z-1153z

- Operating normally

POSS:

- Regular transient noise errors

- Operating normally

MASC:

- Frayed power cord to MASC mount, notes have been made for repairs

- Operating normally

MMCR:

- MMCR-DM computer experiencing issues, the radar seems to be operating normally

- MMCR does not appear to be transferring data

MPL:

- Operating normally

CEILOMETER:

- Operating normally

PAERI:

- Expected to return in September

IceCAM:

- Operating normally.

IcePIC:

- Images taken 2022-09-04

- Camera issues affected photo quality 2022-08-31, resolved same day

MAGIC CPC:

- Operating normally

MSF - OPC:

- Raspberry Pi experiencing issues, Quicklooks pages are not accessible

- Operating normally

AWO - OPC:

- Offline

SKYOPC:

- Offline

Biral PWS :

- Bad self-check 2022-08-26 :2200z – 2022-08-27 :0200z

- Bad self-check 2022-08-29:0045-0115z and 0130-0200z

- Raspberry Pi experiencing issues, Quicklooks pages are not accessible

- Operating normally

Flux Tower instrument suite:

- Upgrades underway

- Operating normally

Radiometers:

- Upgrades underway, resulting in some down time

- Operating normally