## NEPH, CLAP, PSAP
Date | Julian Day | Time (UTC) | Observer's Initials | Check Daily Aerosol Email | NEPH Lamp On/Current(<7A) | NEPH Flag Status | PSAP flow ~0.52 lpm | PSAP Tr>0.75 | CLAP flow ~0.65 lpm | CLAP Tr>0.75 | CLAP Spot # | Bi-Weekly Span Check (Mon) | Comments
2014-10-01 | 274 |  | DP | x | 5.7 | x | 0.51 | 0.9 | 0.58 | 0.99 | 2
2014-10-02 | 275 |  | SMC | x | 5.7 | x | 0.52 | 0.89 | 0.58 | 0.98 | 2
2014-10-03 | 276 |  | DP | x | 5.7 | x | 0.52 | 0.88 | 0.59 | 0.96 | 2
2014-10-04 | 277 |  | SMC | X | 5.6 | x | 0.51 | 0.66 | 0.57 | 0.88 | 3
2014-10-05 | 278 |  | SMC | x | 5.6 | x | 0.53 | 0.64 | 0.6 | 0.87 | 3 |  | *changed PSAP filter
2014-10-06 | 279 |  | DP | x | 5.6 | x | 0.54 | 0.97 | 0.6 | 0.86 | 3
2014-10-07 | 280 |  | DP | x | 5.7 | x | 0.54 | 0.99 | 0.6 | 0.99 | 4
2014-10-08 | 281 |  | SMC | x | 5.6 | x | 0.54 | 0.98 | 0.6 | 0.99 | 4
2014-10-09 | 282 |  | SMC | x | 5.6 | x | 0.54 | 0.97 | 0.6 | 0.97 | 4
2014-10-10 | 283 |  | DP | x | 5.6 | x | 0.54 | 0.97 | 0.61 | 0.96 | 4
2014-10-11 | 284 |  | HJ | x | 5.7 | x | 0.54 | 0.96 | 0.61 | 0.94 | 4
2014-10-12 | 285 |  | HJ | x | 5.7 | x | 0.53 | 0.952 | 0.6 | 0.93 | 4
2014-10-13 | 286 |  | HJ |  | 5.7 | x | 0.53 | 0.952 | 0.6 | 0.93 | 4 | 2.02%
2014-10-14 | 287 |  | DP | x | 5.6 | x | 0.53 | 0.946 | 0.59 | 0.93 | 4
2014-10-15 | 288 |  | HJ |  | 5.7 | x | 0.53 | 0.94 | 0.59 | 0.91 | 4 |  | Planned power outage from 12:10-12:30UTC for turnover training
2014-10-16 | 289 |  | LR | X | 5.6 | x | 0.53 | 0.93 | 0.59 | 0.88 | 4
2014-10-17 | 290 |  | HJ |  | 5.6 | X | 0.53 | 0.943 | 0.59 | 0.89 | 4
2014-10-18 | 291 | 1330 | LR | x | 5.6 | x | 0.52 | 0.935 | 0.59 | 0.85 | 4
2014-10-19 | 292 | 1348 | HJ | x | 5.6 | 100 | 0.52 | 0.928 | 0.58 | 0.83 | 4
2014-10-20 | 293 | 1158 | HJ |  | 5.6 | 2100 | 0.52 | 0.787 | 0.58 | 0.92 | 5
2014-10-21 | 294 | 1146 | HJ | x | 5.6 | 100 | 0.52 | 0.751 | 0.58 | 0.86 | 5
2014-10-22 | 295 | 1156 | HJ | x | 5.6 | 2100 | 0.51 | 0.742 | 0.65 | 0.85 | 5 |  | LR adjusted CLAP flow from .58 to .65
2014-10-23 | 296 | 1142 | HJ | x | 5.6 | 100 | 0.51 | 0.735 | 0.64 | 0.84 | 5
2014-10-24 | 297 | 1238 | HJ |  | 5.6 | 100 | 0.51 | 0.724 | 0.64 | 0.81 | 5
2014-10-25 | 298 | 1251 | LR | * | 5.6 | * | 0.51 | 0.717 | 0.65 | 0.78 | 5
2014-10-26 | 299 | 1315 | LR | * | 5.6 | * | 0.52 | 0.711 | 0.66 | 0.76 | 5
2014-10-27 | 300 | 1326 | HJ |  | 5.6 | 100 | 0.42 | 0.697 | 0.68 | 0.72 | 5 | 2.58%
2014-10-28 | 301 | 1215 | LR | * | 5.6 |  | 0.52 | 0.999 | 0.66 | 0.998 | 6
2014-10-29 | 302 | 1300 | LR | * | 5.6 | 2100 | 0.52 | 1.001 | 0.66 | * | 6
2014-10-30 | 303 | 1418 | HJ | * | 5.6 | 100 | 0.52 | 1.002 | 0.66 | 0.997 | 6
2014-10-31 | 304 | 1245 | LR | * | 5.6 |  | 0.52 | 0.999 | 0.65 | * | 6
2014-11-01 | 305 | 1340 | HJ | * | 5.6 | 100 | 0.52 | 1 | 0.64 | 0.99 | 6
2014-11-02 | 306 | 1410 | LR | * | 5.6 |  | 0.52 | 0.997 | 0.65 | * | 6
2014-11-03 | 307 | 1353 | HJ | * | 5.6 | 100 | 0.52 | 0.997 | 0.65 | 0.975 | 6
2014-11-04 | 308 | 1315 | LR | * | 5.6 |  | 0.52 | 0.975 | 0.64 | *
2014-11-05 | 309 | 1246 | HJ | * | 5.6 | 100 | 0.52 | 0.958 | 0.65 | 0.899 | 6 |  | HJ adjusted clap flow from .62 to .65
2014-11-06 | 310 | 1335 | HJ | * | 5.6 | 100 | 0.51 | 0.954 | 0.65 | 0.892 | 6
2014-11-07 | 311 | 1252 | LR | * | 5.6 |  | 0.51 | 0.949 | 0.65 | * | 6
2014-11-08 | 312 | 1307 | HJ | * | 5.6 | 100 | 0.51 | 1.071 | 0.65 | 0.875 | 6
2014-11-09 | 313 | 1437 | HJ | * | 5.6 | 100 | 0.04 | 1.209 | 0 | 0.875 | 6 |  | pump was off, all flows at 0, email reads ''flow error for 91% of data''. Fixed and working at 18:15 UTC
2014-11-10 | 314 | 1300 | LR | * | 5.6 | 2100 | 0.04/.51 | 1.243! | 0/.65 | * | 6 |  | pump was off. Brief pwr outage at 0600UTC. Pump running at 1426UTC
2014-11-11 | 315 | 1230 | LR | * | 5.6 |  | 0.51 | 0.995 | 0.65 | * | 6 |  | Planned power outages from 1259-1301UTC & 1303-1305UTC
2014-11-12 | 316 | 1255 | LR | * | 5.6 |  | 0.52 | 0.988 | 0.65 | * | 6
2014-11-13 | 317 | 1443 | HJ | * | 5.5 | 100 | 0.52 | 0.984 | 0.65 | 0.831 | 6
2014-11-14 | 318 | 1211 | LR | * | 5.6 |  | 0.52 | 0.976 | 0.65 | * | 6 |  | Planned power outage from 1257-1301UTC
2014-11-15 | 319 | 1319 | HJ | * | 5.6 | 100 | 0.52 | 0.97 | 0.65 | 0.806 | 6
2014-11-16 | 320 | 1418 | HJ | * | 5.5 | 0 | 0.52 | 0.962 | 0.65 | 0.791 | 6
2014-11-17 | 321 | 1315 | LR | * | 5.6 |  | 5.1 | 0.957 | 0.65 | * | 6
2014-11-18 | 322 | 1230 | LR | * | 5.6 |  | 5.1 | 0.951 | 0.65 | * | 6
2014-11-19 | 323 | - | HJ | [9 empty] | 30+ knot winds, did not travel to TAWO
2014-11-20 | 324 | 1247 | HJ | * | * | 0 | 0.51 | 0.941 | 0.65 | 0.731 | 6
2014-11-21 | 325 | - | HJ | [9 empty] | 40+ knot winds, did not travel to TAWO
2014-11-22 | 326 | 1324 | HJ | * | 5.5 | 100 | 0.51 | 0.963 | 0.65 | 0.718 | 6
2014-11-23 | 327 | 1539 | HJ | * | 5.5 | 0 | 0.51 | 0.931 | 0.67 | 0.703 | 6
2014-11-24 | 328 | 1330 | LR | * | 5.5 |  | 0.51 |  | 0.65 | * | 7
2014-11-25 | 329 | 1345 | LR | * | 5.5 |  | 0.51 |  | 0.65 | * | 7
2014-11-26 | 330 | 1449 | HJ | * | 5.5 | 100 | 0.5 | 0.92 | 0.64 | 0.983 | 7
2014-11-27 | 331 | 1356 | HJ | * | 5.5 | 100 | 0.51 | 0.917 | 0.64 | 0.977 | 7
2014-11-28 | 332 | 1327 | HJ | * | 5.5 | 100 | 0.51 | 0.912 | 0.65 | 0.969 | 7
2014-11-29 | 333 | [11 empty] | High winds, did not travel to TAWO
2014-11-30 | 334 | 1609 | HJ | * | 5.5 | 100 | 0.5 | 0.907 | 0.64 | 0.951 | 7
2014-12-01 | 335 | 1337 | HJ | * | 5.5 | 100 | 0.5 | 0.885 | 0.63 | 0.914 | 7
2014-12-02 | 336 | 1400 | HJ | * | 5.5 | 2100 | 0.49 | 0.868 | 0.65 | 0.889 | 7
2014-12-03 | 337 | 1537 | HJ | * | 5.5 | 100 | 0.5 | 0.864 | 0.66 | 0.874 | 7
2014-12-04 | 338 | 1351 | HJ | * | 5.5 | 100 | 0.51 | 0.86 | 0.66 | 0.863 | 7
2014-12-05 | 339 | 1440 | LR | * | 5.6 |  | 0.51 | 0.855 | 0.65 | 0.851 | 7
2014-12-06 | 340 | 1750 | LR | * | 5.5 |  | 0.5 | 0.851 | 0.65 | 0.845 | 7
2014-12-07 | 341 | 1637 | HJ | * | 5.5 | 100 | 0.51 | 0.839 | 0.65 | 0.819 | 7
2014-12-08 | 342 | 1312 | HJ | * | 5.5 | 100 | 0.5 | 0.834 | 0.64 | 0.809 | 7
2014-12-09 | 343 | 1220 | LR | * | 5.5 |  | 0.5 | 0.832 | 0.64 | 0.8 | 7
2014-12-10 | 344 | 1417 | HJ | * | 5.5 | 100 | 0.5 | 0.756 | 0.65 | 0.981 | 8
2014-12-11 | 345 | 1352 | HJ | * | 5.5 | 100 | 0.51 | 0.754 | 0.66 | 0.975 | 8
2014-12-12 | 346 | 1717 | LR | * | 5.5 |  | 0.5 | 0.744 | 0.63 | 0.95 | 8
2014-12-13 | 347 | 1309 | HJ | * | 5.5 | 100 | 0.49 | 0.741 | .62-adjusted | 0.935 | 8
2014-12-14 | 348 | na | na
2014-12-15 | 349 | 1429 | HJ | * | 5.5 | 100 | 0.5 | 0.683 | 0.65 | 0.785 | 8 |  | Changed PSAP filter
2014-12-16 | 350 | 1600 | LR | * | 5.5 |  | 0.49 | * | 0.65 | * | 8
2014-12-17 | 351 | 1357 | HJ | * | 5.5 | 6100 | 0.5 | 0.989 | 0.66 | 0.764 | 8
2014-12-18 | 352
2014-12-19 | 353 | 1305 | LR | * | 5.5 |  | 0.51 | * | .68_adj. | 0.753 | 8
2014-12-20 | 354 | 1535 | LR | * | 5.5 |  | 0.5 | * | 0.64 | .750 changed | 1 |  | Changed CLAP filter 1600UTC
2014-12-21 | 355
2014-12-22 | 356 | 1400 | LR | * | 5.5 |  | 0.5 | * | 0.63 | * | 1
2014-12-23 | 357 | [11 empty] | Did not travel to TAWO
2014-12-24 | 358 | 1411 | HJ | * | 5.5 | 100 | 0.51 | 0.973 | 0.65 | 0.984 | 1
2014-12-25 | 359 | 1545 | LR | * | 5.4 |  | 0.51 | * | 0.64 | * | 1
2014-12-26 | 360 | 1445 | HJ | * | 5.5 | 100 | 0.51 | 0.967 | 0.65 | 0.977 | 1
2014-12-27 | 361 | [11 empty] | Did not travel to TAWO
2014-12-28 | 362 | 1650 | HJ | * | 5.5 | 0 | 0.51 | 0.965 | 0.63 | 0.972 | 1
2014-12-29 | 363 | 1425 | LR | * | 5.4 |  | 0.5 | * | 0.64 | * | 1
2014-12-30 | 364 | 1409 | HJ | * | 5.4 | 100 | 0.5 | 0.966 | 0.64 | 0.971 | 1
2014-12-31 | 365 | 1721 | HJ | * | 5.4 | 100 | 0.49 | 0.965 | 0.64 | 0.967 | 1

## CU GC
Date | Julian Day | Time (UTC) | Observer's Initials | H2 Gen. Act Pressure (PSI) | H2 Generator Water Level | H2 Gen. Flow (ml/min) | Methane Std 1 Pressure | Methane Std 2 Pressure | Inlet Line Flow Rate (L/min) | 4 Pumps On? | NMHC Std 1 Pressure | NMHC Std 2 Pressure | Comments (Inlet Cleaning, Inlet Filter Changes, Water Refill, etc.
2014-10-01 | 274 |  | DP | 60 | x | 45 | 1700 | 1550 | 2 | y | 825 | 750
2014-10-02 | 275 |  | SMC | 60 | x | 46 | 1700 | 1550 | 2 | y | 825 | 750
2014-10-03 | 276 |  | DP | 60 | x | 45 | 1700 | 1550 | 2 | y | 825 | 750
2014-10-04 | 277 |  | SMC | 60 | x | 45 | 1700 | 1550 | 2 | y | 825 | 750
2014-10-05 | 278 |  | SMC | 60 | x | 46 | 1700 | 1550 | 2 | y | 825 | 750
2014-10-06 | 279 |  | DP | 60 | x | 46 | 1700 | 1550 | 2 | y | 825 | 750
2014-10-07 | 280 |  | DP | 60 | x | 46 | 1700 | 1550 | 2 | y | 825 | 750
2014-10-08 | 281 |  | SMC | 60 | x | 47 | 1700 | 1550 | 2 | y | 825 | 750
2014-10-09 | 282 |  | SMC | 60 | x | 46 | 1700 | 1550 | 2 | y | 825 | 750
2014-10-10 | 283 |  | DP | 60 | x | 46 | 1700 | 1550 | 2 | y | 825 | 750
2014-10-11 | 284 |  | HJ | 60 | x | 46 | 1700 | 1550 | 2 | y | 825 | 725 | added water at daily rounds in a.m. alarm went off around 1400, more water was added.Error "H2 Outlet ACT 1.0 PSI *LOCK*" continued at 2030, LR performed a hard reset. System appeards normal -HJ & LR
2014-10-12 | 285 |  | HJ | 60 | x | 47 | 1700 | 1550 | 2 | y | 825 | 725
2014-10-13 | 286 |  | HJ | 60 | x | 46 | 1700 | 1550 | 2 | y | 825 | 725
2014-10-14 | 287 |  | DP | 60 | x | 46 | 1700 | 1550 | 2 | y | 825 | 725
2014-10-15 | 288 |  | HJ | 60 | x | 46 | 1700 | 1550 | 2 | y | 800 | 725 | Planned power outage from 12:10-12:30UTC for turnover training
2014-10-16 | 289 |  | LR | 60 | x | 45 | 1700 | 1550 | 2 | y | 800 | 725
2014-10-17 | 290 |  | HJ | 60 | X | 45 | 1690 | 1550 | 2 | Y | 800 | 725
2014-10-18 | 291 | 1330 | LR | 60 | x | 45 | 1690 | 1550 | 2 | y | 800 | 725
2014-10-19 | 292 | 1350 | HJ | 60 | x | 46 | 1690 | 1550 | 2 | Y | 800 | 725
2014-10-20 | 293 | 1203 | HJ | 60 | x | 46 | 1690 | 1550 | 2 | Y | 800 | 725
2014-10-21 | 294 | 1151 | HJ | 60 | x | 45 | 1690 | 1550 | 2 | Y | 800 | 725
2014-10-22 | 295 | 1156 | HJ | 60 | x | 45 | 1690 | 1550 | 2 | y | 800 | 725
2014-10-23 | 296 | 1142 | HJ | 60 | X | locked | 1690 | 1550 | 2 | y | 800 | 725 | H2 geneator reads "float chamber act1.4psi *lock* set 60psi flow". LR reset at 1140UTC, seems to be running fine since
2014-10-24 | 297 | 1238 | HJ |  |  |  | 1690 | 1550 | 2 | y | 800 | 725 | H2 cyliner at 1000psi/45psi
2014-10-25 | 298 | 1251 | LR |  |  |  | 1690 | 1550 | 2 | y | 800 | 725 | H2 cyliner at 1000psi/45psi
2014-10-26 | 299 | 1315 | LR |  |  |  | 1690 | 1550 | 2 | y | 800 | 725 | H2 cyliner at 1000psi/45psi
2014-10-27 | 300 | 1326 | HJ |  |  |  | 1690 | 1550 | 2 | y | 800 | 725 | H2 cyliner at 1000psi/45psi
2014-10-28 | 301 | 1215 | LR |  |  |  | 1690 | 1550 | 2 | y | 800 | 725 | H2 cyliner at 1000psi/45psi
2014-10-29 | 302 | 1300 | LR |  |  |  | 1690 | 1550 | 2 | y | 800 | 725 | H2 cyliner at 950psi/45psi
2014-10-30 | 303 | 1418 | HJ |  |  |  | 1690 | 1550 | 2 | Y | 800 | 725 | H2 cyliner at 900psi/45psi
2014-10-31 | 304 | 1245 | LR |  |  |  | 1690 | 1550 | 2 | Y | 800 | 725 | H2 cyliner at 900psi/45psi
2014-11-01 | 305 | 1340 | HJ |  |  |  | 1690 | 1550 | 2 | Y | 800 | 725 | H2 cyliner at 890psi/45psi
2014-11-02 | 306 | 1410 | LR |  |  |  | 1690 | 1550 | 2 | Y | 800 | 725 | H2 cyliner at 850psi/45psi
2014-11-03 | 307 | 1353 | HJ |  |  |  | 1690 | 1550 | 2 | y | 800 | 725 | H2 cylinder at 800/45psi
2014-11-04 | 308 | 1315 | LR |  |  |  | 1690 | 1550 | 2 | y | 800 | 725 | H2 cylinder at 800/45psi
2014-11-05 | 309 | 1246 | HJ |  |  |  | 1690 | 1550 | 2 | y | 800 | 725 | H2 cylinder at 800/45psi
2014-11-06 | 310 | 1335 | HJ |  |  |  | 1690 | 1550 | 2 | Y | 800 | 725 | H2 cylinder at 775/45psi
2014-11-07 | 311 | 1252 | LR |  |  |  | 1690 | 1550 | 2 | Y | 800 | 725 | H2 cylinder at 775/45psi
2014-11-08 | 312 | 1307 | HJ |  |  |  | 1690 | 1550 | 2 | Y | 800 | 725 | H2 cylinder at 700/45psi
2014-11-09 | 313 | 1437 | HJ |  |  |  | 1690 | 1550 | 2 | y | 800 | 725 | H2 cylinder at 650/45psi
2014-11-10 | 314 | 1300 | LR |  |  |  | 1690 | 1550 | 2 | y | 800 | 725 | H2 cylinder at 650/45psi. power outage from 0700-0710.
2014-11-11 | 315 | 1230 | LR |  |  |  | 1690 | 1550 | 2 | y | 800 | 725 | H2 cylinder at 650/45psi, Planned power outages from 1259-1301UTC & 1303-1305UTC
2014-11-12 | 316 | 1255 | LR |  |  |  | 1690 | 1550 | 2 | y | 800 | 725 | H2 cylinder at 650/45psi
2014-11-13 | 317 | 1443 | HJ |  |  |  | 1690 | 1550 | 2 | y | 800 | 725 | H2 cylinder at 600/45psi
2014-11-14 | 318 | 1211 | LR |  |  |  | 1690 | 1550 | 2 | y | 800 | 725 | Planned power outage from 1257-1301UCT. H2 cylinder at 600/45psi.
2014-11-15 | 319 | 1319 | HJ |  |  |  | 1690 | 1550 | 2 | Y | 800 | 725 | H2 cylinder at 600/45psi
2014-11-16 | 320 | 1418 | HJ |  |  |  | 1690 | 1550 | 2 | Y | 800 | 725 | H2 cylinder at 575/45psi
2014-11-17 | 321 | 1315 | LR |  |  |  | 1690 | 1550 | 2 | Y | 800 | 725 | H2 cylinder at 575/45psi
2014-11-18 | 322 | 1230 | LR |  |  |  | 1690 | 1550 | 2 | Y | 800 | 725 | H2 cylinder at 550/45psi
2014-11-19 | 323 |  | HJ | [9 empty] | 30+ knot winds, did not travel to TAWO
2014-11-20 | 324 | 1247 | HJ |  |  |  | 1690 | 1550 | 2 | Y | 800 | 725 | H2 cylinder at 500/45psi
2014-11-21 | 325 | - | HJ | [9 empty] | 40+ knot winds, did not travel to TAWO
2014-11-22 | 326 | 1324 | HJ |  |  |  | 1690 | 1550 | 2 | y | 800 | 700 | H2 cylinder at 450/45psi
2014-11-23 | 327 | 1539 | HJ |  |  |  | 1690 | 1550 | 2 | Y | 800 | 725 | H2 cylinder at 425/45psi
2014-11-24 | 328 | 1330 | LR |  |  |  | 1690 | 1550 | 2 | Y | 800 | 725 | H2 cylinder at 425/45psi; One of the fan bearings is making excessive noise.
2014-11-25 | 329 | 1345 | LR |  |  |  | 1690 | 1550 | 2 | N | 800 | 700 | H2 cylinder at 425/45psi; One of the fan bearings is making excessive noise. Fixed at 1930utc
2014-11-26 | 330 | 1449 | HJ |  |  |  | 1690 | 1550 | 2 | Y | 800 | 700 | H2 cylinder at 350/45psi
2014-11-27 | 331 | 1356 | HJ |  |  |  | 1690 | 1550 | 2 | y | 800 | 725 | H2 cylinder at 300/45psi
2014-11-28 | 332 | 1327 | HJ |  |  |  | 1690 | 1550 | 2 | y | 800 | 725 | H2 cylinder at 300/45psi
2014-11-29 | 333 | 0 | 0 | [9 empty] | High winds, did not travel to TAWO
2014-11-30 | 334 | 1609 | HJ |  |  |  | 1690 | 1550 | 2 | Y | 800 | 725 | H2 cylinder at 300/45psi
2014-12-01 | 335 | 1337 | HJ |  |  |  | 1690 | 1550 | 2 | Y | 800 | 700 | H2 cylinder at 250/45psi
2014-12-02 | 336 | 1400 | HJ |  |  |  | 1690 | 1550 | 2 | y | 800 | 725 | H2 cylinder at 200/45psi
2014-12-03 | 337 | 1537 | HJ |  |  |  | 1690 | 1550 | 2 | Y | 800 | 700 | H2 cylinder at 200/45psi
2014-12-04 | 338 | 1351 | HJ |  |  |  | 1690 | 1550 | 2 | Y | 800 | 700 | H2 cylinder at 200/45psi
2014-12-05 | 339 | 1440 | LR |  |  |  | 1690 | 1550 | 2 | Y | 800 | 700 | H2 cylinder at 200/45psi, changed H2 cylinder 2100/45psi
2014-12-06 | 340 | 1750 | LR |  |  |  | 1690 | 1550 | 2 | Y | 800 | 700 | H2 cylinder at 2050/45psi
2014-12-07 | 341 | 1637 | HJ |  |  |  | 1690 | 1550 | 2 | Y | 800 | 700 | H2 cylinder at 2000/50psi
2014-12-08 | 342 | 1312 | HJ |  |  |  | 1690 | 1550 | 2 | Y | 800 | 700 | H2 cylinder at 2000/47psi
2014-12-09 | 343 | 1220 | LR |  |  |  | 1690 | 1550 | 2 | Y | 800 | 700 | H2 cylinder at 2000/47psi
2014-12-10 | 344 | 1417 | HJ |  |  |  | 1690 | 1550 | 2 | Y | 800 | 700 | H2 cylinder at 1950/50psi
2014-12-11 | 345 | 1352 | HJ |  |  |  | 1690 | 1550 | 2 | y | 800 | 700 | H2 cylinder at 1900/50psi
2014-12-12 | 346 | 1717 | LR |  |  |  | 1690 | 1550 | 2 | y | 800 | 700 | H2 cylinder at 1900/50psi
2014-12-13 | 347 | 1309 | HJ |  |  |  | 1690 | 1550 | 2 | Y | 800 | 700 | H2 cylinder at 1950/40psi
2014-12-14 | 348
2014-12-15 | 349 | 1429 | HJ |  |  |  | 1690 | 1550 | 2 | Y | 800 | 700 | H2 cylinder at 1950/40psi
2014-12-16 | 350 | 1600 | LR |  |  |  | 1690 | 1550 | 2 | Y | 800 | 700 | H2 cylinder at 1850/45psi
2014-12-17 | 351 | 1357 | HJ |  |  |  | 1690 | 1550 | 2 | Y | 800 | 700 | H2 cylinder at 1825/45psi
2014-12-18 | 352
2014-12-19 | 353 | 1305 | LR |  |  |  | 1690 | 1550 | 2 | Y | 800 | 700 | H2 cylinder at 1875/40psi
2014-12-20 | 354 | 1535 | LR |  |  |  | 1690 | 1550 | 2 | Y | 800 | 700 | H2 cylinder at 1850/40psi
2014-12-21 | 355
2014-12-22 | 356 | 1400 | LR |  |  |  | 1690 | 1550 | 2 | Y | 800 | 700 | H2 cylinder at 1825/42psi
2014-12-23 | 357
2014-12-24 | 358 | 1411 | HJ |  |  |  | 1650 | 1550 | 2 | Y | 800 | 700 | H2 cylinder at 1650/50psi
2014-12-25 | 359 | 1545 | LR |  |  |  | 1650 | 1550 | 2 | Y | 800 | 700 | H2 cylinder at 1750/42psi
2014-12-26 | 360 | 1445 | HJ |  |  |  | 1690 | 1550 | 2 | Y | 800 | 700 | H2 cylinder at 1700/42psi
2014-12-27 | 361 | 0 | 0 | [9 empty] | did not travel to tawo
2014-12-28 | 362 | 1650 | HJ |  |  |  | 1650 | 1550 | 2 | y | 800 | 700 | H2 cylinder at 1600/45psi
2014-12-29 | 363 | 1425 | LR |  |  |  | 1650 | 1550 | 2 | y | 800 | 700 | H2 cylinder at 1600/45psi
2014-12-30 | 364 | 1409 | HJ |  |  |  | 1690 | 1550 | 2 | Y | 800 | 700 | H2 cylinder at 1590/48psi | 1550 | 2 | y | 800 | 700 | H2 cylinder at 1600/45psi
2014-12-31 | 365 | 1721 | HJ |  |  |  | 1690 | 1550 | 2 | Y | 800 | 700 | H2 cylinder at 1600/43psi

## MET Tower
Date | Julian Day | Time (UTC) | Observer's Initials | 2m temp | 2m RH | 2m V_TD | CATS CG | Surface O11 | Black Carbon | 10m Windbird | 10m temp | 16m Windbird | 16m temp | Offset Zeno Time (if < 5sec, reset) | Comments (Climbing, removing instruments, etc)
2014-10-01 | 274 | 1247 | DP | x | x | x | x | x | x | x | x | x | x | -4
2014-10-02 | 275 | 1601 | SMC | x | x | x | x | x | x | x | x | x | x | reset
2014-10-03 | 276 | 1133 | DP | x | x | x | x | x | x | x | x | x | x | -2
2014-10-04 | 277 | 1221 | SMC | x | x | x | x | x | x | x | x | x | x | -4
2014-10-05 | 278 | 1745 | SMC | x | x | x | x | x | x | x | x | x | x | reset
2014-10-06 | 279 | 1152 | DP | x | x | x | x | x | x | x | x | x | x | -1
2014-10-07 | 280 | 1232 | DP | x | x | x | x | x | x | x | x | x | x | -3
2014-10-08 | 281 | 15:40 | SMC | x | x | x | x | x | x | x | x | x | x | -4
2014-10-09 | 282 | 1821 | SMC | x | x | x | x | x | x | x | x | x | x | reset
2014-10-10 | 283 | 1524 | DP | x | x | x | x | x | x | x | x | x | x | -1
2014-10-11 | 284 | 1130 | HJ | x | x | x | x | x | x | x | x | x | x | -3
2014-10-12 | 285 | 1400 | HJ | rm frost | rm frost | rm frost | hvy frost | hvy frost | hvy frost | hvy frost | hvy frost | hvy frost | hvy frost | -5 | noted 2m humitiy sensor fan not working
2014-10-13 | 286 | 1132 | HJ | rm frost | rm frost | rm frost | hvy frost | hvy frost | hvy frost | hvy frost | hvy frost | hvy frost | hvy frost | -6
2014-10-14 | 287 | 1245 | DP | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | reset | HJ Climbed tower from 1245-1330. All instruments were cleared of snow.
2014-10-15 | 288 | 1140 | HJ | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 0 | Planned power outage from 12:10-12:30UTC for turnover training
2014-10-16 | 289 | 1330 | LR | frost | frost | frost | frost | frost | frost | frost | frost | frost | frost | -4
2014-10-17 | 290 | 1130 | HJ | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | -6 | HJ Climbed tower from 1130-1200UTC. All instruments were cleared of snow.
2014-10-18 | 291 | 1345 | LR | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | reset
2014-10-19 | 292 | 1350 | HJ | frost | frost | frost | frost | frost | frost | frost | frost | frost | frost | -2
2014-10-20 | 293 | 1202 | HJ | rm frost | rm frost | rm frost | frost | frost | frost | frost | frost | frost | frost | -3 | LR cleared 2m instruments
2014-10-21 | 294 | 1149 | HJ | rm frost | rm frost | rm frost | frost | frost | frost | frost | frost | frost | frost | -4 | HJ cleared 2m instruments
2014-10-22 | 295 | 1156 | HJ | frost | frost | frost | frost | frost | frost | frost | frost | frost | frost | reset
2014-10-23 | 296 | 1142 | HJ | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | -1
2014-10-24 | 297 | 1238 | HJ | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | -3
2014-10-25 | 298 | 1251 | LR | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | -4 | climbed tower 1250 - 1310 UTC
2014-10-26 | 299 | 1315 | LR | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | -6
2014-10-27 | 300 | 1326 | HJ | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | reset
2014-10-28 | 301 | 1215 | LR | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | -2 | climbed tower 1325-1345UTC LR
2014-10-29 | 302 | 1300 | LR | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | -4 | climbed tower 1300-1330UTC HJ
2014-10-30 | 303 | 1418 | HJ | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | -5 | climbed tower 1300-1312UTC HJ
2014-10-31 | 304 | 1245 | LR | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | reset
2014-11-01 | 305 | 1340 | HJ | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | 0 | climbed tower 1340-1350 LR
2014-11-02 | 306 | 1410 | LR | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | -1 | climbed tower 1415-1445 HJ
2014-11-03 | 307 | 1353 | HJ | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | -4
2014-11-04 | 308 | 1315 | LR | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | -5
2014-11-05 | 309 | 1246 | HJ | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | reset
2014-11-06 | 310 | 1335 | HJ | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | -2
2014-11-07 | 311 | 1252 | LR | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | -3
2014-11-08 | 312 | 1307 | HJ | frost | frost | frost | frost | frost | frost | frost | frost | frost | frost | -5
2014-11-09 | 313 | 1437 | HJ | frost | frost | frost | frost | frost | frost | frost | frost | frost | frost | reset
2014-11-10 | 314 | 1300 | LR | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | -1 | power outage from 0700-0710. climbed tower 1400-14150HJ
2014-11-11 | 315 | 1230 | LR | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | -3 | Planned power outages from 1259-1301UTC & 1303-1305UTC
2014-11-12 | 316 | 1255 | LR | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | -5 | Removed and fixed humidity fan 1322-1357utc. LR
2014-11-13 | 317 | 1443 | HJ | frost | frost | frost | frost | frost | frost | frost | frost | frost | frost | reset
2014-11-14 | 318 | 1211 | LR | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 0 | Planned power outage from 12:57-13:01UTC
2014-11-15 | 319 | 1319 | HJ | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | -1 | restarted bunting computer at 1339utc
2014-11-16 | 320 | 1418 | HJ | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | -3
2014-11-17 | 321 | 1315 | LR | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | -4
2014-11-18 | 322 | 1230 | LR | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | reset
2014-11-19 | 323 |  | HJ | [11 empty] | 30+ knot winds, did not travel to TAWO
2014-11-20 | 324 | 1247 | HJ | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | -3
2014-11-21 | 325 | - | HJ | [11 empty] | 40+ knot winds, did not travel to TAWO
2014-11-22 | 326 | 1324 | HJ | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | reset
2014-11-23 | 327 | 1539 | HJ | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | -1
2014-11-24 | 328 | 1330 | LR | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | -2
2014-11-25 | 329 | 1345 | LR | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | -3
2014-11-26 | 330 | 1449 | HJ | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | reset
2014-11-27 | 331 | 1356 | HJ | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | -2
2014-11-28 | 332 | 1327 | HJ | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | -3
2014-11-29 | 333 | 0 | 0 | [11 empty] | High winds, did not travel to TAWO
2014-11-30 | 334 | 1609 | HJ | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | reset
2014-12-01 | 335 | 1337 | HJ | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | -2
2014-12-02 | 336 | 1400 | HJ | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | -3
2014-12-03 | 337 | 1537 | HJ | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | -5
2014-12-04 | 338 | 1351 | HJ | rm frost | rm frost | rm frost | frost | frost | frost | frost | frost | frost | frost | reset | cleared 2m instruments at 1340utc
2014-12-05 | 339 | 1440 | LR | rm frost | rm frost | rm frost | frost | frost | frost | frost | frost | frost | frost | -1 | cleared 2m instruments at 1340utc. Humidity sensor has been @61% for several days and does not compare well with the temp/dew point sensor.
2014-12-06 | 340 | 1750 | LR | rm frost | rm frost | rm frost | frost | frost | frost | frost | frost | frost | frost | -2 | removed Humidity cable 1745-1750 UTC. (61 91 100 100 100 60 61).
2014-12-07 | 341 | 1637 | HJ | lt frost | lt frost | lt frost | frost | frost | frost | frost | frost | frost | frost | -5
2014-12-08 | 342 | 1312 | HJ | frost | frost | frost | frost | frost | frost | frost | frost | frost | frost | reset
2014-12-09 | 343 | 1220 | LR | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | 0 | climbed tower 1525-1533utc
2014-12-10 | 344 | 1417 | HJ | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | -2
2014-12-11 | 345 | 1352 | HJ | frost | frost | frost | frost | frost | frost | frost | frost | frost | frost | -4
2014-12-12 | 346 | 1717 | LR | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | -5
2014-12-13 | 347 | 1309 | HJ | clr | clr | clr | cc | clr | clr | clr | clr | clr | clr | reset
2014-12-14 | 348 | na | na
2014-12-15 | 349 | 1429 | HJ | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | -3
2014-12-16 | 350 | 1600 | LR | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | -4
2014-12-17 | 351 | 1357 | HJ | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | reset
2014-12-18 | 352
2014-12-19 | 353 | 1305 | LR | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | reset | date was off a day. LR
2014-12-20 | 354 | 1535 | LR | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 0
2014-12-21 | 355
2014-12-22 | 356 | 1400 | LR | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | -3
2014-12-23 | 357 | [13 empty] | Did not travel to TAWO
2014-12-24 | 358 | 1411 | HJ | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | reset | LR climbed tower 1408-1420utc
2014-12-25 | 359 | 1545 | LR | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | -3
2014-12-26 | 360 | 1445 | HJ | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | -4 | 10m windbird stuck at 309° until 1440utc
2014-12-27 | 361 | 0 | 0 | [11 empty] | Did not travel to TAWO
2014-12-28 | 362 | 1650 | HJ | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | reset
2014-12-29 | 363 | 1425 | LR | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | -1
2014-12-30 | 364 | 1409 | HJ | frost | frost | frost | frost | frost | frost | frost | frost | frost | frost | -3
2014-12-31 | 365 | 1721 | HJ | frost | frost | frost | frost | frost | frost | frost | frost | frost | frost | reset

## Aethalometer AE16-AE33
Date | Julian Day | Time (UTC) | Observer's Initials | Sample inlet clear? | Check date/time (GMT) | Panel Readout BC ng/m3 | Flow rate, lpm AE16/AE33 | Check Disk | Weekly Data Transfer | Weekly Check tape | Comments
2014-10-01 | 274
2014-10-02 | 275
2014-10-03 | 276
2014-10-04 | 277
2014-10-05 | 278
2014-10-06 | 279
2014-10-07 | 280
2014-10-08 | 281
2014-10-09 | 282
2014-10-10 | 283
2014-10-11 | 284
2014-10-12 | 285
2014-10-13 | 286
2014-10-14 | 287
2014-10-15 | 288
2014-10-16 | 289 | 1430 | LR |  | ? | 127 | 7.5 |  |  |  | Restarted w/ new AE33 aethalometer.
2014-10-17 | 290 | 1235 | HJ |  |  | 0 | 8 |  |  |  | Restarted w/ new AE33 aethalometer inside air
2014-10-18 | 291 | 1347 | LR |  | x | 205/207 | 7.6/6.0 |  |  | x/x | sampling room air. Adding T to inlet at 1430 UTC.
2014-10-19 | 292 | 1352 | HJ |  |  | 106 | 6 |  |  |  | AE16 advanced extra tape. Attempted restart at 1352UTC. Screen went blank and tape continued to advance. Currently shutdown. Added serial cables.
2014-10-20 | 293 | 1205 | HJ |  |  | 611/267 | 7.5/6.0 |  |  |  | At 1150UTC restarted AE16
2014-10-21 | 294 | 1145 | HJ |  |  |  | 6 |  |  |  | AE16 CURRENTLY SHUTDOWN
2014-10-22 | 295 | 1156 | HJ |  |  |  | 5.9
2014-10-23 | 296 | 1142 | HJ |  |  | X | 5.9
2014-10-24 | 297 | 1238 | HJ |  |  |  | 5.9 |  |  |  | AE33 sampling ambient air at 13:30UTC. 4.3lmp w/ internal pump, 5.8 w/ external pump
2014-10-25 | 298 | 1251 | LR | * | * | -397 | 5.8
2014-10-26 | 299 | 1315 | LR | * | * | 28 | 5.9
2014-10-27 | 300 | 1326 | HJ | * | * | -107 | 5.9
2014-10-28 | 301 | 1215 | LR | * | * | 173 | 5.9
2014-10-29 | 302 | 1300 | LR | * | * | 220 | 5.9
2014-10-30 | 303 | 1418 | HJ | * | * | 312 | 5.9
2014-10-31 | 304 | 1245 | LR | * | * | 360 | 6
2014-11-01 | 305 | 1340 | HJ | * | * | 64 | 5.9
2014-11-02 | 306 | 1410 | LR | * | * | 347 | 5.9
2014-11-03 | 307 | 1353 | HJ | * | * | -303 | 5.9
2014-11-04 | 308 | 1315 | LR | * | * | -222 | 5.9
2014-11-05 | 309 | 1246 | HJ | * | * | -167 | 5.8
2014-11-06 | 310 | 1335 | HJ | * | * | 158 | 5.7
2014-11-07 | 311 | 1252 | LR | * | * | 432 | 5.8
2014-11-08 | 312 | 1307 | HJ | * | * | -70 | 5.9
2014-11-09 | 313 | 1437 | HJ | * | * | 471 | 5.8
2014-11-10 | 314 | 1300 | LR | * | * | 370 | 4.4 |  |  |  | external pump not functioning, power outage from 0700-0710UTC.
2014-11-11 | 315 | 1230 | LR | * | * | 106/-68 | 6.8/4.0 |  |  |  | both instruments are running great on the same ambient inlet. Planned power outages from 1259-1301UTC & 1303-1305UTC. AE33 was no on UPS, and lost power. It is now on UPS.
2014-11-12 | 316 | 1255 | LR | * | * | 312/185 | 7.1/4.0
2014-11-13 | 317 | 1443 | HJ | * |  | 0.51969 | 7.4/4.0
2014-11-14 | 318 | 1211 | LR | * | * | -120/378 | 7.5/4.0 |  |  |  | Planned power outage from 12:57-13:01UTC
2014-11-15 | 319 | 1319 | HJ | * | * | -1539/608 | 6.1/4.0
2014-11-16 | 320 | 1418 | HJ | * | * | -187/342 | 6.2/4.0
2014-11-17 | 321 | 1315 | LR | * | * | 12/368 | 6.0/4.0 |  |  | *
2014-11-18 | 322 | 1230 | LR | * | * | 13/89 | 4 |  |  |  | AE16 Air Flow getting low
2014-11-19 | 323 |  | HJ | [7 empty] | 30+ knot winds, did not travel to TAWO
2014-11-20 | 324 | 1247 | HJ | * | * | -255/1104 | low flow/4.0 |  |  |  | same message on AE16, "Flow getting low"
2014-11-21 | 325 | - | HJ | [7 empty] | 40+ knot winds, did not travel to TAWO
2014-11-22 | 326 | 1324 | HJ | * | * | -144/38 | low flow/4.0 |  |  |  | same message on AE16, "Flow getting low"
2014-11-23 | 327 | 1539 | HJ | * | * | 147/-244 | 5.9/4.0
2014-11-24 | 328 | 1330 | LR | * | * | 20/665 | 5.6/4.0
2014-11-25 | 329 | 1345 | LR | * | * | -93/179 | 5.6/4.0
2014-11-26 | 330 | 1449 | HJ | * | * | 32/138 | 4.7/4.0
2014-11-27 | 331 | 1356 | HJ | * | * | 32/-599 | 4.7/4.0
2014-11-28 | 332 | 1327 | HJ | * | * | 32/1021 | 4.7/4.0
2014-11-29 | 333 | [9 empty] | High winds, did not travel to TAWO
2014-11-30 | 334 | 1609 | HJ | * | * | no comms/-146 | no comms/4.0
2014-12-01 | 335 | 1337 | HJ | * | * | no comms/-710 | no comms/4.0 |  |  |  | at 13:44:47 on Nov 26 AE16 froze. At 1420utc system was restarted. 6.1lpm/0ng/m3
2014-12-02 | 336 | 1400 | HJ | * | * | -454/-506 | 6.1/4.0
2014-12-03 | 337 | 1537 | HJ | * | * | -155/733 | 6.2/4.0
2014-12-04 | 338 | 1351 | HJ | * | * | -29/269 | 6.5/4.0
2014-12-05 | 339 | 1440 | LR | * | * | 126/186 | 6.3/4.0
2014-12-06 | 340 | 1750 | LR | * | * | 0/401 | 6.3/4.0
2014-12-07 | 341 | 1637 | HJ | * | * | 225/-24 | 6.2/4.0
2014-12-08 | 342 | 1312 | HJ | * | * | -197/-856 | 6.3/4.0
2014-12-09 | 343 | 1220 | LR | * | * | 6/175 | 6.0/4.0
2014-12-10 | 344 | 1417 | HJ | * | * | -905/-7648 | 6.2/4.0
2014-12-11 | 345 | 1352 | HJ | * | * | -437/-1549 | 6.3/4.0
2014-12-12 | 346 | 1717 | LR | * | * | 3/-6 | 6.3/4.0
2014-12-13 | 347 | 1309 | HJ | * | * | 604/-762 | 5.8/4.0
2014-12-14 | 348
2014-12-15 | 349 | 1429 | HJ | * | * | 298/-988 | 6.4/4.0
2014-12-16 | 350 | 1600 | LR | * | * | -15/346 | 6.3/4.0
2014-12-17 | 351 | 1357 | HJ | * | * | -115/478 | 6.1/4.0
2014-12-18 | 352 |  |  | * | * | -68/-413 | 6.5/4.0
2014-12-19 | 353 | 1305 | LR | * | * | 92/969 | 6.4/4.0
2014-12-20 | 354 | 1535 | LR | * | * | -131/28 | 6.6/4.0
2014-12-21 | 355
2014-12-22 | 356 | 1400 | LR | * | * | -131/28 | 6.6/4.0
2014-12-23 | 357 | [9 empty] | Did not travel to TAWO
2014-12-24 | 358 | 1411 | HJ | * | * | -833/155 | 6.3/4.0
2014-12-25 | 359 | 1545 | LR | * | * | 14/572 | 6.7/4.0
2014-12-26 | 360 | 1445 | HJ | * | * | -2068/235 | 6.7/4.0
2014-12-27 | 361 | 0 | 0 | [7 empty] | Did not travel to TAWO
2014-12-28 | 362 | 1650 | HJ | * | * | -18/-288 | 6.6/4.0
2014-12-29 | 363 | 1425 | LR | * | * | -61/-77 | 6.5/4.0
2014-12-30 | 364 | 1409 | HJ | * | * | -98/735 | 6.3/4.0
2014-12-31 | 365 | 1721 | HJ | * | * | 241/761 | 6.2/4.0

## Surface TEI 03
Date | Julian Day | Time (UTC) | Observer's Initials | Inlet | Panel Display | Bench Temp | Pressure (mb) | Flow rate (LPM) | Intensity Cell A | Intensity Cell B | Comments (Alarms)
2014-10-01 | 274 | 1229 | DP | x | 38.3 | 36 | 484.9 | 438/446 | 83 | 86
2014-10-02 | 275 | 1555 | SMC | x | 35.8 | 35.9 | 488.4 | 441/449 | 83 | 86
2014-10-03 | 276 | 1123 | DP | x | 33.9 | 37.3 | 490.4 | 441/453 | 83 | 86
2014-10-04 | 277 | 1215 | SMC | x | 33.5 | 37.3 | 486.2 | 439/450 | 84 | 86
2014-10-05 | 278 | 1742 | SMC | x | 35.3 | 34.8 | 502 | 450/455 | 83 | 85
2014-10-06 | 279 | 1142 | DP | x | 37.3 | 36 | 504.5 | 451/459 | 83 | 86
2014-10-07 | 280 | 1222 | DP | x | 38.7 | 35.3 | 504.7 | 451/457 | 83 | 85
2014-10-08 | 281 | 1530 | SMC | x | 48.1 | 35.6 | 503 | 451/458 | 83 | 86
2014-10-09 | 282 | 1816 | SMC | x | 40.9 | 33 | 503.2 | 450/453 | 83 | 85
2014-10-10 | 283 | 1515 | DP | x | 38.2 | 33.2 | 506.3 | 451/456 | 83 | 85
2014-10-11 | 284 | 1300 | HJ | X | 45.9 | 36.2 | 506 | 453/460 | 83 | 85
2014-10-12 | 285 | 1429 | HJ | frost | 43 | 37.5 | 502.4 | 451/460 | 83 | 86
2014-10-13 | 286 | 1127 | HJ | frost | 49.6 | 38 | 502.4 | 451/462 | 83 | 86
2014-10-14 | 287 | 1112 | DP | x | 43.3 | 33.5 | 500.2 | 447/452 | 83 | 85
2014-10-15 | 288 | 1208 | HJ | clr | 44.5 | 39.3 | 497.4 | 448/458 | 83 | 85 | Planned power outage from 12:10-12:30UTC for turnover training
2014-10-16 | 289 | 1300 | LR | frost | 40.7 | 37.1 | 497.8 | 447/457 | 83 | 85
2014-10-17 | 290 | 1210 | HJ | frost | 47.5 | 37.7 | 497.3 | 447/457 | 83 | 85
2014-10-18 | 291 | 1350 | LR | lt frost | 40.4 | 38.7 | 492.3 | 444/456 | 83 | 86
2014-10-19 | 292 | 1358 | HJ | lt frost | 38.7 | 37.6 | 490.3 | 443/453 | 83 | 85
2014-10-20 | 293 | 1208 | HJ | lt frost | 41.1 | 38.7 | 491.5 | 443/456 | 83 | 85
2014-10-21 | 294 | 1147 | HJ | lt frost | 38.7 | 38 | 492.5 | 444/456 | 83 | 85
2014-10-22 | 295 | 1156 | HJ | frost | 47.7 | 38.8 | 485.8 | 440/452 | 83 | 85
2014-10-23 | 296 | 1142 | HJ | frost | 45 | 38.3 | 483.5 | 437/450 | 83 | 86
2014-10-24 | 297 | 1238 | HJ | frost | 40.2 | 38.8 | 480.5 | 436/449 | 83 | 86
2014-10-25 | 298 | 1251 | LR | clr | 36.5 | 37.5 | 484.5 | 438/449 | 83 | 85
2014-10-26 | 299 | 1315 | LR | clr | 32.9 | 39 | 488.9 | 442/454 | 83 | 85
2014-10-27 | 300 | 1326 | HJ | lt frost | 37.8 | 38.6 | 495.9 | 446/459 | 83 | 85
2014-10-28 | 301 | 1215 | LR | frost | 40.4 | 36.7 | 497.4 | 447/456 | 83 | 85
2014-10-29 | 302 | 1300 | LR | clr | 38.5 | 37.1 | 496.6 | 446/457 | 83 | 86
2014-10-30 | 303 | 1418 | HJ | clr | 42.1 | 37.6 | 499.1 | 448/459 | 83 | 85
2014-10-31 | 304 | 1245 | LR | clr | 45 | 36.5 | 500.5 | 449/457 | 83 | 85
2014-11-01 | 305 | 1340 | HJ | clr | 50.5 | 37.2 | 496.1 | 446/456 | 83 | 85
2014-11-02 | 306 | 1410 | LR | clr | 49.3 | 38.2 | 497.8 | 446/459 | 83 | 85
2014-11-03 | 307 | 1353 | HJ | lt frost | 51.5 | 37.6 | 496.9 | 447/458 | 83 | 85
2014-11-04 | 308 | 1315 | LR | clr | 50.5 | 38.4 | 491.1 | 442/455 | 83 | 85
2014-11-05 | 309 | 1246 | HJ | clr | 43.8 | 38.7 | 485.3 | 438/451 | 83 | 85
2014-11-06 | 310 | 1335 | HJ | lt frost | 43.1 | 38.4 | 486.4 | 439/452 | 83 | 85
2014-11-07 | 311 | 1252 | LR | clr | 43.3 | 38.8 | 484.4 | 438/451 | 83 | 85
2014-11-08 | 312 | 1307 | HJ | frost | 39.9 | 37.7 | 490.6 | 442/453 | 83 | 85
2014-11-09 | 313 | 1437 | HJ | frost | 41.6 | 37.4 | 490.1 | 441/452 | 83 | 85
2014-11-10 | 314 | 1300 | LR | rm frost | 41.5 | 38 | 487.1 | 440/452 | 83 | 85 | brief power outage from 0700-0710
2014-11-11 | 315 | 1230 | LR | clr | 41.8 | 38.1 | 488.8 | 441/453 | 83 | 85 | Planned power outages from 1259-1301UTC & 1303-1305UTC
2014-11-12 | 316 | 1255 | LR | clr | 39 | 37.6 | 494.1 | 445/456 | 83 | 85
2014-11-13 | 317 | 1443 | HJ | frost | 42.6 | 38.3 | 498.5 | 448/460 | 83 | 85
2014-11-14 | 318 | 1211 | LR | clr | 39.9 | 37.3 | 499.6 | 448/458 | 82 | 85 | Planned power outage from 12:57-13:01UTC
2014-11-15 | 319 | 1319 | HJ | clr | 37.7 | 38.4 | 498.2 | 448/459 | 82 | 85
2014-11-16 | 320 | 1418 | HJ | lt frost | 36.6 | 36.4 | 496.8 | 446/454 | 82 | 85
2014-11-17 | 321 | 1315 | LR | clr | 36.1 | 37 | 491.5 | 443/453 | 82 | 85
2014-11-18 | 322 | 1230 | LR | clr | 37.7 | 37.9 | 492 | 443/454 | 82 | 85
2014-11-19 | 323 | [9 empty] | 30+ knot winds, did not travel to TAWO
2014-11-20 | 324 | 1247 | HJ | clr | 44.1 | 37.1 | 489.1 | 440/452 | 82 | 85
2014-11-21 | 325 | - | HJ | [7 empty] | 40+ knot winds, did not travel to TAWO
2014-11-22 | 326 | 1324 | HJ | clr | 46.3 | 36 | 489.8 | 441/449 | 82 | 85
2014-11-23 | 327 | 1539 | HJ | clr | 38.5 | 37.3 | 490.6 | 442/452 | 82 | 85
2014-11-24 | 328 | 1330 | LR | clr | 29.9 | 38.8 | 485.6 | 438/451 | 82 | 85
2014-11-25 | 329 | 1345 | LR | clr | 40.5 | 37.9 | 486.2 | 439/450 | 82 | 85
2014-11-26 | 330 | 1449 | HJ | clr | 39.9 | 39.6 | 480.2 | 435/450 | 82 | 85
2014-11-27 | 331 | 1356 | HJ | lt frost | 40.2 | 37.7 | 484.3 | 437/449 | 82 | 85
2014-11-28 | 332 | 1327 | HJ | clr | 43.5 | 38.7 | 484.9 | 438/450 | 82 | 85
2014-11-29 | 333 | 0 | 0 | [7 empty] | High winds, did not travel to TAWO
2014-11-30 | 334 | 1609 | HJ | lt frost | 41 | 38.4 | 480.6 | 435/447 | 82 | 85
2014-12-01 | 335 | 1337 | HJ | lt frost | 43.9 | 38.1 | 477.7 | 433/445 | 82 | 85
2014-12-02 | 336 | 1400 | HJ | lt frost | 39.9 | 39.5 | 473.6 | 430/445 | 82 | 85
2014-12-03 | 337 | 1537 | HJ | lt frost | 42.1 | 39 | 484 | 437/451 | 82 | 85
2014-12-04 | 338 | 1351 | HJ | frost | 36.3 | 39 | 492 | 443/456 | 82 | 85
2014-12-05 | 339 | 1440 | LR | frost | 35.4 | 38.5 | 487.8 | 440/452 | 82 | 85
2014-12-06 | 340 | 1750 | LR | frost | 37.3 | 38.5 | 483.6 | 437/450 | 82 | 85
2014-12-07 | 341 | 1637 | HJ | frost | 37.9 | 37.9 | 486.3 | 439/451 | 82 | 85
2014-12-08 | 342 | 1312 | HJ | frost | 41.3 | 38.4 | 482.9 | 436/450 | 82 | 85
2014-12-09 | 343 | 1220 | LR | rm frost | 38.8 | 39 | 480.7 | 435/449 | 82 | 85
2014-12-10 | 344 | 1417 | HJ | lt frost | 38.7 | 38.9 | 486.3 | 439/452 | 82 | 85
2014-12-11 | 345 | 1352 | HJ | frost | 39.2 | 38 | 489.9 | 441/454 | 82 | 85
2014-12-12 | 346 | 1717 | LR | clr | 41.1 | 34.5 | 478.9 | 432/439 | 82 | 84
2014-12-13 | 347 | 1309 | HJ | clr | 36.9 | 39.2 | 477.4 | 433/446 | 82 | 85
2014-12-14 | 348 | na | na
2014-12-15 | 349 | 1429 | HJ | clr | 33.6 | 38.3 | 482.6 | 436/448 | 82 | 85
2014-12-16 | 350 | 1600 | LR | clr | 40.9 | 38.5 | 477.1 | 433/445 | 82 | 84
2014-12-17 | 351 | 1357 | HJ | lt frost | 41.7 | 38.2 | 478 | 433/446 | 82 | 84
2014-12-18 | 352 |  |  | lt frost | 37.8 | 38.7 | 489.3 | 441/454 | 82 | 84
2014-12-19 | 353 | 1305 | LR | clr | 38.1 | 37.3 | 486.3 | 438/449 | 81 | 84
2014-12-20 | 354 | 1535 | LR | lt frost | 39.3 | 36.7 | 485.8 | 438/447 | 81 | 84
2014-12-21 | 355
2014-12-22 | 356 | 1400 | LR | lt frost | 39.3 | 36.7 | 485.8 | 438/447 | 81 | 84
2014-12-23 | 357 | [9 empty] | Did not travel to TAWO
2014-12-24 | 358 | 1411 | HJ | rm frost | 25.6 | 38.4 | 496.5 | 447/459 | 82 | 84 | pump left on and calibration not closed out from Monday. Displsy still read Level 2. Cycled power and reset time at 1444. Redid calibration starting at 1145.
2014-12-25 | 359 | 1545 | LR | clr | 39.6 | 35.9 | 491.2 | 441/451 | 81 | 84
2014-12-26 | 360 | 1445 | HJ | lt frost | 35.9 | 35.3 | 494 | 443/450 | 81 | 83
2014-12-27 | 361 | 0 | 0 | [7 empty] | Did not travel to TAWO
2014-12-28 | 362 | 1650 | HJ | clr | 42.1 | 35.3 | 486.8 | 438/445 | 83 | 84
2014-12-29 | 363 | 1425 | LR | clr | 42.1 | 35.2 | 479.2 | 433/440 | 81 | 84
2014-12-30 | 364 | 1409 | HJ | frost | 41.1 | 38 | 479.3 | 434/445 | 82 | 84
2014-12-31 | 365 | 1721 | HJ | frost | 46.2 | 38.5 | 480.7 | 435/447 | 81 | 84

## Kramer NOx
Date | Julian Day | Time (UTC) | Observer's Initials | NOxIV computer running? | All data values within range? | O2, NO, and Air flows ok? | PO-air-O3>P1-O2-O3 (~0.5)? | Purge pump (~0.6 bar) | PMT cool temperature (<-17C) | NO Cal tank pressure | O2 1 tank pressure | O2 2 tank pressure | Breathing Air tank pressure | N2 Standard tank pressure | Comments
2014-10-01 | 274 | 1241 | DP | x | x | x | x | x | x | 825/24 | 2150/15 | 400/16 | 1400/25 | 500/10
2014-10-02 | 275 | 1557 | SMC | x | x | x | x | x | x | 825/24 | 2150/15 | 400/16 | 1400/25 | 500/10
2014-10-03 | 276 | 1134 | DP | x | x | x | x | x | x | 825/24 | 2100/15 | 250/16 | 1350/25 | 500/10
2014-10-04 | 277 | 1219 | SMC | x | x | x | x | x | x | 825/24 | 2100/15 | 2100/16 | 1300/25 | 500/10 | *changed tank
2014-10-05 | 278 | 1744 | SMC | x | x | x | x | x | x | 825/24 | 2100/15 | 2050/16 | 1300/25 | 500/10
2014-10-06 | 279 | 1155 | DP | x | x | x | x | x | x | 800/24 | 2100/15 | 2050/16 | 1300/25 | 500/10
2014-10-07 | 280 | 1234 | DP | x | x | x | x | x | x | 750/24 | 2100/15 | 2000/16 | 1300/25 | 500/10
2014-10-08 | 281 | 1536 | SMC | x | x | x | x | x | x | 750/24 | 2100/15 | 2000/16 | 1300/25 | 500/10
2014-10-09 | 282 | 1818 | SMC | x | x | x | x | x | x | 750/24 | 2100/15 | 1800/16 | 1250/25 | 500/10
2014-10-10 | 283 | 1526 | DP | x | x | x | x | x | x | 750/24 | 2100/15 | 1700/16 | 1250/25 | 500/10
2014-10-11 | 284 | 1308 | HJ | x | x | x | x | x | x | 750/24 | 2100/15 | 1600/16 | 1220/25 | 500/10
2014-10-12 | 285 | 1425 | HJ | x | x | x | x | x | x | 700/24 | 2100/15 | 1550/17 | 1200/25 | 500/10
2014-10-13 | 286 | 1133 | HJ | x | x | x | x | x | x | 700/24 | 2050/12 | 1450/15 | 1200/25 | 500/10
2014-10-14 | 287 | 1121 | DP | x | x | x | x | x | x | 700/24 | 2050/12 | 1450/15 | 1200/25 | 500/10
2014-10-15 | 288 | 1236 | HJ | x | x | x | x | x | x | 700/24 | 2000/12 | 1250/14 | 1100/25 | 500/10 | power outage for turnover prior to readings (1215-1230UTC)
2014-10-16 | 289 | 1300 | LR | x | x | x | x | x | x | 700/24 | 1900/12 | 1100/14 | 1100/25 | 400/10
2014-10-17 | 290 | 1213 | HJ | x | x | x | x | x | x | 700/24 | 1800/12 | 1000/14 | 1050/24 | 500/10
2014-10-18 | 291 | 1400 | LR | x | x | x | x | x | x | 700/24 | 2000/12 | 1000/14 | 1100/25 | 500/10
2014-10-19 | 292 | 1401 | HJ | x | x | x | x | x | x | 700/24 | 1950/12 | 950/14 | 1100/24 | 500/10
2014-10-20 | 293 | 1211 | HJ | x | x | x | x | x | x | 700/24 | 1950/12 | 850/14 | 1100/24 | 500/10
2014-10-21 | 294 | 1159 | HJ | x | x | x | x | x | x | 700/24 | 1900/12 | 750/14 | 1000/24 | 500/10
2014-10-22 | 295 | 1156 | HJ | x | x | x | adjusted | x | x | 690/24 | 1950/12 | 700/14 | 1100/24 | 500/10 | PO-airO3/P1-O2-O3 difference was .3, so adjusted
2014-10-23 | 296 | 1142 | HJ | x | x | x | adjusted | x | x | 690/24 | 1900/12 | 700/14 | 1100/24 | 500/10 | PO-airO3/P1-O2-O3 difference was still .3, so adjusted
2014-10-24 | 297 | 1238 | HJ | x | x | x | x | x | x | 690/24 | 1890/12 | 500/14 | 1000/24 | 500/10
2014-10-25 | 298 | 1251 | LR | * | * | * | * | * | * | 690/24 | 1800/12 | 400/14 | 1000/24 | 500/10
2014-10-26 | 299 | 1315 | LR | * | * | * | * | * | * | 690/24 | 1800/12 | 350/14 | 1000/24 | 500/10
2014-10-27 | 300 | 1326 | HJ | * | * | * | * | * | * | 650/24 | 1890/12 | 250/14 | 1000/24 | 500/10
2014-10-28 | 301 | 1215 | LR | * | * | * | * | * | * | 650/24 | 1890/12 | 150/14 | 1000/24 | 500/10
2014-10-29 | 302 | 1300 | LR | * | * | * | * | * | * | 650/24 | 1850/12 | 100/14 | 1000/24 | 500/10
2014-10-30 | 303 | 1418 | HJ | * | * | * | adjusted | * | * | 600/24 | 1800/12 | 2100/14 | 900/24 | 500/10 | changed O2 #2 tank. PO-airO3/P1-O2-O3 difference was still .3, so adjusted
2014-10-31 | 304 | 1245 | LR | * | * | * | * | * | * | 625/24 | 1800/12 | 2000/14 | 900/24 | 500/10
2014-11-01 | 305 | 1340 | HJ | * | * | * | * | * | * | 600/24 | 1800/12 | 2000/14 | 900/24 | 500/10
2014-11-02 | 306 | 1410 | LR | * | * | * | * | * | * | 600/24 | 1800/12 | 1900/14 | 900/24 | 500/10
2014-11-03 | 307 | 1353 | HJ | y | no | no | no | * | display off | 550/24 | 1800/12 | 1825/14 | 800/24 | 500/10 | vacuum pump off, display photo sent in email
2014-11-04 | 308 | 1315 | LR | * | * | * | * | * | * | 550/24 | 1700/12 | 1700/14 | 750/24 | 500/10
2014-11-05 | 309 | 1246 | HJ | * | * | * | * | * | * | 550/24 | 1700/12 | 1625/14 | 750/24 | 500/10
2014-11-06 | 310 | 1335 | HJ | * | * | * | * | * | * | 550/24 | 1700/12 | 1500/14 | 750/24 | 500/10
2014-11-07 | 311 | 1252 | LR | * | * | * | * | * | * | 550/24 | 1700/12 | 1400/14 | 700/24 | 500/10
2014-11-08 | 312 | 1307 | HJ | * | * | * | * | * | * | 500/24 | 1650/12 | 1300/14 | 700/24 | 500/10
2014-11-09 | 313 | 1437 | HJ | * | * | * | * | * | * | 500/24 | 1600/12 | 1200/14 | 600/24 | 500/10
2014-11-10 | 314 | 1300 | LR | * | * | * | * | * | * | 500/24 | 1600/12 | 1100/14 | 600/24 | 500/10 | brief pwr outage at 0700 UTC.
2014-11-11 | 315 | 1230 | LR | * | * | * | * | <0.6 | * | 500/24 | 1550/12 | 1100/14 | 600/24 | 500/10 | Planned power outages from 1259-1301UTC & 1303-1305UTC
2014-11-12 | 316 | 1255 | LR | * | * | * | * | <0.6 | * | 500/24 | 1550/12 | 1000/14 | 600/24 | 500/10 | Changed out CAL gas 1400UTC. LR
2014-11-13 | 317 | 1443 | HJ | * | * | * | * | 0.56 | * | 2000/24 | 1650/12 | 950/14 | 600/24 | 500/10
2014-11-14 | 318 | 1211 | LR | * | * | * | * | <.6 | * | 2000/24 | 1650/12 | 900/14 | 600/24 | 500/10 | Planned power outage from 12:57-13:01UTC
2014-11-15 | 319 | 1319 | HJ | * | * | * | * | 0.58 | * | 2000/24 | 1650/12 | 800/14 | 600/24 | 500/10
2014-11-16 | 320 | 1418 | HJ | * | * | * | * | 0.58 | * | 1990/24 | 1700/12 | 700/14 | 600/24 | 500/10
2014-11-17 | 321 | 1315 | LR | * | * | * | * | * | * | 1990/24 | 1625/12 | 650/14 | 600/24 | 500/10
2014-11-18 | 322 | 1230 | LR | * | * | * | * | * | * | 1990/24 | 1600/12 | 550/14 | 600/24 | 500/10
2014-11-19 | 323 |  | HJ | [11 empty] | 30+ knot winds, did not travel to TAWO
2014-11-20 | 324 | 1247 | HJ | * | * | * | * | * | * | 1950/24 | 1600/12 | 350/14 | 500/24 | 500/10
2014-11-21 | 325 | - | HJ | [11 empty] | 40+ knot winds, did not travel to TAWO
2014-11-22 | 326 | 1324 | HJ | * | * | * | * | * | * | 1900/24 | 1600/12 | 200/15 | 500/24 | 500/10
2014-11-23 | 327 | 1539 | HJ | * | * | * | * | * | * | 1900/24 | 1600/14 | 100/14 | 500/24 | 500/10
2014-11-24 | 328 | 1330 | LR | * | * | * | * | * | * | 1900/24 | 1550/12 | >100/14 | 500/24 | 500/10
2014-11-25 | 329 | 1345 | LR | * | * | * | * | * | * | 1925/24 | 1500/12 | 2100/14 | 500/24 | 500/10
2014-11-26 | 330 | 1449 | HJ | * | * | * | no- same | * | * | 1925/24 | 1425/12 | 1950/14 | 500/24 | 500/10
2014-11-27 | 331 | 1356 | HJ | * | * | * | * | * | * | 1900/24 | 1300/12 | 1800/14 | 500/24 | 500/10
2014-11-28 | 332 | 1327 | HJ | * | * | * | * | * | * | 1900/24 | 1400/12 | 1850/14 | 500/24 | 500/10
2014-11-29 | 333 | 0 | 0 | [11 empty] | High winds, did not travel to TAWO
2014-11-30 | 334 | 1609 | HJ | * | * | * | * | * | * | 1900/24 | 1400/12 | 1700/24 | 500/24 | 500/10
2014-12-01 | 335 | 1337 | HJ | * | * | * | * | * | * | 1875/24 | 1300/12 | 1600/14 | 450/24 | 500/10
2014-12-02 | 336 | 1400 | HJ | * | * | * | 0.2 | * | * | 1875/24 | 1300/12 | 1550/14 | 450/24 | 500/10
2014-12-03 | 337 | 1537 | HJ | * | * | * | * | * | * | 1875/24 | 1250/12 | 1450/14 | 450/24 | 500/10
2014-12-04 | 338 | 1351 | HJ | * | * | * | * | * | * | 1850/24 | 1250/12 | 1375/14 | 400/24 | 500/10
2014-12-05 | 339 | 1440 | LR | * | * | * | * | * | * | 1850/24 | 1210/12 | 1350/14 | 400/24 | 500/10
2014-12-06 | 340 | 1750 | LR | * | * | * | * | * | * | 1850/24 | 1200/12 | 1300/14 | 400/24 | 500/10
2014-12-07 | 341 | 1637 | HJ | * | * | * | * | * | * | 1825/24 | 1100/14 | 1150/14 | 350/24 | 500/10
2014-12-08 | 342 | 1312 | HJ | * | * | * | * | * | * | 1800/24 | 1150/14 | 1150/14 | 350/24 | 500/10
2014-12-09 | 343 | 1220 | LR | * | * | * | * | * | * | 1800/24 | 1100/13 | 1100/14 | 350/24 | 500/10
2014-12-10 | 344 | 1417 | HJ | * | * | * | * | * | * | 1800/24 | 1000/13 | 1000/14 | 200/25 | 500/10
2014-12-11 | 345 | 1352 | HJ | * | * | * | * | * | * | 1800/24 | 1000/14 | 900/14 | 200/25 | 500/10
2014-12-12 | 346 | 1717 | LR | * | * | * | * | * | * | 1800/24 | 950/14 | 800/14 | 150/25 | 500/10 | Changed out Breathing Air (2650psi/26psi).
2014-12-13 | 347 | 1309 | HJ | * | * | * | adjusted | * | * | 1800/24 | 1000/13 | 800/14 | 2650/25 | 500/10
2014-12-14 | 348 | na | na
2014-12-15 | 349 | 1429 | HJ | * | * | * | adjusted | * | * | 1750/24 | 950/13 | 700/14 | 2550/25 | 500/10
2014-12-16 | 350 | 1600 | LR | * | * | * | * | * | * | 1750/24 | 850/13 | 2050/14 | 2400/25 | 500/10
2014-12-17 | 351 | 1357 | HJ | * | * | * | adjusted | * | * | 1750/24 | 800/13 | 2025/14 | 2300/25 | 500/10
2014-12-18 | 352
2014-12-19 | 353 | 1305 | LR | * | * | * | * | * | * | 1750/24 | 750/13 | 2100/14 | 2400/25 | 500/10
2014-12-20 | 354 | 1535 | LR | * | * | * | * | * | * | 1725/24 | 700/13 | 2100/14 | 2350/25 | 500/10
2014-12-21 | 355
2014-12-22 | 356 | 1400 | LR | * | * | * | * | * | * | 1700/24 | 550/13 | 1950/14 | 2200/25 | 500/10
2014-12-23 | 357 | [13 empty] | Did not travel to TAWO
2014-12-24 | 358 | 1411 | HJ | * | * | * | * | * | * | 1700/24 | 400/14 | 1650/14 | 1900/26 | 500/10
2014-12-25 | 359 | 1545 | LR | * | * | * | * | * | * | 1700/24 | 400/14 | 1800/14 | 2100/26 | 500/10
2014-12-26 | 360 | 1445 | HJ | * | * | * | * | * | * | 1690/24 | 350/13 | 1750/13 | 2050/26 | 500/10
2014-12-27 | 361 | 0 | 0 | [11 empty] | Did not travel to TAWO
2014-12-28 | 362 | 1650 | HJ | * | * | * | * | * | * | 1650/24 | 200/14 | 1650/14 | 1950/26 | 500/10
2014-12-29 | 363 | 1425 | LR | * | * | * | * | * | * | 1650/24 | 150/14 | 1600/14 | 1975/26 | 500/10
2014-12-30 | 364 | 1409 | HJ | * | * | * | * | * | * | 1650/24 | 100/14 | 1500/14 | 1900/26 | 500/10
2014-12-31 | 365 | 1721 | HJ | * | see note | * | * | * | * | 1625/24 | 100/12 | 1550/13 | 2000/26 | 500/10 | P-O2 flashing between green and red. Values of 9.96-10.06

## NOAA CATS GC
|  |  |  | Cylinder Pressures |  |  |  |  | CATS Gas Chromatograph | [9 empty] | CATS Data Computer | [8 empty] | Chromatograms
Date | Julian Day | Time (UTC) | Observer's Initials | N2 (high/low psig) | P5 (high psig) | Cal. 1 (high/low psig) | Cal. 2 (high/low psig) | N2O (high/low psig) |  | N2 Purifier Hot? | ECD1 Temp (350C) | ECD2 TEMP (314C) | ECD3 TEMP (350C) | ECD4 TEMP (350C) | Column 1 Temp (80C) | Column 2 Temp (85C) | Column 3 Temp (70C) | Column 4 Temp (62C) |  | Main Col. 1 Flow (22±5 sccm) | Main Col. 2 Flow (34±5 sccm) | Main Col. 3 Flow (19±5 sccm) | Main Col. 4 Flow (45±5 sccm) | B.F. Col. 1 Flow (28±5 sccm) | B.F. Col. 2 Flow (30±5 sccm) | B.F. Col. 3 Flow (18±5 sccm) | B.F. Col. 4 Flow (40±5 sccm) | Channel 1 | Channel 2 | Channel 3 | Channel 4 | Comments
2014-10-01 | 274 | 1246 | DP | 2000/61 | 1350 | 400/25 | 700/26 | 650/19
2014-10-02 | 275 | 1556 | SMC | 1900/61 | 1300 | 400/25 | 700/26 | 650/19
2014-10-03 | 276 | 1136 | DP | 1900/61 | 1300 | 400/25 | 700/26 | 650/19
2014-10-04 | 277 | 1217 | SMC | 1900/61 | 1300 | 400/24 | 700/26 | 650/19
2014-10-05 | 278 | 1743 | SMC | 1600/61 | 1200 | 400/24 | 700/26 | 650/19
2014-10-06 | 279 | 1205 | DP | 1500/61 | 1190 | 400/25 | 700/26 | 650/19
2014-10-07 | 280 | 1237 | DP | 1450/61 | 1190 | 400/25 | 700/26 | 650/19
2014-10-08 | 281 | 1535 | SMC | 1400/61 | 1190 | 400/25 | 700/26 | 650/19
2014-10-09 | 282 | 1817 | SMC | 1300/61 | 1100 | 400/25 | 700/26 | 650/19
2014-10-10 | 283 | 1528 | DP | 1190/61 | 1000 | 400/25 | 700/26 | 650/19
2014-10-11 | 284 | 1303 | HJ | 1100/61 | 950 | 400/25 | 700/26 | 650/19
2014-10-12 | 285 | 1447 | HJ | 1000/61 | 900 | 400/25 | 700/26 | 650/19
2014-10-13 | 286 | 1148 | HJ | 1000/60 | 900 | 400/25 | 700/28 | 650/19
2014-10-14 | 287 | 1125 | DP | 850/61 | 890 | 400/25 | 700/28 | 650/19
2014-10-15 | 288 | 1140 | HJ | 800/61 | 850 | 400/25 | 690/28 | 650/19 | [23 empty] | **Note: On Wed, 10/15, there was a planned power outage for turnover training from 12:10-12:30 UTC***
2014-10-16 | 289 | 1310 | LR | 700/61 | 800 | 400/25 | 675/25 | 650/19
2014-10-17 | 290 | 1226 | HJ | 600/60 | 750 | 350/25 | 675/25 | 650/19
2014-10-18 | 291 | 1400 | LR | 500/60 | 700 | 390/25 | 675/25 | 650/19
2014-10-19 | 292 | 1430 | HJ | 400/60 | 650 | 390/25 | 690/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-10-20 | 293 | 1219 | HJ | 300/60 | 625 | 350/25 | 650/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-10-21 | 294 | 1220 | HJ | 250/60 | 600 | 350/25 | 650/25 | 650/19
2014-10-22 | 295 | 1156 | HJ | 100/60 | 600 | 350/25 | 650/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-10-23 | 296 | 1142 | HJ | 2500/62 | 500 | 350/25 | 650/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | computer was shut down, found in position 2. after replacing N2 cylinder, system was back to normal.
2014-10-24 | 297 | 1238 | HJ | 2500/62 | 500 | 350/25 | 650/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-10-25 | 298 | 1251 | LR | 2400/62 | 400 | 350/25 | 650/25 | 650/19 |  | 😉 | 😉 | 😉 | 😉 | 😉 | 😉 | 😉 | 😉 | 😉 |  | 😉 | 😉 | 😉 | 😉 | 😉 | 😉 | 😉 | 😉 | 😉 | 😉 | 😉 | 😉
2014-10-26 | 299 | 1315 | LR | 2300/62 | 350 | 350/25 | 650/25 | 650/19 |  | 😉 | 😉 | 😉 | 😉 | 😉 | 😉 | 😉 | 😉 | 😉 |  | 😉 | 😉 | 😉 | 😉 | 😉 | 😉 | 😉 | 😉 | 😉 | 😉 | 😉 | 😉
2014-10-27 | 300 | 1326 | HJ | 2200/62 | 350 | 350/25 | 650/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-10-28 | 301 | 1215 | LR | 2100/62 | 300 | 350/25 | 650/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😉 | 😉 | 😉 | 😉
2014-10-29 | 302 | 1300 | LR | 2000/62 | 300 | 350/25 | 650/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-10-30 | 303 | 1418 | HJ | 1900/61 | 200 | 350/25 | 650/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-10-31 | 304 | 1245 | LR | 1900/62 | 200 | 350/25 | 650/24 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | changed out P5 tank @1330utc
2014-11-01 | 305 | 1340 | HJ | 1800/61 | 2150 | 350/25 | 650/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-11-02 | 306 | 1410 | LR | 1600/61 | 2100 | 350/25 | 650/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-11-03 | 307 | 1353 | HJ | 1500/61 | 2050 | 350/25 | 625/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-11-04 | 308 | 1315 | LR | 1400/61 | 2000 | 350/25 | 625/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-11-05 | 309 | 1246 | HJ | 1400/61 | 2000 | 350/25 | 650/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 82.4😥 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-11-06 | 310 | 1335 | HJ | 1400/61 | 2000 | 350/25 | 650/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-11-07 | 311 | 1252 | LR | 1200/61 | 1950 | 350/25 | 650/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-11-08 | 312 | 1307 | HJ | 1100/61 | 1900 | 350/25 | 625/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-11-09 | 313 | 1437 | HJ | 1000/61 | 1800 | 350/61 | 600/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-11-10 | 314 | 1300 | LR | 950/61 | 1800 | 350/25 | 600/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | brief power outage from 0700-0710
2014-11-11 | 315 | 1230 | LR | 900/61 | 1800 | 300/25 | 600/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | Planned power outages from 1259-1301UTC & 1303-1305UTC
2014-11-12 | 316 | 1255 | LR | 800/61 | 1700 | 300/25 | 600/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-11-13 | 317 | 1443 | HJ | 650/61 | 1700 | 300/25 | 600/25 | 650/19
2014-11-14 | 318 | 1211 | LR | 600/61 | 1650 | 300/25 | 600/25 | 650/19 | [23 empty] | Planned power outage from 12:57-13:01UTC
2014-11-15 | 319 | 1319 | HJ | 500/60 | 1600 | 300/25 | 625/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-11-16 | 320 | 1418 | HJ | 400/61 | 1550 | 300/25 | 625/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-11-17 | 321 | 1315 | LR | 300/61 | 1500 | 300/25 | 600/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-11-18 | 322 | 1230 | LR | 200/61 | 1500 | 300/25 | 600/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-11-19 | 323 |  | HJ | [28 empty] | 30+ knot winds, did not travel to TAWO
2014-11-20 | 324 | 1247 | HJ | 100/61 | 1400 | 300/25 | 600/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | changed out N2 tank. Ending 100psi, new tank 2500psi
2014-11-21 | 325 | - | HJ | [28 empty] | 40+ knot winds, did not travel to TAWO
2014-11-22 | 326 | 1324 | HJ | 2300/61 | 1300 | 300/25 | 600/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😃 | 😃 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-11-23 | 327 | 1539 | HJ | 2200/61 | 1300 | 300/25 | 600/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😃 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😃 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😃
2014-11-24 | 328 | 1330 | LR | 2100/61 | 1250 | 300/25 | 600/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😃 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😃 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😃
2014-11-25 | 329 | 1345 | LR | 2000/61 | 1200 | 300/25 | 600/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😃 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😃 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😃
2014-11-26 | 330 | 1449 | HJ | 2000/61 | 1200 | 300/25 | 600/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-11-27 | 331 | 1356 | HJ | 1850/61 | 1150 | 300/25 | 600/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😃 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-11-28 | 332 | 1327 | HJ | 1750/61 | 1100 | 290/25 | 600/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-11-29 | 333 | 0 | 0 | [28 empty] | High winds, did not travel to TAWO
2014-11-30 | 334 | 1609 | HJ | 1600/61 | 1000 | 290/25 | 575/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-12-01 | 335 | 1337 | HJ | 1500/61 | 1000 | 290/25 | 575/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | changed cal1 tank. New pressure at 1900.
2014-12-02 | 336 | 1400 | HJ | 1400/61 | 950 | 1950/25 | 575/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-12-03 | 337 | 1537 | HJ | 1300/61 | 900 | 1950/25 | 575/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-12-04 | 338 | 1351 | HJ | [24 empty] | 😊 | 😊 | 😊 | 😊
2014-12-05 | 339 | 1440 | LR | 1100/61 | 800 | 1900/25 | 575/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-12-06 | 340 | 1750 | LR | 1050/61 | 800 | 1900/25 | 575/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-12-07 | 341 | 1637 | HJ | 1000/61 | 750 | 1950/25 | 575/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-12-08 | 342 | 1312 | HJ | 900/61 | 700 | 1950/25 | 550/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-12-09 | 343 | 1220 | LR | 800/61 | 650 | 1900/25 | 550/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-12-10 | 344 | 1417 | HJ | 700/61 | 625 | 1950/25 | 550/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-12-11 | 345 | 1352 | HJ | 650/61 | 600 | 1925/25 | 550/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-12-12 | 346 | 1717 | LR | 550/61 | 550 | 1925/25 | 550/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-12-13 | 347 | 1309 | HJ | 400/61 | 500 | 1925/25 | 550/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-12-14 | 348 | NA | NA
2014-12-15 | 349 | 1429 | HJ | 300/61 | 425 | 1900/25 | 550/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-12-16 | 350 | 1600 | LR | 200/61 | 400 | 1925/25 | 550/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-12-17 | 351 | 1357 | HJ | 2500/61 | 350 | 1900/25 | 550/27 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | changed N2 tank 1420-1422utc
2014-12-18 | 352
2014-12-19 | 353 | 1305 | LR | 2350/60 | 250 | 1925/25 | 550/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-12-20 | 354 | 1535 | LR | 2200/60 | 200 | 1925/25 | 550/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | changed P5 tank 1550UTC. LR
2014-12-21 | 355
2014-12-22 | 356 | 1400 | LR | 2000/60 | 2150 | 1900/25 | 550/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-12-23 | 357 | [30 empty] | Did not travel to TAWO
2014-12-24 | 358 | 1411 | HJ | 1850/61 | 2050 | 1900/25 | 550/24 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-12-25 | 359 | 1545 | LR | 1750/60 | 2000 | 1900/25 | 550/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-12-26 | 360 | 1445 | HJ | 1650/61 | 1950 | 1900/25 | 550/24 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-12-27 | 361 | 0 | 0 | [28 empty] | Did not travel to TAWO
2014-12-28 | 362 | 1650 | HJ | 1500/61 | 1890 | 1900/25 | 550/24 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-12-29 | 363 | 1425 | LR | 1400/60 | 1850 | 1900/25 | 550/24 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-12-30 | 364 | 1409 | HJ | 1350/61 | 1800 | 1900/25 | 550/24 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊
2014-12-31 | 365 | 1721 | HJ | 1250/61 | 1790 | 1900/25 | 525/25 | 650/19 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 |  | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊 | 😊

## DRUM Sampler
Date | Julian Day | Time (UTC) | Observer's Initials | Inlet Clear? | Pump running/temp OK? | Drum Rotating? | Vacuum (in HG) | "Flow" (WC)
2014-10-01 | 274 | 1241 | DP | * | * | *
2014-10-02 | 275 | 1600 | SMC | * | * | *
2014-10-03 | 276 | 1140 | DP | * | * | *
2014-10-04 | 277 | 1218 | SMC | * | * | *
2014-10-05 | 278 | 1745 | SMC | * | * | *
2014-10-06 | 279 | 1200 | DP | * | * | * | 5 | 12
2014-10-07 | 280 | 1237 | DP | * | * | *
2014-10-08 | 281 | 1539 | SMC | * | * | *
2014-10-09 | 282 | 1820 | SMC | * | * | *
2014-10-10 | 283 | 1530 | DP | * | * | *
2014-10-11 | 284 | 1320 | HJ | * | * | *
2014-10-12 | 285 | 1456 | HJ | * | * | *
2014-10-13 | 286 | 1125 | HJ | * | * | * | 5 | 12
2014-10-14 | 287 | 1124 | DP | * | * | *
2014-10-15 | 288 | 1145 | HJ | * | * | *
2014-10-16 | 289 | 1300 | LR | * | * | *
2014-10-17 | 290 | 1231 | HJ | * | * | *
2014-10-18 | 291 | 1345 | LR | * | * | *
2014-10-19 | 292 | 1358 | HJ | * | * |  | 5 | 12
2014-10-20 | 293 | 1210 | HJ | * | * | * | 6 | 12
2014-10-21 | 294 | 1148 | HJ | * | * |  | 5 | 12
2014-10-22 | 295 | 1156 | HJ | * | * | * | 5 | 12
2014-10-23 | 296 | 1142 | HJ | * | * |  | 5 | 11
2014-10-24 | 297 | 1238 | HJ | * | * |  | 5 | 11
2014-10-25 | 298 | 1251 | LR | * | * |  | 5 | 12
2014-10-26 | 299 | 1315 | LR | * | * |  | 5 | 12
2014-10-27 | 300 | 1326 | HJ | * | * |  | 5 | 12
2014-10-28 | 301 | 1215 | LR | * | * |  | 5 | 12
2014-10-29 | 302 | 1300 | LR | * | * |  | 5 | 12
2014-10-30 | 303 | 1418 | HJ | * | * |  | 5 | 12
2014-10-31 | 304 | 1245 | LR | * | * |  | 5 | 12
2014-11-01 | 305 | 1340 | HJ | * | * |  | 5 | 12
2014-11-02 | 306 | 1410 | LR | * | * |  | 5 | 12
2014-11-03 | 307 | 1353 | HJ | * | * |  | 5 | 12
2014-11-04 | 308 | 1315 | LR | * | * |  | 5 | 12
2014-11-05 | 309 | 1246 | HJ | * | * |  | 6 | 11
2014-11-06 | 310 | 1335 | HJ | * | * |  | 5 | 11
2014-11-07 | 311 | 1252 | LR | * | * |  | 5 | 11
2014-11-08 | 312 | 1307 | HJ | * | * |  | 5 | 11
2014-11-09 | 313 | 1437 | HJ | * | * |  | 5 | 11
2014-11-10 | 314 | 1300 | LR | * | * |  | 5 | 11 | brief power outage from 0700-0710
2014-11-11 | 315 | 1230 | LR | * | * |  | 5 | 11 | Planned power outages from 1259-1301UTC & 1303-1305UTC
2014-11-12 | 316 | 1255 | LR | * | * |  | 5 | 11.5
2014-11-13 | 317 | 1443 | HJ | * | * |  | 5 | 11.5
2014-11-14 | 318 | 1211 | LR | * | * |  | 5 | 12 | Planned power outage from 12:57-13:01UTC
2014-11-15 | 319 | 1319 | HJ | * | * |  | 5 | 12
2014-11-16 | 320 | 1418 | HJ | * | * |  | 5 | 12
2014-11-17 | 321 | 1315 | LR | * | * |  | 5 | 12
2014-11-18 | 322 | 1230 | LR | * | * |  | 5 | 12
2014-11-19 | 323 | [7 empty] | 30+ knot winds, did not travel to TAWO
2014-11-20 | 324 | 1247 | HJ | * | * |  | 5 | 12
2014-11-21 | 325 | - | HJ | [5 empty] | 40+ knot winds, did not travel to TAWO
2014-11-22 | 326 | 1324 | HJ | * | * |  | 5 | 12
2014-11-23 | 327 | 1539 | HJ | * | * |  | 5 | 12
2014-11-24 | 328 | 1330 | LR | * | * |  | 5 | 12
2014-11-25 | 329 | 1345 | LR | * | * |  | 5 | 12
2014-11-26 | 330 | 1449 | HJ | * | * |  | 5 | 12
2014-11-27 | 331 | 1356 | HJ | * | * |  | 5 | 12 | seems to be quite loud from outside of TAWO
2014-11-28 | 332 | 1327 | HJ | * | * |  | 5 | 12
2014-11-29 | 333 | 0 | 0 | [5 empty] | High winds, did not travel to TAWO
2014-11-30 | 334 | 1609 | HJ | * | * |  | 5 | 12
2014-12-01 | 335 | 1337 | HJ | * | * |  | 5 | 11
2014-12-02 | 336 | 1400 | HJ | * | * |  | 5 | 12
2014-12-03 | 337 | 1537 | HJ | * | * |  | 5 | 12
2014-12-04 | 338 | 1351 | HJ | * | * |  | 5 | 12
2014-12-05 | 339 | 1440 | LR | * | * |  | 5 | 12
2014-12-06 | 340 | 1750 | LR | * | * |  | 5 | 12
2014-12-07 | 341 | 1637 | HJ | * | * |  | 5 | 11
2014-12-08 | 342 | 1312 | HJ | * | * |  | 5 | 11
2014-12-09 | 343 | 1220 | LR | * | * |  | 5 | 12
2014-12-10 | 344 | 1417 | HJ | * | * |  | 5 | 12
2014-12-11 | 345 | 1352 | HJ | * | * |  | 5 | 11
2014-12-12 | 346 | 1717 | LR | * | * |  | 5 | 11
2014-12-13 | 347 | 1309 | HJ | * | * |  | 5 | 11
2014-12-14 | 348 | na | na
2014-12-15 | 349 | 1429 | HJ | * | * |  | 5 | 12
2014-12-16 | 350 | 1600 | LR | * | * |  | 5 | 11
2014-12-17 | 351 | 1357 | HJ | * | * |  | 5 | 12
2014-12-18 | 352
2014-12-19 | 353 | 1305 | LR | * | * |  | 5 | 12
2014-12-20 | 354 | 1535 | LR | * | * |  | 5 | 12
2014-12-21 | 355
2014-12-22 | 356 | 1400 | LR | * | * |  | 5 | 12
2014-12-23 | 357 | 0 | 0 | [5 empty] | Did not travel to TAWO
2014-12-24 | 358 | 1411 | HJ | * | * |  | 5 | 11
2014-12-25 | 359 | 1545 | LR | * | * |  | 5 | 12
2014-12-26 | 360 | 1445 | HJ | * | * |  | 5 | 12
2014-12-27 | 361 | 0 | 0 | [5 empty] | Did not travel to TAWO
2014-12-28 | 362 | 1650 | HJ | * | * |  | 5 | 11
2014-12-29 | 363 | 1425 | LR | * | * |  | 5 | 11
2014-12-30 | 364 | 1409 | HJ | could not remove | * |  | 5 | 11
2014-12-31 | 365 | 1721 | HJ |  | * |  | 5 | 12