## NEPH & CLAP
Date | Julian Day | Time (UTC) | Observer's Initials | Sample inlet clear? | Check Daily Aerosol Email | NEPH Lamp On/Current (<7A) | NEPH Flag Status | NEPH Flow (~18 lpm) | CLAP flow ~0.60 lpm | CLAP tr>0.7 | CLAP Spot # | Bi-Weekly Span Check (Mon) | Comments
2016-01-01 | 1 | 1510 | JG | clr | * | 5.5 | 100 | 18 | 0.61 | 0.797 | 8
2016-01-02 | 2 | 1229 | NB | rm frost | * | 5.5 | 100 | 18 | 0.62 | 0.795 | 8 |  | Adjusted CLAP flow to .60.
2016-01-03 | 3 | 1310 | JG | rm frost | * | 5.5 | 100 | 18 | 0.61 | 0.795 | 8
2016-01-04 | 4 | 1228 | NB | rm frost | * | 5.5 | 100 | 18 | 0.62 | 0.795 | 8 | * | Adjusted CLAP flow to .60. HEPA filter on 1225-1326z. CO2 tank pressure 420 psi. Span check error 1.81%.
2016-01-05 | 5 | 1230 | JG | clr | * | 5.5 | 100 | 18 | 0.6 | 0.795 | 8
2016-01-06 | 6 | 1307 | NB | rm frost | * | 5.5 | 100 | 18 | 0.6 | 0.79 | 8
2016-01-07 | 7 | 1235 | JG | rm frost | * | 5.5 | 100 | 18 | 0.61 | 0.788 | 8
2016-01-08 | 8 | 1347 | NB | clr | * | 5.5 | 100 | 18 | 0.6 | 0.78 | 8
2016-01-09 | 9 | 1240 | JG | lt frost | * | 5.5 | 100 | 18 | 0.6 | 0.764 | 8 |  | Started CLAP filter change at 1245z. System did not autodetect end of change; tried another filter and waited 20 minutes before pressing red button to force it to finish at 1354z.
2016-01-10 | 10 | 1504 | NB | rm frost | * | 5.5 | 100 | 18 | 0.59 | 0.987 | 1
2016-01-11 | 11 | 1240 | JG | rm frost | * | 5.5 | 100 | 18 | 0.6 | 0.885 | 1
2016-01-12 | 12 | 1232 | NB | rm frost | * | 5.5 | 100 | 18 | 0.62 | 0.88 | 1 |  | Performed CLAP white filter change 1238-1244z. Adjusted CLAP flow to .60
2016-01-13 | 13 | 1255 | JG | rm frost | * | 5.5 | 100 | 18 | 0.6 | 0.843 | 2 |  | Planned power outages for e-gen test 1307-1309 and 1321-1324 UTC. Carbon vane pump did not come back on after power outage - was quite hot. Restarted pump at 1357 UTC; Neph and CLAP flow came back up to normal.
2016-01-14 | 14 | 1225 | NB | rm frost | * | 5.5 | 100 | 18 | 0.59 | 0.747 | 2
2016-01-15 | 15 | 1255 | JG | rm frost | * | 5.5 | 100 | 18 | 0.59 | 0.851 | 3 |  | N winds last 4 days.
2016-01-16 | 16 | 1304 | NB | rm frost | * | 5.5 | 100 | 18 | 0.6 | 0.848 | 3
2016-01-17 | 17 | 1725 | JG | rm frost | * | 5.5 | 100 | 18 | 0.6 | 0.844 | 3
2016-01-18 | 18 | 1229 | NB | clr | * | 5.5 | 100 | 18 | 0.59 | 0.841 | 3 | * | HEPA filter installed 1229-1330z. Span check started 1332z. CO2 tank pres 400psi. Span check error 2.10%.
2016-01-19 | 19 | 1310 | JG | rm frost | * | 5.5 | 100 | 18 | 0.59 | 0.838 | 3
2016-01-20 | 20 | 1232 | NB | clr | * | 5.5 | 0 | 18 | 0.59 | 0.834 | 3
2016-01-21 | 21 | 1250 | JG | rm frost | * | 5.5 | 100 | 18 | 0.59 | 0.832 | 3
2016-01-22 | 22 | 1234 | NB | lt frost | * | 5.5 | 0 | 18 | 0.57 | 0.829 | 3 |  | Adjusted CLAP flow to .60 LPM.
2016-01-23 | 23 | 1210 | JG | rm frost | * | 5.5 | 100 | 18 | 0.59 | 0.826 | 3
2016-01-24 | 24 | 1258 | NB | lt frost | no | 5.5 | 6100 | 18 | 0.59 | 0.822 | 3
2016-01-25 | 25 | 1230 | JG | rm frost | * | 5.5 | 100 | 18 | 0.58 | 0.821 | 3 |  | Adjusted CLAP flow to .60 LPM.
2016-01-26 | 26 | 1254 | NB | rm frost | * | 5.5 | 100 | 18 | 0.6 | 0.819 | 3
2016-01-27 | 27 | 1340 | JG | rm frost | * | 5.5 | 100 | 18 | 0.6 | 0.803 | 3 |  | End piece for inlet found on snow sfc. Brought in to thaw, reinstalled at 1400z.
2016-01-28 | 28 | 1234 | NB | lt frost | * | 5.5 | 100 | 18 | 0.61 | 0.782 | 3
2016-01-29 | 29 | 1255 | JG | rm frost | * | 5.5 | 100 | 18 | 0.61 | 0.762 | 3
2016-01-30 | 30 | 1225 | NB | rm frost | * | 5.5 | 100 | 18 | 0.6 | 0.717 | 3
2016-01-31 | 31 | 1455 | JG | rm frost | * | 5.4 | 0 | 18 | 0.62 | 0.96 | 4 |  | Adjusted CLAP flow to .60 LPM.
2016-02-01 | 32 | 1225 | NB | lt frost | * | 5.4 | 100 | 18 | 0.6 | 0.932 | 4 | no | Bi-weekly span check moved to next Monday 2/8/16 in order to coincide with turnover.
2016-02-02 | 33 | 1240 | JG | rm frost | * | 5.4 | 100 | 18 | 0.61 | 0.981 | 1 |  | Spot changed from 4 to 1; system appears to have detected a filter change yesterday and reverted to spot 1. Changed filter today 1713-1719z to give it a fresh start.
2016-02-03 | 34 | 1248 | NB | rm frost | * | 5.4 | 100 | 18 | 0.6 | 0.996 | 1
2016-02-04 | 35 | 1220 | JG | rm frost | * | 5.4 | 100 | 18 | 0.59 | 0.993 | 1
2016-02-05 | 36 | 1333 | NB | rm frost | * | 5.4 | 100 | 18 | 0.59 | 0.991 | 1
2016-02-06 | 37 | 1355 | JG | rm frost | * | 5.4 | 100 | 18 | 0.56 | 0.989 | 1 |  | Adjusted CLAP flow to .60 LPM.
2016-02-07 | 38 | 1455 | MB/JJ | rm frost | * | 5.6 | 2100 | 18 | 0.6 | 0.984 | 1
2016-02-08 | 39 | 1240 | MB/JJ/SD | rm frost | * | 5.7 | 100 | 18 | 0.62 | 0.981 | 1 | * | HEPA filter installed 1233-1340z. Span check started 1340z. CO2 tank pres 400psi. Span check error 2.78%. Adjusted CLAP flow to 0.60.
2016-02-09 | 40 | 1230 | MB/JJ/NB | rm frost | * | 5.6 | 100 | 18 | 0.59 | 0.979 | 1
2016-02-10 | 41 | 1226 | SWD/JG/NB | rm frost | * | 5.6 | 100 | 18 | 0.6 | 0.983 | 1
2016-02-11 | 42 | 1340 | MB/JJ | rm frost | * | 5.6 | 100 | 18 | 0.59 | 0.981 | 1
2016-02-12 | 43 | 1625 | MB/JJ | rm frost | * | 5.6 | 100 | 18 | 0.59 | 0.972 | 1
2016-02-13 | 44 | 1220 | MB/JJ | rm frost | * | 5.6 | 100 | 18 | 0.6 | 0.952 | 1
2016-02-14 | 45 | 1700 | MB/JJ | rm frost | * | 5.6 | 100 | 18 | 0.6 | 0.912 | 1
2016-02-15 | 46 | 1641 | JJ | rm frost | * | 5.6 | 100 | 18 | 0.59 | 0.905 | 1
2016-02-16 | 47 | 1321 | jj/mb | rm frost | * | 5.6 | 100 | 18 | 0.58 | 0.903 | 1
2016-02-17 | 48 | 1415 | MB/JJ | rm frost | * | 5.6 | 100 | 18 | 0.59 | 0.898 | 1
2016-02-18 | 49 | 1405 | MB/JJ | clr | * | 5.6 | 100 | 18 | 0.59 | 0.89 | 1
2016-02-19 | 50 | 1349 | JJ/MB | rm frost | * | 5.5 | 100 | 17 | 0.58 | 0.888 | 1 |  | Adjusted CLAP flow to 0.60 LPM.
2016-02-20 | 51 | 1311 | JJ/MB | rm frost | * | 5.5 | 100 | 17 | 0.6 | 0.887 | 1
2016-02-21 | 52 | 1510 | JJ/MB | rm frost | * | 5.5 | 100 | 17 | 0.62 | 0.885 | 1 |  | Adjusted CLAP flow to 0.60 LPM.
2016-02-22 | 53 | 1244 | JJ/MB | rm frost | * | 5.6 | 100 | 17 | 0.6 | 0.876 | 1 | * | HEPA filter installed 1240-1343z. Span check started 1349z. CO2 tank pres 420psi. Span check error 1.88%.
2016-02-23 | 54 | 1704 | MB | rm frost | * | 5.9 | 0 | 18 | 0.62 | 0.867 | 1 |  | Adjusted CLAP flow to 0.60 LPM.
2016-02-24 | 55 | 1314 | JJ/MB | clr | * | 6 | 100 | 18 | 0.6 | 0.855 | 1
2016-02-25 | 56 | 1336 | jj/mb | clr | * | 5.9 | 100 | 18 | 0.6 | 0.833 | 1
2016-02-26 | 57 | 1348 | JJ/MB | rm frost | * | 5.9 | 100 | 18 | 0.6 | 0.816 | 1
2016-02-27 | 58 | 1305 | MB | rm frost | * | 6 | 100 | 18 | 0.6 | 0.803 | 1
2016-02-28 | 59 | 1708 | JJ | rm frost | * | 5.9 | 100 | 18 | 0.59 | 0.7991 | 1
2016-02-29 | 60 | 1339 | jj/mb | rm frost | * | 5.9 | 100 | 18 | 0.58 | 0.796 | 1 |  | adjusted CLAP flow to 0.60 lpm
2016-03-01 | 61 | 1248 | JJ/MB | rm frost | * | 6 | 100 | 18 | 0.6 | 0.789 | 1
2016-03-02 | 62 | 1711 | MB/JJ | rm frost | * | 5.9 | 100 | 18 | 0.6 | 0.782 | 1
2016-03-03 | 63 | 1338 | MB/JJ | rm frost | * | 5.9 | 100 | 18 | 0.61 | 0.78 | 1
2016-03-04 | 64 | 1309 | jj/mb | rm frost | * | 5.9 | 100 | 18 | 0.61 | 0.777 | 1
2016-03-05 | 65 | 1235 | jj/mb | rm frost | * | 5.9 | 100 | 18 | 0.61 | 0.772 | 1
2016-03-06 | 66 | 1342 | MB | rm frost | * | 6.2 | 100 | 17 | 0.6 | 0.76 | 1
2016-03-07 | 67 | 1321 | jj/mb | rm frost | * | 6.2 | 100 | 17 | 0.6 | 0.75 | 1 | * | HEPA filter installed 1317-1417. Span check started 1418z. CO2 tank pres 400psi. Span check error 2.04%. Adjusted CLAP flow to .60lpm 1322 UTC
2016-03-08 | 68 | 1331 | jj/mb | clr | * | 6.2 | 100 | 18 | 0.61 | 0.746 | 1
2016-03-09 | 69 | 1235 | JJ | clr | * | 6.1 | 100 | 17 | 0.62 | 0.743 | 1 |  | Adjusted CLAP flow to 0.60 lpm 1355 UTC
2016-03-10 | 70 | [11 empty] | White out conditions (Condition 1) did not travel to TAWO this day.
2016-03-11 | 71 | 1352 | JJ | clr | * | 6.4 | 100 | 17 | 0.57 | 0.7337 | 1 |  | Adjusted CLAP flow to 0.60 lpm 1238 and 1417 UTC (before and after power outage)
2016-03-12 | 72 | 1437 | MB | rm frost | * | 6.5 | 100 | 17 | 0.62 | 0.723 | 1 |  | Adjusted CLAP flow to 0.60 lpm 1439z
2016-03-13 | 73 | 1346 | JJ | clr | * | 6.5 | 100 | 17 | 0.6 | 0.718 | 1 |  | unexpected power outage today from 13:16 to 13:23 UTC
2016-03-14 | 74 | 1353 | jj/mb | clr | * | 6.5 | 100 | 17 | 0.59 | 0.711 | 1
2016-03-15 | 75 | 1533 | jj/mb | rm frost | * | 6.5 | 100 | 17 | 0.64 | 0.998 | 2 |  | Adjusted CLAP flow to .60 lpm at 15:45 UTC
2016-03-16 | 76 | 1319 | MB | clr | * | 6.4 | 100 | 18 | 0.61 | 0.986 | 2
2016-03-17 | 77 | 1338 | jj/mb | clr | * | 6.8 | 0 | 18 | 0.61 | 0.975 | 2
2016-03-18 | 78 | 1305 | MB/JJ | clr | * | 6.7 | 0 | 18 | 0.61 | 0.967 | 2
2016-03-19 | 79 | 1854 | MB/JJ | rm frost | * | 6.7 | 100 | 17 | 0.6 | 0.936 | 2
2016-03-20 | 80 | 1334 | MB | rm frost | * | 6.7 | 100 | 18 | 0.6 | 0.932 | 2
2016-03-21 | 81 | 1307 | jj/mb | rm frost | * | 6.7 | 100 | 18 | 0.6 | 0.931 | 2 | * | HEPA filter installed 1302-1402z. Span check started 1405z. CO2 tank pres 400psi. Span check error 3.53%.
2016-03-22 | 82 | 1320 | MB | rm frost | * | 6.7 | 100 | 18 | 0.6 | 0.929 | 2
2016-03-23 | 83 | 1744 | MB/JJ | rm frost | * | 6.7 | 100 | 18 | 0.6 | 0.924 | 2
2016-03-24 | 84 | 1329 | jj/mb | rm frost | * | 6.6 | 100 | 18 | 0.59 | 0.92 | 2
2016-03-25 | 85 | 1303 | JJ | rm frost | * | 6.6 | 100 | 18 | 0.59 | 0.918 | 2
2016-03-26 | 86 | 1303 | MB | rm frost | * | 6.6 | 100 | 18 | 0.59 | 0.769 | 2
2016-03-27 | 87 | 1715 | JJ | clr | * | 6.6 | 100 | 18 | 0.59 | 0.752 | 2
2016-03-28 | 88 | 1327 | MB | clr | * | 6.6 | 100 | 17 | 0.59 | 0.745 | 2
2016-03-29 | 89 | 1151 | JJ | rm frost | * | 6.6 | 100 | 17 | 0.6 | 0.721 | 2
2016-03-30 | 90 | 1134 | MB | rm frost | * | 6.6 | 100 | 18 | 0.6 | 0.707 | 2
2016-03-31 | 91 | 1718 | jj/mb | rm frost | * | 6.6 | 100 | 18 | 0.61 | 0.995 | 3
2016-04-01 | 92 | 1304 | JJ | clr | * | 6.8 | 100 | 17 | 0.6 | 0.983 | 3
2016-04-02 | 93 | 1213 | jj/mb | clr | * | 6.8 | 100 | 18 | 0.61 | 0.972 | 3
2016-04-03 | 94 | 1550 | MB | rm frost | * | 6.8 | 100 | 18 | 0.61 | 0.969 | 3
2016-04-04 | 95 | 1613 | jj/mb | rm frost | * | 6.8 | 100 | 18 | 0.62 | 0.907 | 3 | * | HEPA filter installed 1611-1711 Span check started 1715z. CO2 tank pres 380psi. Span check error 2.12%. Adjusted CLAP flow to 0.60 lpm @ 1624z
2016-04-05 | 96 | 1251 | MB/JJ | rm frost | * | 6.7 | 100 | 18 | 0.59 | 0.894 | 3
2016-04-06 | 97 | 1332 | JJ | clr | * | 6.7 | 100 | 18 | 0.59 | 0.775 | 3 |  | Planned power outage today. Adjusted CLAP flow to .60lpm
2016-04-07 | 98 | 1224 | MB/JJ | clr | * | 6.7 | 100 | 18 | 0.6 | 0.762 | 3
2016-04-08 | 99 | 1139 | JJ/MB | clr | * | 6.7 | 100 | 18 | 0.61 | 0.755 | 3
2016-04-09 | 100 | 1254 | JJ/MB | clr | * | 6.7 | 100 | 18 | 0.62 | 0.713 | 3 |  | adjusted CLAP flow to 0.60 lpm
2016-04-10 | 101 | 1707 | JJ | clr | * | 6.6 | 0 | 18 | 0.6 | 0.993 | 4
2016-04-11 | 102 | 1321 | JJ/MB | clr | * | 6.6 | 0 | 18 | 0.61 | 0.974 | 4
2016-04-12 | 103 | 1554 | JJ/MB | clr | * | 6.7 | 0 | 18 | 0.61 | 0.96 | 4
2016-04-13 | 104 | 1317 | JJ/MB | clr | * | 6.6 | 100 | 18 | 0.6 | 0.933 | 4
2016-04-14 | 105 | 1329 | JJ/MB | rm frost | * | 6.8 | 0 | 18 | 0.6 | 0.914 | 4 |  | Missing curved intake tube for Neph, CLAP, and AE33.
2016-04-15 | 106 | 1711 | JJ/MB/SWD | clr | * | 6.7 | 0 | 18 | 0.59 | 0.894 | 4 |  | Missing curved intake tube for Neph, CLAP, and AE33.
2016-04-16 | 107 | 1309 | JJ/MB | rm frost | * | 6.8 | 100 | 18 | 0.58 | 0.884 | 4 |  | Missing curved intake tube for Neph, CLAP, and AE33. Adjusted CLAP flow to 0.60 lpm.
2016-04-17 | 108 | 1628 | MB | clr | * | 6.8 | 100 | 18 | 0.6 | 0.876 | 4 |  | Missing curved intake tube for Neph, CLAP, and AE33.
2016-04-18 | 109 | 1249 | MB | rm frost | * | 6.7 | 100 | 18 | 0.6 | 0.862 | 4 | * | Missing curved intake tube for Neph, CLAP, and AE33. HEPA filter installed without curved intake tube 1234-1334. Span check started 1335z. CO2 tank pres 360psi. Span check error 13.39%. Redoing Span check due to high percent error. Span check started 1408z. CO2 tank pres 320 psi. Span check error 2.25%.
2016-04-19 | 110 | 1152 | JJ/MB | clr | * | 6.7 | 100 | 18 | 0.6 | 0.852 | 4 |  | Missing curved intake tube for Neph, CLAP, and AE33.
2016-04-20 | 111 | 1143 | JJ/MB | clr | * | 6.6 | 100 | 18 | 0.61 | 0.845 | 4 |  | Missing curved intake tube for Neph, CLAP, and AE33.
2016-04-21 | 112 | 1122 | JJ/MB/SWD | rm frost | * | 6.8 | 100 | 18 | 0.62 | 0.839 | 4 |  | Missing curved intake tube for Neph, CLAP, and AE33. Adjusted CLAP flow to 0.60 lpm.
2016-04-22 | 113 | 1838 | JJ/MB | clr | * | 6.3 | 0 | 18 | 0.6 | 0.819 | 4 |  | Missing curved intake tube for Neph, CLAP, and AE33.
2016-04-23 | 114 | 1536 | [10 empty] | White out conditions (Condition 1) did not travel to TAWO this day.
2016-04-24 | 115 | 1536 | JJ/MB | clr | * | 8.1 | 0 | 18 | 0.6 | 0.813 | 4 |  | Missing curved intake tube for Neph, CLAP, and AE33. Changed Neph lamp bulb 1612-1629z. Neph lamp current 5.7 amps, lamp voltage 13.2 V.
2016-04-25 | 116 | 1207 | JJ/MB | clr | * | 5.7 | 0 | 19 | 0.61 | 0.806 | 4 |  | Missing curved intake tube for Neph, CLAP, and AE33. Adjusted CLAP flow to .60 lpm
2016-04-26 | 117 | 1130 | JJ | rm frost | * | 5.8 | 0 | 19 | 0.62 | 0.797 | 4 |  | Missing curved intake tube for Neph, CLAP, and AE33.
2016-04-27 | 118 | 1549 | JJ/MB | clr | * | 5.7 | 0 | 19 | 0.6 | 0.851 | 5 |  | Missing curved intake tube for Neph, CLAP, and AE33.
2016-04-28 | 119 | 1315 | JJ/MB | rm frost | * | 5.7 | 0 | 19 | 0.6 | 0.848 | 5 |  | Missing curved intake tube for Neph, CLAP, and AE33.
2016-04-29 | 120 | 1300 | MB/JJ | rm frost | * | 5.7 | 0 | 19 | 0.59 | 0.841 | 5 |  | Attached teflon tape and new curved intake tube made out of 1/2" electrical conduit to Neph/CLAP/AE33 intake cyclone.
2016-04-30 | 121 | 1336 | MB/JJ | rm frost | * | 5.7 | 100 | 19 | 0.6 | 0.827 | 5
2016-05-01 | 122 | 1624 | MB/JJ | clr | * | 5.7 | 100 | 18 | 0.59 | 0.793 | 5
2016-05-02 | 123 | 1322 | JJ/MB | rm frost | * | 5.8 | 100 | 18 | 0.59 | 0.779 | 5 | * | HEPA filter installed 1320-1422 Span check started 1605z. CO2 tank pres 310psi. Span check error 2.02%. Planned power outages starting from 1635 to 1700z
2016-05-03 | 124 | 1230 | MB/JJ | rm frost | * | 5.7 | 100 | 18 | 0.58 | 0.767 | 5 |  | Adjusted CLAP flow to 0.60.
2016-05-04 | 125 | 1451 | MB/JJ | clr | * | 5.8 | 100 | 18 | 0.6 | 0.726 | 5
2016-05-05 | 126 | 1528 | MB | clr | * | 5.7 | 100 | 18 | 0.6 | 0.949 | 6
2016-05-06 | 127 | 1845 | JJ/MB | clr | * | 5.7 | 100 | 18 | 0.6 | 0.924 | 6
2016-05-07 | 128 | 1315 | JJ/MB | rm frost | * | 5.7 | 100 | 18 | 0.6 | 0.918 | 6
2016-05-08 | 129 | 1558 | JJ | rm frost | * | 5.7 | 2100 | 18 | 0.61 | 0.911 | 6
2016-05-09 | 130 | 1635 | JJ/MB | clr | * | 5.8 | 0 | 18 | 0.61 | 0.895 | 6
2016-05-10 | 131 | 1648 | JJ/MB | clr | * | 5.7 | 0 | 18 | 0.59 | 0.881 | 6
2016-05-11 | 132 | 1753 | MB/JJ | clr | * | 5.7 | 0 | 18 | 0.61 | 0.84 | 6
2016-05-12 | 133 | 1622 | MB/JJ | rm frost | * | 5.8 | 0 | 19 | 0.61 | 0.834 | 6
2016-05-13 | 134 | 1308 | MB/JJ | clr | * | 5.7 | 0 | 19 | 0.61 | 0.828 | 6
2016-05-14 | 135 | 1149 | JJ/MB | clr | * | 5.8 | 0 | 19 | 0.59 | 0.811 | 6
2016-05-15 | 136 | 1730 | MB | clr | * | 5.7 | 100 | 18 | 0.59 | 0.783 | 6
2016-05-16 | 137 | 1609 | JJ/MB | clr | * | 5.7 | 100 | 18 | 0.59 | 0.73 | 6
2016-05-17 | 138 | 1722 | JJ/MB | clr | * | 5.7 | 0 | 18 | 0.58 | 0.705 | 6 |  | Ajusted CLAP flow to .60 slpm
2016-05-18 | 139 | 1309 | JJ/MB | clr | * | 5.7 | 0 | 18 | 0.59 | 0.96 | 7
2016-05-19 | 140 | 1322 | JJ | clr | * | 5.7 | 0 | 18 | 0.59 | 0.904 | 7
2016-05-20 | 141 | 1850 | JJ/MB | clr | * | 5.7 | 0 | 18 | 0.59 | 0.864 | 7
2016-05-21 | 142 | 1228 | JJ/MB | rm frost | * | 5.7 | 0 | 18 | 0.59 | 0.845 | 7
2016-05-22 | 143 | 2327 | JJ | clr | * | 5.7 | 100 | 18 | 0.59 | 0.747 | 7
2016-05-23 | 144 | 1656 | JJ | clr | * | 5.7 | 6000 | 18 | 0.59 | 0.732 | 7
2016-05-24 | 145 | 1211 | JJ/MB | rm frost | * | 5.7 | 100 | 19 | 0.59 | 0.717 | 7
2016-05-25 | 146 | 1726 | JJ | clr | * | 5.7 | 0 | 18 | 0.57 | 0.706 | 7 |  | Adjusted CLAP flow to .60 slpm
2016-05-26 | 147 | 1232 | JJ/MB | clr | * | 5.7 | 100 | 18 | 0.6 | 0.988 | 8
2016-05-27 | 148 | 1845 | JJ/mb | clr | * | 5.7 | 100 | 18 | 0.6 | 0.882 | 8 |  | Changed CLAP filter.
2016-05-28 | 149 | 1221 | MB | clr | * | 5.7 | 100 | 18 | 0.63 | 0.981 | 1 |  | Adjusted CLAP flow to 0.60 slpm.
2016-05-29 | 150 | 1530 | MB | clr | * | 5.7 | 100 | 18 | 0.6 | 0.892 | 1
2016-05-30 | 151 | 1226 | JJ | clr | * | 5.7 | 100 | 18 | 0.6 | 0.878 | 1 |  | Planned power outage today from 2130 to 2230 Aerosol instruments lost flow during outage.
2016-05-31 | 152 | 1731 | JJ | clr | * | 5.8 | 100 | 18 | 0.6 | 0.856 | 1
2016-06-01 | 153 | 1840 | JJ | clr | * | 5.7 | 100 | 18 | 0.6 | 0.835 | 1
2016-06-02 | 154 | 1831 | MB/JJ | clr | * | 5.7 | 0 | 19 | 0.61 | 0.817 | 1
2016-06-03 | 155 | 1737 | MB/JJ/HJ | clr | * | 5.8 | 0 | 18 | 0.61 | 0.81 | 1
2016-06-04 | 156 | 1156 | MB | clr | * | 5.7 | 6000 | 19 | 0.6 | 0.805 | 1
2016-06-05 | 157 | 1403 | JJ/TG/EB | clr | * | 5.7 | 0 | 18 | 0.59 | 0.756 | 1 |  | Attached new curved intake tube
2016-06-06 | 158 | 1613 | MB/JJ/TG/EB | clr | * | 5.7 | 100 | 18 | 0.6 | 0.719 | 1 |  | HEPA filter installed 1610-1713 Span check started 1718z. CO2 tank pres 300psi. Span check error 1.93%.
2016-06-07 | 159 | 1536 | TG/EB/JJ | clr | * | 5.75 | 0 | 18 | 0.61 | 0.998 | 2
2016-06-08 | 160 | 1309 | TG/EB/JJ/MB | clr | * | 5.7 | 0 | 18 | 0.61 | 0.976 | 2
2016-06-09 | 161 | 1553 | TG/EB | clr | * | 5.7 | 0 | 18 | 0.61 | 0.959 | 2
2016-06-10 | 162 | 1239 | TG/EB | clr | * | 5.8 | 0 | 18 | 0.6 | 0.942 | 2 |  | Work on ventilation duct from 1105 to 1330, higher room temps.
2016-06-11 | 163 | 1203 | TG/EB | clr | * | 5.7 | 0 | 18 | 0.6 | 0.8562 | 2
2016-06-12 | 164 | 1509 | TG/EB | clr | * | 5.7 | 0 | 18 | 0.6 | 0.844 | 2
2016-06-13 | 165 | 1637 | TG/EB | clr | * | 5.7 | 0 | 18 | 0.6 | 0.836 | 2
2016-06-14 | 166 | 1153 | TG/EB | clr | * | 5.7 | 0 | 18 | 0.59 | 0.831 | 2
2016-06-15 | 167 | 1202 | TG/EB | clr | * | 5.7 | 0 | 19 | 0.59 | 0.8267 | 2
2016-06-16 | 168 | 1156 | TG | clr | * | 5.7 | 6000 | 19 | 0.59 | 0.82 | 2 |  | Flags went to 0000 shortly after checking
2016-06-17 | 169 | 1314 | TG | clr | * | 5.7 | 0 | 18 | 0.59 | 0.8128 | 2
2016-06-18 | 170 | 1241 | TG/EB | clr | * | 5.7 | 0 | 18 | 0.59 | 0.806 | 2
2016-06-19 | 171 | 1808 | EB | clr | * | 5.7 | 0 | 18 | 0.59 | 0.788 | 2
2016-06-20 | 172 | 1812 | TG/EB | clr | * | 5.7 | 0 | 18 | 0.59 | 0.7765 | 2
2016-06-21 | 173 | 1204 | TG/EB | clr | * | 5.7 | 0 | 18 | 0.58 | 0.7692 | 2
2016-06-22 | 174 | 1332 | TG/EB | clr | * | 5.7 | 0 | 18 | 0.58 | 0.763 | 2
2016-06-23 | 175 | 1313 | EB | rm frost | * | 5.7 | 0 | 18 | 0.6 | 0.759 | 2
2016-06-24 | 176 | 1329 | TG | clr | * | 5.7 | 0 | 18 | 0.59 | 0.7562 | 2
2016-06-25 | 177 | 1717 | EB/TG | clr | * | 5.7 | 0 | 18 | 0.6 | 0.752 | 2
2016-06-26 | 178 | 1828 | TG | clr | * | 5.7 | 0 | 18 | 0.6 | 0.7506 | 2
2016-06-27 | 179 | 1603 | EB/TG | clr | * | 5.7 | 0 | 18 | 0.6 | 0.749 | 2
2016-06-28 | 180 | 1452 | EB/TG | clr | * | 5.7 | 0 | 18 | 0.6 | 0.746 | 2
2016-06-29 | 181 | 1135 | EB/TG | clr | * | 5.7 | 0 | 18 | 0.6 | 0.742 | 2
2016-06-30 | 182 | 1143 | TG/EB | clr | * | 5.7 | 0 | 18 | 0.6 | 0.735 | 2 |  | Not receiving/sending data? Computer restarted?
2016-07-01 | 183 | 1202 | TG/EB | clr | * | 5.7 | 0 | 18 | 0.6 | 0.91 | 4 |  | Programs had closed. Re-opened and appear to be reading. 1224z neph and clap turned off in prep for outside TAWO construction, TAWO drilling began at 1730z
2016-07-02 | 184 | 1239 | TG/EB | clr | * | - | - | - | - | - | - |  | Neph and Clap remained off for tower construction 20:01 turned neph and clap back on. Communication regained. However, pump colgan wouldn't work so there was no flow. Turned neph and clap back off. Cutting of metal grate at TAWO began at 1209z
2016-07-03 | 185 | 1835 | EB | clr | * | 5.7 | 0 | 18 | 0.61 | 0.902 | 4
2016-07-04 | 186 | 1029 | TG | rm frost | * | 5.7 | 0 | 18 | 0.6 | 0.904 | 4 |  | Neph and Clap turned off 1036z due to construction at TAWO instruments and pumps turned back on 14:03z
2016-07-05 | 187 | 1207 | TG/EB | clr | * | 5.7 | 0 | 18 | 0.6 | 0.732 | 5
2016-07-06 | 188 | 11:19 | TG | clr | * | 5.7 | 0 | 18 | 0.6 | 0.7162 | 5
2016-07-07 | 189 | 1214 | TG/EB | clr | * | 5.7 | 0 | 18 | 0.6 | 0.9067 | 6 |  | Neph span check: HEPA filter on from 1131-1231, span check at 1232, CO2 tank pressure 270 psi. Error 2.27
2016-07-08 | 190 | 1604 | TG/EB | clr | * | - | - | - | - | - | - |  | Instruments have disappeared we suspect Kathy in the kitchen with a blender
2016-07-09 | 191 | 1342 | TG | clr | * | - | - | - | - | - | - |  | Instruments have disappeared we suspect Thano in the galley with scrabble pieces
2016-07-10 | 192 | 1319 | TG | clr | * | - | - | - | - | - | - |  | Colgan is acting up and Andy and Christy have it turned off, will be working on it today/ Brought out pump Dignana to be changed out
2016-07-11 | 193 | 1323 | EB | clr | * | 5.65 | 0 | 18 | 0.6 | 0.999 | 8
2016-07-12 | 194 | 1228 | TG/EB | clr | * | 5.6 | 0 | 18 | 0.6 | 0.855 | 8
2016-07-13 | 195 | 1137 | TG | clr | * | 5.7 | 0 | 18 | 0.58 | 0.8714 | 8
2016-07-14 | 196 | 1147 | TG/EB | clr | * | 5.6 | 0 | 18 | 0.57 | 0.863 | 8 |  | Clap flow set to 0.6
2016-07-15 | 197 | 1235 | TG/EB | clr | * | 5.6 | 0 | 20 | 0.61 | 0.854 | 8
2016-07-16 | 198 | 1301 | TG/EB | clr | * | 5.6 | 2000 | 19 | 0.6 | 0.844 | 8
2016-07-17 | 199 | 1627 | EB | clr | * | 5.6 | 0 | 19 | 0.6 | 1 | 1 |  | Clap filter changed, flow adjusted to 0.60
2016-07-18 | 200 | 1207 | TG | clr | * | 5.6 | 0 | 19 | 0.6 | 0.9921 | 1 |  | in and out of N winds, but construction has continued on station power outage 2048-54z
2016-07-19 | 201 | 1237 | TG | clr | * | 5.6 | 0 | 20 | 0.6 | 0.9628 | 1
2016-07-20 | 202 | 1609 | TG/EB | clr | * | 5.6 | 0 | 19 | 0.6 | 0.942 | 1
2016-07-21 | 203 | 1224 | TG/EB | clr | * | 5.6 | 0 | 19 | 0.6 | 0.917 | 1
2016-07-22 | 204 | 1141 | TG | rm frost | * | 5.6 | 0 | 19 | 0.6 | 0.8619 | 1
2016-07-23 | 205 | 1317 | TG | clr | * | 5.6 | 0 | 19 | 0.6 | 0.8406 | 1
2016-07-24 | 206 | 1747 | TG | clr | * | 5.6 | 0 | 19 | 0.6 | 0.8269 | 1
2016-07-25 | 207 | 1234 | TG/EB | clr | * | 5.6 | 0 | 19 | 0.6 | 0.793 | 1
2016-07-26 | 208 | 1214 | TG/EB | clr | * | 5.6 | 0 | 19 | 0.6 | 0.777 | 1
2016-07-27 | 209 | [11 empty] | Didn't make it out to TAWO
2016-07-28 | 210 | 1313 | TG/EB | clr | * | 5.6 | 0 | 19 | 0.59 | 0.733 | 1
2016-07-29 | 211 | 1611 | TG/EB | clr | * | 5.6 | 0 | 19 | 0.6 | 0.728 | 1
2016-07-30 | 212 | 1236 | TG/EB | clr | * | 5.6 | 0 | 19 | 0.6 | 0.943 | 2
2016-07-31 | 213 | 1521 | EB | clr | * | 5.6 | 0 | 19 | 0.6 | 0.709 | 2
2016-08-01 | 214 | 1254 | EB | clr | * | 5.6 | 2000 | 19 | 0.6 | 0.704 | 2
2016-08-02 | 215 | 1259 | EB | clr | * | 5.6 | 2100 | 19 | 0.6 | 0.872 | 3
2016-08-03 | 216 | 1121 | TG | clr | * | 5.6 | 0 | 19 | 0.6 | 0.801 | 3
2016-08-04 | 217 | 1400 | EB | clr | * | 5.6 | 2000 | 19 | 0.6 | 0.965 | 4
2016-08-05 | 218 | 1353 | EB/TG | clr | * | 5.6 | 0 | 19 | 0.59 | 0.805 | 4
2016-08-06 | 219 | 1717 | EB | clr | * | 5.6 | 0 | 19 | 0.59 | 0.694 | 4
2016-08-07 | 220 | 1632 | TG | clr | * | 5.6 | 0 | 19 | 0.59 | 0.789 | 4
2016-08-08 | 221 | 1541 | TG | clr | * | 5.6 | 0 | 19 | 0.6 | 0.785 | 4
2016-08-09 | 222 | 1832 | TG | clr | * | 5.6 | 0 | 19 | 0.6 | 0.781 | 4
2016-08-10 | 223 | 1400 | EB | clr | * | 5.6 | 2000 | 19 | 0.6 | 0.771 | 4 |  | HEPA filter on at 15:50z, but disconnected at unknown time. No span check performed.
2016-08-11 | 224 | 1243 | EB | clr | * | 5.6 | 0 | 19 | 0.6 | 0.761 | 4
2016-08-12 | 225 | 1534 | TG | clr | *` | 5.6 | 0 | 19 | 0.59 | 0.7508 | 4 |  | Power outage 1609-1611z. 1642-1644z | H2: 500psi Power outage 1609-1611z. 1642-1644z
2016-08-13 | 226 | 1755 | EB | clr | * | 5.6 | 0 | 19 | 0.59 | 0.745 | 4
2016-08-14 | 227 | 1759 | EB | clr | * | 5.5 | 6000 | 19 | 0.59 | 0.738 | 4
2016-08-15 | 228 | 1907 | TG/EB | clr | * | 5.6 | 0 | 19 | 0.59 | 0.7266 | 4
2016-08-16 | 229 | 1810 | TG | clr | * | 5.5 | 0 | 19 | 0.59 | 0.724 | 4
2016-08-17 | 230 | 1652 | TG | rm frost | * | 5.6 | 0 | 19 | 0.59 | 0.722 | 4
2016-08-18 | 231 | 1731 | TG | clr | * | 5.6 | 0 | 19 | 0.59 | 0.718 | 4
2016-08-19 | 232 | 1612 | EB/TG | clr | * | 5.6 | 0 | 19 | 0.6 | 0.7111 | 4 | * | HEPA filter on at 16:03z, off at 17:15z. Span check begun at 17:17z. CO2 cylinder pressure: 200psi, starting another SPAN check 1824-1855z. Final span check error 1.63.
2016-08-20 | 233 | 1222 | TG | clr | * | 5.5 | 0 | 19 | 0.6 | 0.706 | 4
2016-08-21 | 234 | 1545 | EB | clr | * | 5.5 | 0 | 19 | 0.61 | 0.705 | 4
2016-08-22 | 235 | 1634 | TG | clr | * | 5.5 | 0 | 19 | 0.61 | 0.9962 | 5
2016-08-23 | 236 | 1255 | EB | rm frost | * | 5.5 | 0 | 19 | 0.6 | 0.979 | 5 |  | Adjusted CLAP flow from 0.63 to 0.60
2016-08-24 | 237 | 1330 | TG | rm frost | * | 5.6 | 0 | 19 | 0.59 | 0.9686 | 5
2016-08-25 | 238 | 1223 | EB | clr | * | 5.5 | 0 | 19 | 0.59 | 0.959 | 5
2016-08-26 | 239 | 1352 | TG | clr | * | 5.5 | 0 | 19 | 0.59 | 0.9546 | 5
2016-08-27 | 240 | 1744 | EB | clr | * | 5.5 | 100 | 19 | 0.59 | 0.901 | 5
2016-08-28 | 241 | 1715 | TG | clr | * | 5.5 | 0 | 19 | 0.58 | 0.8538 | 5
2016-08-29 | 242 | 1354 | EB | clr | * | 5.5 | 0 | 19 | 0.58 | 0.827 | 5
2016-08-30 | 243 | 1300 | TG | rm frost | * | 5.5 | 2000 | 19 | 0.6 | 0.7929 | 5
2016-08-31 | 244 | 1306 | EB | clr | * | 5.5 | 0 | 19 | 0.58 | 0.79 | 5
2016-09-01 | 245 | 1316 | TG | rm frost | * | 5.5 | 0 | 19 | 0.59 | 0.7868 | 5
2016-09-02 | 246 | 1700 | EB | clr | * | 5.5 | 2100 | 19 | 0.59 | 0.783 | 5
2016-09-03 | 247 | 1323 | TG | rm frost | * | 5.5 | 100 | 19 | 0.6 | 0.7753 | 5 |  | changed to spot 1 while I was here. Weird! Why not spot #6?? Moved clap below Neph as per Betsy's instructions. Manually advanced to spot #6 1410z
2016-09-04 | 248 | 1419 | EB | rm frost | * | 5.5 | 100 | 19 | 0.6 | 0.999 | 6
2016-09-05 | 249 | 1250 | TG | rm frost | * | 5.5 | 0 | 19 | 0.6 | 0.9973 | 6 |  | Adjusted rotameter so that Q_Q11 flow reads reads 9.8lpm 1351z. Stopped AE33 at 1400z, pulled inlet tube, restarted at 1401z, stopped AE33 at 1405z, put inlet tube back in, restarted at AE33 at 1406z
2016-09-06 | 250 | 1342 | EB | clr | * | 5.5 | 2000 | 19 | 0.6 | 0.996 | 6
2016-09-07 | 251 | 1243 | TG | rm frost | * | 5.5 | 6000 | 19 | 0.6 | 0.9941 | 6
2016-09-08 | 252 | 1237 | EB | clr | * | 5.5 | 100 | 19 | 0.6 | 0.992 | 6
2016-09-09 | 253 | 1249 | TG | rm hv frost | * | 5.5 | 100 | 19 | 0.6 | 0.9903 | 6
2016-09-10 | 254 | 1233 | EB | clr | * | 5.5 | 0 | 19 | 0.6 | 0.988 | 6
2016-09-11 | 255 | 1742 | TG | rm frost | * | 5.5 | 0 | 19 | 0.6 | 0.9766 | 6
2016-09-12 | 256 | 1237 | EB | clr | * | 5.5 | 0 | 19 | 0.6 | 0.976 | 6
2016-09-13 | 257 | 1237 | TG | rm hv frost | * | 5.4 | 100 | 19 | 0.6 | 0.9747 | 6
2016-09-14 | 258 | 1246 | EB | rm lt frost | * | 5.4 | 0 | 19 | 0.6 | 0.972 | 6
2016-09-15 | 259 | 1220 | TG | rm frost | * | 5.6 | 0 | 19 | 0.61 | 0.9706 | 6
2016-09-16 | 260 | 1340 | EB | rm frost | * | 5.7 | 2000 | 19 | 0.58 | 0.97 | 6
2016-09-17 | 261 | 1247 | TG | rm frost | * | 5.6 | 100 | 19 | 0.59 | 0.9671 | 6
2016-09-18 | 262 | 1457 | EB | rm frost | * | 5.6 | 100 | 19 | 0.59 | 0.94 | 6
2016-09-19 | 263 | 1323 | TG | rm frost | * | 5.6 | 100 | 19 | 0.6 | 0.9371 | 6
2016-09-20 | 264 | 1324 | EB | rm frost | * | 5.6 | 100 | 19 | 0.6 | 0.935 | 6
2016-09-21 | 265 | 1226 | TG | rm frost | * | 5.6 | 0 | 19 | 0.6 | 0.932 | 6 |  | HEPA filter on at 1206z. Off at 132. CO2 tank pressure = 150psi Span Check Error = 2.33%
2016-09-22 | 266 | 1318 | EB | rm frost | * | 5.6 | 100 | 19 | 0.6 | 0.916 | 6
2016-09-23 | 267 | 1216 | TG | rm lt frost | * | 5.6 | 100 | 19 | 0.6 | 0.8847 | 6
2016-09-24 | 268 | 1409 | EB | clr | * | 5.6 | 0 | 19 | 0.6 | 0.881 | 6
2016-09-25 | 269 | 1337 | TG | clr | * | 5.6 | 100 | 19 | 0.6 | 0.8784 | 6
2016-09-26 | 270 | 1257 | EB | rm frost | * | 5.6 | 2100 | 19 | 0.6 | 0.876 | 6
2016-09-27 | 271 | 1207 | TG | rm frost | * | 5.6 | 0 | 19 | 0.6 | 0.8207 | 6
2016-09-28 | 272 | 1325 | EB | rm frost | * | 5.6 | 0 | 19 | 0.6 | 0.834 | 6
2016-09-29 | 273 | 1247 | TG | rm lt frost | * | 5.6 | 100 | 19 | 0.6 | 0.824 | 6
2016-09-30 | 274 | 1402 | EB | clr | * | 5.6 | 2100 | 19 | 0.6 | 0.806 | 6
2016-10-01 | 275 | 1323 | TG | clr | * | 5.6 | 100 | 19 | 0.6 | 0.803 | 6
2016-10-02 | 276 | 1403 | EB | clr | * | 5.6 | 100 | 19 | 0.6 | 0.801 | 6
2016-10-03 | 277 | 1510 | TG | clr | * | 5.6 | 100 | 19 | 0.6 | 0.7995 | 6
2016-10-04 | 278 | 1219 | EB | rm lt frost | * | 5.6 | 100 | 19 | 0.6 | 0.798 | 6
2016-10-05 | 279 | 1408 | TG | rm hv frost | * | 5.6 | 100 | 19 | 0.6 | 0.7943 | 6
2016-10-06 | 280 | 1314 | EB | clr | * | 5.6 | 0 | 19 | 0.6 | 0.792 | 6
2016-10-07 | 281 | 1303 | TG | rm windblown snow | * | 5.6 | 100 | 19 | 0.6 | 0.7892 | 6
2016-10-08 | 282 | [11 empty] | High winds/low visibility, didn't make it out to TAWO
2016-10-09 | 283 | 1459 | TG | rm frost | * | 5.5 | 2100 | 19 | 0.6 | 0.7785 | 6
2016-10-10 | 284 | 1644 | EB | lt frost | * | 5.8 | 100 | 19 | 0.6 | 0.777 | 6
2016-10-11 | 285 | 1338 | TG | rm frost | * | 5.8 | 100 | 19 | 0.6 | 0.7756 | 6
2016-10-12 | 286 | [11 empty] | High winds/low visitbility
2016-10-13 | 287 | 1303 | TG | clr | * | 5.7 | 0 | 19 | 0.59 | 0.7477 | 6
2016-10-14 | 288 | 1716 | EB | clr | * | 5.9 | 0 | 19 | 0.59 | 0.746 | 6
2016-10-15 | 289 | 1300 | TG | clr | * | 5.8 | 0 | 19 | 0.59 | 0.7426 | 6
2016-10-16 | 290 | [11 empty] | Sunday ICESat- didn't make it to TAWO
2016-10-17 | 291 | 1431 | TG | rm frost | * | 5.8 | 100 | 19 | 0.6 | 0.7167 | 6
2016-10-18 | 292 | 1259 | EB | rm frost | * | 5.8 | 2100 | 19 | 0.6 | 0.706 | 6
2016-10-19 | 293 | 1257 | TG | rm lt frost | * | 5.8 | 2000 | 19 | 0.59 | 0.9703 | 7
2016-10-20 | 294 | [11 empty] | Con 1. Didn't go to TAWO.
2016-10-21 | 295 | 1619 | TG | rm lt frost | * | 5.8 | 100 | 19 | 0.6 | 0.8326 | 7
2016-10-22 | 296 | 1213 | EB | rm frost | * | 5.8 | 0 | 19 | 0.6 | 0.826 | 7
2016-10-23 | 297 | 1331 | TG | rm lt frost | * | 6.1 | 100 | 19 | 0.6 | 0.8166 | 7
2016-10-24 | 298 | 1544 | TG/MG/NB | rm frost | * | 6.1 | 100 | 19 | 0.58 | 0.812 | 7
2016-10-25 | 299 | 1133 | MG/TG | lt frost | * | 6.1 | 100 | 19 | 0.58 | 0.809 | 7 | * | Error in span check was 1.97, CO2 tank pressure is less than 100, flow rate was 4.2 lpm
2016-10-26 | 300 | 1213 | TG | rm frost | * | 5.8 | 100 | 19 | 0.59 | 0.8051 | 7 |  | 1253z started computer replacement, new computer up and running at 1307z. 1327z white filter change initiated. 1333z White filter change finished. 1343z changed clap flow from 0.54 to 0.60
2016-10-27 | 301 | 1128 | MG/NB | rm frost | * | 6.1 | 0 | 19 | 0.59 | 1.0006 | 2 |  | Computer found to be frozen, restarted at 1126z,
2016-10-28 | 302 | 1135 | MG/NB | rm frost | * | 6.2 | 0 | 19 | 0.6 | 0.822 | 2
2016-10-29 | 303 | 1150 | MG/NB | clr | * | 6.2 | 100 | 19 | 0.61 | 0.809 | 2
2016-10-30 | 304 | 1133 | NB | clr | * | 6.1 | 100 | 19 | 0.6 | 0.801 | 2
2016-10-31 | 305 | 1318 | NB | rm frost | * | 6.1 | 100 | 19 | 0.59 | 0.789 | 2 |  | Removed frost from sample inlet at 1315z.
2016-11-01 | 306 | 1253 | MG | rm lt frost | * | 6.1 | 0 | 19 | 0.6 | 0.778 | 2 |  | Removed frost from sample inlet at 1250z, sample inlet appears to be a little loose
2016-11-02 | 307 | 1251 | NB | clr | * | 6.1 | 100 | 19 | 0.6 | 0.772 | 2
2016-11-03 | 308 | 1245 | MG | rm frost | * | 6 | 100 | 19 | 0.6 | 0.698? | 2
2016-11-04 | 309 | 1309 | MG | rm frost | * | 6 | 100 | 19 | 0.6 | 0.694? | 2
2016-11-05 | 310 | 1230 | MG | rm lt frost | * | 6 | 0 | 19 | 0.6 | 0.691? | 2
2016-11-06 | 311 | 1325 | NB | clr | * | 5.8 | 0 | 20 | 0.6 | 0.7 | 2 |  | Found primary pump, Colgan, shut down upon arrival at the TAWO this morning. Moved sample lines over to spare pump, Dignan, and restarted sample flow around 1338z. Colgan shut down due to carbon vane failure. Replaced carbon vanes.
2016-11-07 | 312 | 1230 | MG | rm frost |  | 5.7 | 0 | 20 | 0.59 | 0.698 | 2
2016-11-08 | 313 | 1225 | MG | rm lt frost |  | 6.1 | 100 | 20 | 0.58 | 0.691 | 2
2016-11-09 | 314 | 1216 | MG | clr |  | 6.4 | 100 | 20 | 0.5 | 0.688 | 2
2016-11-10 | 315 | 1242 | MG | clr | * | 6.7 | 100 | 20 | 0.61 | 0.688 | 2 |  | Clap flag 0004, TrB<0.7
2016-11-11 | 316 |  |  |  | * | [7 empty] | Did not go to TAWO condition 1
2016-11-12 | 317 | 1240 | MG/NB | clr | * | 6.8 | 0 | 20 | 0.6 | 0.984 | 3
2016-11-13 | 318 | 1250 | MG | rm lt frost | * | 6.8 | 100 | 20 | 0.55 | 0.977 | 3
2016-11-14 | 319 | 1900 | NB | rm frost | * | 6.6 | 100 | 20 | 0.6 | 0.977 | 3
2016-11-15 | 320 | 1400 | MG | rm lt frost | * | 6.2 | 100 | 20 | 0.58 | 0.975 | 3
2016-11-16 | 321 | 1302 | MG | rm lt frost | * | 5.7 | 100 | 20 | 0.61 | 0.966 | 3 |  | Changed neph lamp bulb.
2016-11-17 | 322 | 1218 | MG/NB | rm lt frost | * | 5.7 | 100 | 20 | 0.62 | 0.965 | 3
2016-11-18 | 323 | 1323 | MG | rm frost | * | 5.8 | 0 | 20 | 0.62 | 0.947 | 3
2016-11-19 | 324 | 1235 | NB | rm lt frost | * | 5.8 | 100 | 20 | 0.61 | 0.946 | 3
2016-11-20 | 325 | 1312 | NB | rm frost | * | 5.7 | 100 | 20 | 0.61 | 0.945 | 3
2016-11-21 | 326 | 1305 | NB | rm frost | * | 5.8 | 100 | 20 | 0.61 | 0.885 | 3
2016-11-22 | 327 | 1301 | MG/NB | rm frost | * | 5.8 | 100 | 20 | 0.61 | 0.881 | 3
2016-11-23 | 328 | 1230 | NB | lt frost | * | 5.8 | 100 | 20 | 0.6 | 0.88 | 3
2016-11-24 | 329 | 1250 | NB | rm frost | * | 5.8 | 100 | 20 | 0.6 | 0.877 | 3
2016-11-25 | 330 | 1249 | MG | rm frost | * | 5.8 | 100 | 19 | 0.61 | 0.861 | 3
2016-11-26 | 331 | 1335 | NB | clr | * | 5.8 | 100 | 20 | 0.61 | 0.859 | 3
2016-11-27 | 332 | 1420 | MG | clr | * | 5.8 | 100 | 20 | 0.6 | 0.854 | 3
2016-11-28 | 333 | 1235 | NB | rm frost | * | 5.8 | 100 | 20 | 0.61 | 0.852 | 3 | * | Attached HEPA filter at 1232z for NEPH span check. Removed HEPA filter and started span check at 1345z. Aerosol computer froze during last filtered air flush of span check. Rebooted computer at 1407z. Did not attempt a second span check because of limited CO2 supply.
2016-11-29 | 334 | 1245 | NB | rm frost | * | 5.8 | 100 | 20 | 0.62 | 0.832 | 3 | * | Attached HEPA filter at 1245z for repeat of span check, sans CO2. Span check ended at 1435. Error: 61.1%
2016-11-30 | 335 | 1325 | MG | rm lt frost | * | 5.8 | 100 | 20 | 0.61 | 0.805 | 3
2016-12-01 | 336 | 1250 | NB | rm frost | * | 5.8 | 100 | 20 | 0.61 | 0.804 | 3
2016-12-02 | 337 | 1321 | MG | rm frost | *
2016-12-03 | 338 | 1235 | NB | rm frost | * | 5.8 | 100 | 20 | 0.6 | 0.788 | 3
2016-12-04 | 339 | 1348 | NB | rm frost | * | 5.8 | 100 | 20 | 0.61 | 0.785 | 3
2016-12-05 | 340 | 1242 | NB | clr | * | 5.8 | 100 | 20 | 0.61 | 0.777 | 3
2016-12-06 | 341 | 1244 | MG | rm frost | * | 5.8 | 0 | 20 | 0.61 | 0.769 | 3
2016-12-07 | 342 | 1240 | MG | clr | * | 5.7 | 0 | 20 | 0.62 | 0.765 | 3
2016-12-08 | 343 | 1234 | MG/NB | clr | * | 5.7 | 100 | 20 | 0.61 | 0.762 | 3
2016-12-09 | 344 | 1328 | MG | rm frost | * | 5.8 | 100 | 20 | 0.62 | 0.76 | 3
2016-12-10 | 345 | 1920 | MG | rm frost | * | 5.8 | 100 | 20 | 0.6 | 0.758 | 3
2016-12-11 | 346 | 1340 | MG | rm lt frost | * | 5.8 | 100 | 20 | 0.62 | 0.758 | 3
2016-12-12 | 347 | 1230 | NB | rm lt frost |  | 5.8 | 100 | 20 | 0.62 | 0.757 | 3 |  | Daily email missing.
2016-12-13 | 348 | 1225 | NB | rm lt frost | * | 5.8 | 100 | 20 | 0.61 | 0.75 | 3
2016-12-14 | 349 | 1304 | MG | rm lt frost | * | 5.7 | 100 | 20 | 0.63 | 0.748 | 3
2016-12-15 | 350 | 1415 | MG | rm lt frost | * | 5.8 | 100 | 20 | 0.63 | 0.757 | 3
2016-12-16 | 351 | 1308 | MG | rm lt frost | * | 5.7 | 100 | 20 | 0.61 | 0.746 | 3
2016-12-17 | 352 | 1252 | MG | rm lt frost | * | 5.8 | 100 | 20 | 0.59 | 0.745 | 3
2016-12-18 | 353 | 1328 | NB | clr | * | 5.7 | 100 | 20 | 0.61 | 0.742 | 3
2016-12-19 | 354 | 1220 | NB | clr | * | 5.7 | 100 | 20 | 0.59 | 0.739 | 3
2016-12-20 | 355 | 1506 | MG | rm lt frost | * | 5.7 | 100 | 20 | 0.59 | 0.894 | 4
2016-12-21 | 356 | 1325 | MG | rm lt frost | * | 5.7 | 100 | 19 | 0.55 | 0.689 | 4 |  | Lots of north wind, clap trasmitance makes sense
2016-12-22 | 357 | 1350 | MG | rm lt frost | * | 5.8 | 0 | 19 | 0.62 | 0.834 | 5
2016-12-23 | 358 | 1244 | NB | clr | * | 5.7 | 100 | 20 | 0.62 | 0.819 | 5 |  | Power outages for e-gen swap: 1330-1333z and 1645-1648z.
2016-12-24 | 359 | 1326 | NB | clr | * | 5.7 | 100 | 20 | 0.6 | 0.816 | 5
2016-12-25 | 360 | 1456 | MG | rm lt frost | * | 5.7 | 100 | 20 | 0.61 | 0.813 | 5
2016-12-26 | 361 | 1235 | NB | rm lt frost |  | 5.7 | 100 | 20 | 0.61 | 0.812 | 5 |  | Daily email missing. Attached HEPA filter 1235z-1337z. CO2 tank pressure 0 psi. Flow rate 4.2 lpm at 25psi. Span check error: 3.2%.
2016-12-27 | 362 | 1034 | MG | - |  | 5.7 | 100 | 20 | 0.59 | 0.809 | 5 |  | Did not go on the roof due to 32 knot winds
2016-12-28 | 363 | 1358 | MG | clr |  | 5.7 | 100 | 20 | 0.61 | 0.808 | 5
2016-12-29 | 364 | 1308 | MG/NB | clr |  | 5.7 | 100 | 20 | 0.61 | 0.798 | 5
2016-12-30 | 365 | 1314 | MG/NB | clr |  | 5.7 | 100 | 20 | 0.63 | 0.777 | 5
2016-12-31 | 366 | 1332 | NB | clr |  | 5.7 | 100 | 20 | 0.65 | 0.77 | 5 |  | Adjusted CLAP flow from 0.65 - 0.60 lpm.

## CU GC
Date | Julian Day | Time (UTC) | Observer's Initials | H2 Gen. Act Pressure (PSI) | H2 Generator Water Level | H2 Gen. Flow (ml/min) | Methane Std 1 Pressure | Methane Std 2 Pressure | Inlet Line Flow Rate (L/min) | 4 Pumps On? | NMHC Std 1 Pressure | NMHC Std 2 Pressure | Trap Housing Purge Flow (ml/min) | Graphs Look Good? | Daily H2 Gen Desiccant Check | Daily H2 Line Filter Desiccant Check | Daily Oxygen indicator (green) | Comments (Inlet Cleaning, Inlet Filter Changes, Water Refill, etc.)
2016-01-01 | 1 | 1510 | JG | off | off | off | off | 960 | 2 | Y | 560 | 1110 | 90 | Y |  |  |  | H2 tank pres 2300 psi; line pres 49 psi.
2016-01-02 | 2 | 1229 | NB | off | off | off | off | 960 | 2 | Y | 560 | 1110 | 90 | Y |  |  |  | H2 tank pres 2300 psi; line pres 49 psi. Cleaned frost from inlet 1410z.
2016-01-03 | 3 | 1310 | JG | off | off | off | off | 950 | 2 | Y | 560 | 1110 | 90 | Y |  |  |  | H2 tank pres 2250 psi; line pres 50 psi.
2016-01-04 | 4 | 1228 | NB | off | off | off | off | 950 | 2 | Y | 560 | 1110 | 90 | Y |  | * |  | H2 tank pres 2250 psi; line pres 50 psi. JG climbed tower 1355-1417z and removed frost from inlet.
2016-01-05 | 5 | 1230 | JG | off | off | off | off | 950 | 2 | Y | 560 | 1110 | 95 | Y |  |  |  | H2 tank pres 2250 psi; line pres 50 psi.
2016-01-06 | 6 | 1307 | NB | off | off | off | off | 940 | 2 | Y | 560 | 1110 | 95 | Y |  |  |  | H2 tank pres 2240 psi; line pres 49 psi.
2016-01-07 | 7 | 1235 | JG | off | off | off | off | 930 | 2 | Y | 550 | 1110 | 95 | Y |  |  |  | H2 tank pres 2200 psi; line pres 50 psi.
2016-01-08 | 8 | 1347 | NB | off | off | off | off | 930 | 2 | Y | 550 | 1110 | 95 | Y |  |  |  | H2 tank pres 2150 psi; line pres 51 psi.
2016-01-09 | 9 | 1240 | JG | off | off | off | off | 930 | 2 | Y | 550 | 1110 | 95 | Y |  |  |  | H2 tank pres 2070 psi; line pres 52 psi.
2016-01-10 | 10 | 1504 | NB | off | off | off | off | 920 | 2 | Y | 550 | 1110 | 95 | Y |  |  |  | H2 tank pres 2050 psi; line pres 51 psi. JG climbed tower 1513-1530z and removed frost from inlet.
2016-01-11 | 11 | 1240 | JG | off | off | off | off | 910 | 2 | Y | 550 | 1110 | 95 | Y |  | * |  | H2 tank pres 2070 psi; line pres 52 psi.
2016-01-12 | 12 | 1232 | NB | off | off | off | off | 910 | 2 | Y | 550 | 1110 | 95 | Y |  |  |  | H2 tank pres 2060 psi; line pres 50 psi. JG climbed tower 1243-1307z and removed frost from inlet. Installed new H2 generator at 1615z. H2 tank turned off - tank pres 2060 psi.
2016-01-13 | 13 | 1255 | JG | 60 | * | 52 | off | 900 | 2 | Y | 550 | 1110 | 98 | Y |  |  |  | Planned power outage for e-gen test 1307-1309 and 1321-1324 UTC.
2016-01-14 | 14 | 1225 | NB | 60 | added 1L | 53 | off | 900 | 2 | Y | 550 | 1110 | 97 | Y
2016-01-15 | 15 | 1255 | JG | 60 | * | 54 | off | 900 | 2 | Y | 550 | 1110 | 98 | Y |  |  |  | JG climbed tower 1327-1347 and cleaned frost from inlet.
2016-01-16 | 16 | 1304 | NB | 60 | * | 52 | off | 890 | 2 | Y | 550 | 1110 | 98 | Y
2016-01-17 | 17 | 1725 | JG | 60 | * | 55 | off | 880 | 2 | Y | 550 | 1110 | 98 | Y
2016-01-18 | 18 | 1229 | NB | 60 | * | 53 | off | 880 | 2 | Y | 550 | 1110 | 97 | Y | * | *
2016-01-19 | 19 | 1310 | JG | 60 | * | 53 | off | 880 | 2 | Y | 550 | 1110 | 98 | Y
2016-01-20 | 20 | 1232 | NB | 60 | * | 55 | off | 880 | 2 | Y | 550 | 1110 | 98 | Y
2016-01-21 | 21 | 1250 | JG | 60 | * | 54 | off | 870 | 2 | Y | 550 | 1110 | 98 | Y
2016-01-22 | 22 | 1234 | NB | 60 | * | 53 | off | 870 | 2 | Y | 550 | 1110 | 98 | Y
2016-01-23 | 23 | 1210 | JG | 60 | * | 54 | off | 860 | 2 | Y | 550 | 1110 | 98 | Y |  |  |  | NB climbed tower and removed frost from inlets 1217-1234z.
2016-01-24 | 24 | 1258 | NB | 60 | * | 53 | off | 860 | 2 | Y | 550 | 1100 | 98 | Y
2016-01-25 | 25 | 1230 | JG | 60 | * | 53 | off | 860 | 2 | Y | 550 | 1100 | 98 | Y | * | *
2016-01-26 | 26 | 1254 | NB | 60 | * | 54 | off | 850 | 2 | Y | 550 | 1100 | 98 | Y
2016-01-27 | 27 | 1340 | JG | 60 | * | 54 | off | 840 | 2 | Y | 550 | 1100 | 98 | Y
2016-01-28 | 28 | 1234 | NB | 60 | * | 54 | off | 840 | 2 | Y | 550 | 1100 | 98 | Y
2016-01-29 | 29 | 1255 | JG | 60 | added 1L | 54 | off | 840 | 2 | Y | 550 | 1100 | 98 | Y
2016-01-30 | 30 | 1225 | NB | 60 | * | 54 | off | 840 | 2 | Y | 550 | 1100 | 98 | Y
2016-01-31 | 31 | 1455 | JG | 60 | * | 55 | off | 830 | 2 | Y | 550 | 1100 | 98 | Y
2016-02-01 | 32 | 1225 | NB | 60 | * | 54 | off | 820 | 2 | Y | 550 | 1100 | 98 | Y | * | * |  | JG climbed tower 1310-1327z and cleaned frost from inlet.
2016-02-02 | 33 | 1240 | JG | 60 | * | 55 | off | 810 | 2 | Y | 550 | 1100 | 100 | Y
2016-02-03 | 34 | 1248 | NB | 60 | * | 56 | off | 810 | 2 | Y | 550 | 1100 | 100 | Y
2016-02-04 | 35 | 1220 | JG | 60 | * | 55 | off | 810 | 2 | Y | 550 | 1100 | 100 | Y |  |  |  | NB climbed tower 1224-1239z and removed frost from inlet.
2016-02-05 | 36 | 1333 | NB | 60 | * | 54 | off | 810 | 2 | Y | 550 | 1100 | 100 | Y |  |  |  | JG climbed tower 1342-1359z and removed frost from inlet.
2016-02-06 | 37 | 1355 | JG | 60 | * | 54 | off | 800 | 2 | Y | 550 | 1100 | 100 | Y |  |  |  | NB climbed tower 1407-1421z and removed frost from inlet.
2016-02-07 | 38 | 1455 | MB/JJ | 60 | * | 54 | off | 800 | 2 | y | 530 | 1100 | 100 | Y
2016-02-08 | 39 | 1240 | MB/JJ/SD | 60 | * | 55 | off | 800 | 2 | Y | 530 | 1100 | 100 | Y | * | *
2016-02-09 | 40 | 1230 | MB/JJ/NB | 60 | * | 55 | off | 790 | 2 | Y | 530 | 1100 | 100 | Y
2016-02-10 | 41 | 1226 | SWD/JG/NB | 60 | * | 54 | off | 780 | 2 | Y | 530 | 1100 | 100 | Y
2016-02-11 | 42 | 1340 | MB/JJ | 60 | * | 54 | off | 770 | 2 | Y | 530 | 1100 | 100 | Y
2016-02-12 | 43 | 1625 | MB/JJ | 60 | * | 55 | off | 770 | 2 | Y | 530 | 1080 | 100 | Y
2016-02-13 | 44 | 1220 | MB/JJ | 60 | * | 55 | off | 770 | 2 | Y | 530 | 1080 | 100 | Y
2016-02-14 | 45 | 1700 | MB/JJ | 60 | added 1L | 55 | off | 760 | 2 | y | 530 | 1080 | 100 | Y |  |  |  | Added 1L DI water to H2 generator.
2016-02-15 | 46 | 1641 | JJ | 60 | * | 54 | off | 760 | 2 | Y | 525 | 1070 | 100 | Y | * | *
2016-02-16 | 47 | 1321 | jj/mb | 60 | * | 54 | off | 760 | 2 | y | 530 | 1080 | 100 | y
2016-02-17 | 48 | 1415 | MB/JJ | 60 | * | 53 | off | 750 | 2 | Y | 530 | 1080 | 100 | Y
2016-02-18 | 49 | 1405 | MB/JJ | 60 | * | 53 | off | 740 | 2 | Y | 520 | 1080 | 100 | Y
2016-02-19 | 50 | 1349 | JJ/MB | 60 | * | 55 | off | 730 | 2 | Y | 510 | 1080 | 85 | Y
2016-02-20 | 51 | 1311 | JJ/MB | 60 | * | 54 | off | 730 | 2 | Y | 530 | 1080 | 100 | Y |  |  |  | JJ climbed tower 1332-1358z and 1457-1526z to troubleshoot 2m RTD cable.
2016-02-21 | 52 | 1510 | JJ/MB | 60 | * | 55 | off | 720 | 2 | Y | 520 | 1080 | 100 | Y
2016-02-22 | 53 | 1244 | JJ/MB | 60 | * | 55 | off | 720 | 2 | Y | 510 | 1080 | 100 | Y | * | *
2016-02-23 | 54 | 1704 | MB | 60 | * | 55 | off | 730 | 2 | Y | 520 | 1080 | 100 | Y
2016-02-24 | 55 | 1314 | JJ/MB | 60 | * | 54 | off | 710 | 2 | Y | 510 | 1070 | 100 | Y
2016-02-25 | 56 | 1336 | jj/mb | 60 | * | 55 | off | 710 | 2 | Y | 520 | 1070 | 100 | Y
2016-02-26 | 57 | 1348 | JJ/MB | 60 | * | 55 | off | 710 | 2 | Y | 510 | 1070 | 100 | Y
2016-02-27 | 58 | 1305 | MB | 60 | * | 55 | off | 700 | 2 | Y | 520 | 1080 | 100 | Y |  |  |  | SWD climbed tower to clean instruments and troubleshoot 2m-RTD cable from 1444-1542z.
2016-02-28 | 59 | 1708 | JJ | 60 | * | 55 | off | 700 | 2 | y | 510 | 1070 | 100 | Y
2016-02-29 | 60 | 1339 | jj/mb | 60 | * | 57 | off | 700 | 2 | y | 510 | 1070 | 100 | Y | * | *
2016-03-01 | 61 | 1248 | JJ/MB | 60 | * | 55 | off | 680 | 2 | Y | 510 | 1060 | 100 | Y
2016-03-02 | 62 | 1711 | MB/JJ | 60 | added 1L | 56 | off | 680 | 2 | Y | 510 | 1060 | 100 | Y |  |  |  | Added 1L DI water to H2 generator. JJ climbed tower to clean instruments from 1737-1833z.
2016-03-03 | 63 | 1338 | MB/JJ | 60 | * | 55 | off | 670 | 2 | Y | 500 | 1060 | 100 | Y
2016-03-04 | 64 | 1309 | jj/mb | 60 | * | 55 | off | 660 | 2 | Y | 510 | 1060 | 100 | Y
2016-03-05 | 65 | 1235 | jj/mb | 60 | * | 56 | off | 660 | 2 | Y | 510 | 1070 | 100 | Y
2016-03-06 | 66 | 1342 | MB | 60 | * | 55 | off | 660 | 2 | Y | 510 | 1070 | 100 | Y
2016-03-07 | 67 | 1321 | jj/mb | 60 | * | 56 | off | 660 | 2 | Y | 510 | 1060 | 100 | Y | * | *
2016-03-08 | 68 | 1331 | jj/mb | 60 | * | 56 | off | 650 | 2 | Y | 510 | 1060 | 100 | Y
2016-03-09 | 69 | 1235 | JJ | 60 | * | 55 | off | 650 | 2 | y | 510 | 1070 | 100 | Y
2016-03-10 | 70 | 0 | 0 | [14 empty] | White out conditions (Condition 1) did not travel to TAWO this day.
2016-03-11 | 71 | 1352 | JJ | 60 | * | 55 | off | 630 | 2 | y | 510 | 1060 | 100 | Y
2016-03-12 | 72 | 1437 | MB | 60 | * | 55 | off | 620 | 2 | Y | 510 | 1060 | 100 | Y
2016-03-13 | 73 | 1346 | JJ | 60 | * | 55 | off | 620 | 2 | y | 510 | 1060 | 100 | Y |  |  |  | unexpected power outage today from 13:16 to 13:23 UTC
2016-03-14 | 74 | 1353 | jj/mb | 60 | * | 55 | off | 620 | 2 | y | 510 | 1070 | 100 | Y | * | *
2016-03-15 | 75 | 1533 | jj/mb | 60 | * | 57 | off | 610 | 2 | y | 500 | 1060 | 100 | Y
2016-03-16 | 76 | 1319 | MB | 60 | * | 57 | off | 600 | 2 | Y | 500 | 1050 | 100 | Y
2016-03-17 | 77 | 1338 | jj/mb | 60 | * | 57 | off | 600 | 2 | Y | 500 | 1060 | 100 | Y
2016-03-18 | 78 | 1305 | MB/JJ | 60 | * | 56 | off | 590 | 2 | Y | 500 | 1060 | 100 | Y
2016-03-19 | 79 | 1854 | MB/JJ | 60 | * | 55 | off | 590 | 2 | Y | 500 | 1060 | 100 | Y
2016-03-20 | 80 | 1334 | MB | 60 | * | 57 | off | 580 | 2 | Y | 500 | 1060 | 100 | Y
2016-03-21 | 81 | 1307 | jj/mb | 60 | added 1L | 55 | off | 580 | 2 | Y | 500 | 1060 | 100 | Y | * | *
2016-03-22 | 82 | 1320 | MB | 60 | * | 56 | off | 570 | 2 | Y | 500 | 1060 | 100 | Y
2016-03-23 | 83 | 1744 | MB/JJ | 60 | * | 55 | off | 560 | 2 | Y | 500 | 1060 | 100 | Y
2016-03-24 | 84 | 1329 | jj/mb | 60 | * | 56 | off | 570 | 2 | Y | 500 | 1060 | 100 | Y
2016-03-25 | 85 | 1303 | JJ | 60 | * | 56 | off | 560 | 2 | y | 500 | 1050 | 95 | Y
2016-03-26 | 86 | 1303 | MB | 60 | * | 56 | off | 550 | 2 | Y | 500 | 1050 | 95 | Y
2016-03-27 | 87 | 1715 | JJ | 60 | * | 55 | off | 550 | 2 | Y | 500 | 1060 | 95 | Y
2016-03-28 | 88 | 1327 | MB | 60 | * | 55 | off | 540 | 2 | Y | 500 | 1050 | 100 | Y | * | *
2016-03-29 | 89 | 1151 | JJ | 60 | * | 56 | off | 540 | 2 | Y | 500 | 1050 | 85 | Y
2016-03-30 | 90 | 1134 | MB | 60 | * | 55 | off | 540 | 2 | Y | 500 | 1050 | 100 | Y
2016-03-31 | 91 | 1718 | jj/mb | 60 | * | 55 | off | 540 | 2 | Y | 500 | 1050 | 95 | Y
2016-04-01 | 92 | 1304 | JJ | 60 | * | 56 | off | 530 | 2 | Y | 500 | 1050 | 95 | Y
2016-04-02 | 93 | 1213 | jj/mb | 60 | * | 53 | off | 520 | 2 | Y | 500 | 1050 | 100 | Y
2016-04-03 | 94 | 1550 | MB | 60 | * | 55 | off | 520 | 2 | Y | 500 | 1050 | 100 | Y
2016-04-04 | 95 | 1613 | jj/mb | 60 | * | 55 | off | 510 | 2 | Y | 500 | 1050 | 100 | Y | * | *
2016-04-05 | 96 | 1251 | MB/JJ | 60 | * | 56 | off | 500 | 2 | Y | 500 | 1050 | 100 | Y
2016-04-06 | 97 | 1332 | JJ | 60 | added 1L | 56 | off | 500 | 2 | Y | 500 | 1050 | 100 | Y |  |  |  | Planned power outage today
2016-04-07 | 98 | 1224 | MB/JJ | 60 | * | 56 | off | 500 | 2 | Y | 500 | 1050 | 100 | Y
2016-04-08 | 99 | 1139 | JJ/MB | 60 | * | 56 | off | 500 | 2 | Y | 500 | 1050 | 100 | Y |  |  |  | SWD climbed the tower from 1715-1800z to measure 10m and 16m windbird alignment. SWD reattached turnbuckle to strength cable from CUGC exhaust line to tower.
2016-04-09 | 100 | 1254 | JJ/MB | 60 | * | 57 | off | 480 | 2 | Y | 500 | 1050 | 100 | Y
2016-04-10 | 101 | 1707 | JJ | 60 | * | 56 | off | 480 | 2 | Y | 490 | 1050 | 100 | Y
2016-04-11 | 102 | 1321 | jj | 60 | * | 56 | off | 470 | 2 | Y | 500 | 1050 | 100 | Y | * | *
2016-04-12 | 103 | 1554 | JJ/MB | 60 | * | 55 | off | 470 | 2 | Y | 500 | 1050 | 100 | Y
2016-04-13 | 104 | 1317 | JJ/MB | 60 | * | 57 | off | 460 | 2 | Y | 500 | 1040 | 100 | Y
2016-04-14 | 105 | 1329 | JJ/MB | 60 | * | 57 | off | 460 | 2 | Y | 490 | 1040 | 100 | Y |  |  |  | JJ climbed tower from 1356-1422z to work on surface ozone inlet and clean instruments.
2016-04-15 | 106 | 1711 | JJ/MB/SWD | 60 | * | 56 | off | 450 | 2 | Y | 480 | 1040 | 100 | Y
2016-04-16 | 107 | 1309 | JJ/MB | 60 | * | 55 | off | 450 | 2 | Y | 480 | 1040 | 100 | Y
2016-04-17 | 108 | 1628 | MB | 60 | * | 56 | off | 440 | 2 | Y | 480 | 1040 | 100 | Y
2016-04-18 | 109 | 1249 | MB | 60 | * | 57 | off | 440 | 2 | Y | 480 | 1040 | 100 | Y | * | *
2016-04-19 | 110 | 1152 | JJ/MB | 60 | * | 57 | off | 430 | 2 | Y | 500 | 1040 | 100 | Y
2016-04-20 | 111 | 1143 | JJ/MB | 60 | * | 56 | off | 440 | 2 | Y | 480 | 1030 | 100 | Y
2016-04-21 | 112 | 1122 | JJ/MB/SWD | 60 | * | 56 | off | 410 | 2 | Y | 480 | 1040 | 100 | Y |  |  |  | 1426-1458z and 1513-1527z SWD climbed tower to clean instruments and realign 10m windbird.
2016-04-22 | 113 | 1838 | JJ/MB | 60 | * | 56 | off | 400 | 2 | Y | 480 | 1030 | 100 | Y
2016-04-23 | 114 | 1536 | 0 | [14 empty] | White out conditions (Condition 1) did not travel to TAWO this day.
2016-04-24 | 115 | 1536 | JJ/MB | 60 | added 1L | 57 | off | 400 | 2 | Y | 480 | 1030 | 100 | Y
2016-04-25 | 116 | 1207 | JJ/MB | 60 | * | 56 | off | 390 | 2 | Y | 480 | 1030 | 100 | Y | * | *
2016-04-26 | 117 | 1130 | JJ | 60 | * | 57 | off | 390 | 2 | Y | 480 | 1030 | 100 | Y
2016-04-27 | 118 | 1549 | JJ/MB | 60 | * | 74 | 1650 | 390 | 2 | Y | 490 | 1030 | 130 | High H2 flow |  |  |  | Detlev working on the GC.
2016-04-28 | 119 | 1315 | MB | 60 | * | 55 | 1580 | 380 | 2 | Y | 480 | 1030 | 100 | Y |  |  |  | Detlev working on the GC. Trap housing >250, reset to 100 ml/min. 1255-1305z JJ climbed the tower to clean instruments.
2016-04-29 | 120 | 1300 | MB/JJ | 60 | * | 55 | 1450 | 370 | 2 | Y | 480 | 1040 | 110 | Y
2016-04-30 | 121 | 1336 | MB/JJ | 60 | * | 53 | 1530 | 370 | 2 | Y | 470 | 1050 | 110 | Y |  |  |  | Detlev working with tanks.
2016-05-01 | 122 | 1624 | MB/JJ | 60 | * | 50 | 1620 | 370 | 2 | Y | 470 | 480 | 100 | switched tanks |  |  |  | 1743-1815z JJ and MB climbed tower to replace filter on CUGC sample inlet.
2016-05-02 | 123 | 1322 | JJ/MB | 60 | * | 50 | 1540 | 360 | 2 | Y | 450 | 480 | 105 | Y | * | * |  | Planned power outages starting from 1635 to 1700z
2016-05-03 | 124 | 1230 | MB/JJ | 60 | * | 65 | 1650 | 350 | 2 | Y | 490 | 640 | 110 | High H2 flow; switched tanks |  |  |  | 1153-1217z JJ climbed tower to clean instruments and attach ties to CUGC intake hose.
2016-05-04 | 125 | 1451 | MB/JJ | 60 | * | 64 | 1650 | 350 | 2 | Y | 490 | 1040 | 110 | Y |  |  |  | Adjusted trap housing purge to 100 ml/min.
2016-05-05 | 126 | 1528 | MB | 60 | * | 62 | 1630 | 350 | 2 | Y | 480 | 1040 | 105 | Y |  |  |  | Adjusted trap housing purge to 100 ml/min.
2016-05-06 | 127 | 1845 | JJ/MB | 60 | * | 65 | 1620 | 350 | 2 | Y | 480 | 1040 | 110 | Y |  |  |  | Adjusted trap housing purge to 100 ml/min.
2016-05-07 | 128 | 1315 | JJ/MB | 60 | * | 65 | 1620 | 350 | 2 | Y | 480 | 1030 | 90 | Y |  |  |  | Adjusted trap housing purge to 100 ml/min.
2016-05-08 | 129 | 1558 | JJ | 60 | * | 62 | 1620 | 350 | 2 | Y | 470 | 1040 | 90 | Y |  |  |  | Adjusted trap housing purge to 100 ml/min.
2016-05-09 | 130 | 1635 | JJ/MB | 60 | * | 65 | 1640 | 350 | 2 | Y | 470 | 1050 | 100 | Y | * | *
2016-05-10 | 131 | 1648 | JJ/MB | 60 | * | 66 | 1640 | 350 | 2 | Y | 470 | 1030 | 100 | Y
2016-05-11 | 132 | 1753 | MB/JJ | 60 | * | 65 | 1610 | 340 | 2 | Y | 470 | 1030 | 100 | Y
2016-05-12 | 133 | 1622 | MB/JJ | 60 | added 1L | 65 | 1620 | 340 | 2 | Y | 480 | 1030 | 100 | Y
2016-05-13 | 134 | 1308 | MB/JJ | 60 | * | 64 | 1620 | 340 | 2 | Y | 480 | 1030 | 100 | Y
2016-05-14 | 135 | 1149 | JJ/MB | 60 | * | 64 | 1620 | 340 | 2 | Y | 480 | 1030 | 100 | Y
2016-05-15 | 136 | 1730 | MB | 60 | * | 65 | 1620 | 340 | 2 | Y | 480 | 1030 | 100 | Y
2016-05-16 | 137 | 1609 | JJ/MB | 60 | * | 65 | 1620 | 340 | 2 | Y | 480 | 1030 | 100 | Y | * | *
2016-05-17 | 138 | 1722 | JJ/MB | 60 | * | 66 | 1610 | 330 | 2 | Y | 470 | 1030 | 100 | Y
2016-05-18 | 139 | 1309 | JJ/MB | 60 | * | 64 | 1620 | 330 | 2 | Y | 470 | 1030 | 100 | Y
2016-05-19 | 140 | 1322 | JJ | 60 | * | 64 | 1610 | 330 | 2 | Y | 470 | 1030 | 95 | Y |  |  |  | Adjusted trap housing purge to 100 ml/min
2016-05-20 | 141 | 1850 | JJ/MB | 60 | * | 65 | 1620 | 330 | 2 | Y | 470 | 1030 | 95 | Y
2016-05-21 | 142 | 1228 | JJ/MB | 60 | * | 65 | 1610 | 330 | 2 | Y | 470 | 1030 | 95 | Y
2016-05-22 | 143 | 2327 | JJ | 60 | * | 66 | 1620 | 330 | 2 | Y | 460 | 1030 | 95 | Y
2016-05-23 | 144 | 1656 | JJ | 60 | * | 65 | 1620 | 330 | 2 | Y | 460 | 1030 | 95 | Y | * | *
2016-05-24 | 145 | 1211 | JJ/MB | 60 | * | 65 | 1620 | 320 | 2 | Y | 470 | 1030 | 95 | Y
2016-05-25 | 146 | 1726 | JJ | 60 | * | 62 | 1620 | 330 | 2 | Y | 450 | 1030 | 95 | Y
2016-05-26 | 147 | 1232 | JJ/MB | 60 | * | 65 | 1610 | 320 | 2 | Y | 460 | 1030 | 90 | Y |  |  |  | Adjusted trap housing purge to 100 ml/min
2016-05-27 | 148 | 1845 | JJ/mb | 60 | * | 65 | 1610 | 320 | 2 | Y | 460 | 1030 | 100 | Y
2016-05-28 | 149 | 1221 | MB | 60 | added 1L | 65 | 1610 | 310 | 2 | Y | 460 | 1030 | 100 | Y |  |  |  | Added 1L of DI water.
2016-05-29 | 150 | 1530 | MB | 60 | * | 65 | 1610 | 310 | 2 | Y | 460 | 1030 | 100 | Y
2016-05-30 | 151 | 1226 | JJ | 60 | * | 65 | 1630 | 320 | 2 | Y | 460 | 1030 | 100 | Y | * | * |  | Planned power outage today from 2130 to 2230
2016-05-31 | 152 | 1731 | JJ | 60 | * | 65 | 1620 | 320 | 2 | Y | 450 | 1030 | 100 | Y
2016-06-01 | 153 | 1840 | JJ | 60 | * | 66 | 1620 | 320 | 2 | Y | 450 | 1020 | 105 | Y
2016-06-02 | 154 | 1831 | MB/JJ | 60 | * | 65 | 1620 | 310 | 2 | Y | 460 | 1030 | 105 | Y
2016-06-03 | 155 | 1737 | MB/JJ/HJ | 60 | * | 54 | 1660 | 310 | 2 | Y | 460 | 1030 | 105 | Y |  |  |  | H2 line filter desiccant pale yellow. Turned off H2 generator and replaced with H2 backup tank. Switched valve to receive H2 tank at 2012z. H2 tank pressure 2100 psi.
2016-06-04 | 156 | 1156 | MB | offline | offline | offline | 1610 | 300 | 2 | Y | 460 | 1030 | 95 | Y |  |  |  | H2 tank pressure 2060psi.
2016-06-05 | 157 | 1403 | JJ/TG/EB | offline | offline | offline | 1620 | 310 | 2 | Y | 450 | 1020 | 100 | Y
2016-06-06 | 158 | 1613 | MB/JJ/TG/EB | offline | offline | offline | 1620 | 300 | 2 | Y | 450 | 1040 | 105 | Y | * | * |  | H2 Line filter desiccant very pale/clear yellow; H2 tank pressure 2100 psi.
2016-06-07 | 159 | 1536 | TG/EB/JJ | offline | offline | offline | 1610 | 300 | 2 | Y | 450 | 1020 | 104 | Y |  |  |  | H2 tank pressure 2060psi.
2016-06-08 | 160 | 1309 | TG/EB/JJ/MB | offline | offline | offline | 1620 | 300 | 2 | Y | 450 | 1000 | 105 | Y |  |  |  | H2 tank pressure 2000psi.
2016-06-09 | 161 | 1553 | TG/EB | offline | offline | offline | 1630 | 300 | 2 | Y | 420 | 1020 | 105 | Y |  |  |  | H2 tank pressure 2000psi.
2016-06-10 | 162 | 1239 | TG/EB | offline | offline | offline | 1630 | 300 | 2 | Y | 450 | 1030 | 105 | Y |  |  |  | H2 tank pressure 1940psi. Work on ventilation duct from 1105 to 1330, higher room temps.
2016-06-11 | 163 | 1203 | TG/EB | offline | offline | offline | 1620 | 300 | 2 | Y | 440 | 1040 | 105 | Y |  |  |  | H2 tank pressure 1950psi.
2016-06-12 | 164 | 1509 | TG/EB | offline | offline | offline | 1630 | 300 | 2 | Y | 450 | 1040 | 105 | Y |  |  |  | H2 tank pressure 1900psi
2016-06-13 | 165 | 1637 | TG/EB | offline | offline | offline | 1620 | 300 | 2 | Y | 450 | 1020 | 105 | Y | * | * |  | H2 Line filter desiccant very pale/clear yellow; H2 tank pressure 1900 psi.
2016-06-14 | 166 | 1153 | TG/EB | offline | offline | offline | 1620 | 300 | 2 | Y | 450 | 1030 | 103 | Y |  |  |  | H2 tank pressure 1900psi
2016-06-15 | 167 | 1202 | TG/EB | offline | offline | offline | 1610 | 300 | 2 | Y | 440 | 1020 | 105 | Y |  |  |  | H2 tank pressure 1895psi
2016-06-16 | 168 | 1156 | TG | offline | offline | offline | 1610 | 295 | 2 | Y | 430 | 1010 | 105 | Y |  |  |  | H2 tank pressure 1870psi
2016-06-17 | 169 | 1314 | TG | offline | offline | offline | 1620 | 295 | 2 | Y | 420 | 1010 | 105 | Y |  |  |  | H2 tank pressure 1860psi
2016-06-18 | 170 | 1241 | TG/EB | offline | offline | offline | 1625 | 290 | 2 | Y | 450 | 1020 | 105 | Y |  |  |  | H2 tank pressure 1825psi
2016-06-19 | 171 | 1808 | EB | offline | offline | offline | 1620 | 300 | 2 | Y | 450 | 1020 | 105 | Y |  |  |  | H2 tank pressure 1795psi
2016-06-20 | 172 | 1812 | TG/EB | offline | offline | offline | 1625 | 290 | 2 | Y | 450 | 1010 | 105 | Y | * | * |  | H2 tank pressure 1740psi. New H2 Gen installed w/o input to GC; new filters installed.
2016-06-21 | 173 | 1204 | TG/EB | offline | offline | offline | 1620 | 290 | 2 | Y | 440 | 1010 | 105 | Y |  |  |  | H2 tank pressure 1750psi.
2016-06-22 | 174 | 1332 | TG/EB | offline | offline | offline | 1620 | 300 | 2 | Y | 450 | 1020 | 105 | Y |  |  |  | H2 tank pressure 1700psi.
2016-06-23 | 175 | 1313 | EB | offline | offline | offline | 1520 | 290 | 2 | Y | 450 | 1025 | 105 | Y |  |  |  | H2 tank pressure 1680psi.
2016-06-24 | 176 | 1329 | TG | offline | offline | offline | 1620 | 280 | 2 | Y | 440 | 1010 | 105 | Y |  |  |  | H2 tank pressure 1640psi
2016-06-25 | 177 | 1517 | EB/TG | offline | offline | offline | 1640 | 290 | 2 | Y | 440 | 1010 | 105 | Y |  |  |  | H2 tank pressure 1650psi
2016-06-26 | 178 | 1828 | TG | offline | offline | offline | 1610 | 270 | 2 | Y | 440 | 1010 | 105 | Y |  |  |  | H2 tank pressure 1650psi
2016-06-27 | 179 | 1603 | EB/TG | offline | offline | offline | 1610 | 290 | 2 | Y | 440 | 1000 | 105 | Y |  |  |  | H2 tank pressure 1610psi.
2016-06-28 | 180 | 1452 | EB/TG | offline | offline | offline | 1610 | 290 | 2 | Y | 450 | 1010 | 105 | Y |  |  |  | H2 tank pressure 1600 psi.
2016-06-29 | 181 | 1135 | EB/TG | offline | offline | offline | 1610 | 270 | 2 | Y | 420 | 1000 | 105 | Y |  |  |  | H2 tank pressure 1590psi.
2016-06-30 | 182 | 1143 | TG/EB | offline | offline | offline | 1620 | 275 | 2 | Y | 450 | 1020 | 105 | Y |  |  |  | H2 tank pressure 1550psi
2016-07-01 | 183 | 1202 | TG/EB | offline | offline | offline | 1620 | 260 | 2 | Y | 430 | 1010 | 105 | Y |  |  |  | H2 tank pressure 1510psi
2016-07-02 | 184 | 1239 | TG/EB | offline | offline | offline | 1610 | 260 | 2 | Y | 420 | 1000 | 105 | Y |  |  |  | H2 tank pressure 1380psi
2016-07-03 | 185 | 1835 | EB | offline | offline | offline | 1620 | 260 | 2 | Y | 450 | 1000 | 105 | Y |  |  |  | H2 tank pressure 1410psi
2016-07-04 | 186 | 1029 | TG | offline | offline | offline | 1620 | 260 | 2 | y | 420 | 1010 | 105 | Y |  |  |  | H2 tank pressure 1410psi
2016-07-05 | 187 | 1207 | TG/EB | offline | offline | offline | 1620 | 260 | 2 | Y | 450 | 1000 | 105 | Y |  |  |  | H2 tank pressure 1380psi
2016-07-06 | 188 | 0.47153 | TG | offline | offline | offline | 1620 | 260 | 2 | Y | 420 | 1000 | 105 | Y |  |  |  | H2 tank pressure 1390psi
2016-07-07 | 189 | 1214 | TG/EB | offline | offline | offline | 1620 | 260 | 2 | Y | 420 | 1000 | 105 | Y |  |  |  | H2 tank pressure 1360psi
2016-07-08 | 190 | 1604 | TG/EB | offline | offline | offline | 1620 | 260 | 2 | Y | 450 | 1000 | 105 | Y |  |  |  | H2 tank pressure 1290psi
2016-07-09 | 191 | 1342 | TG | offline | offline | offline | 1620 | 260 | 2 | Y | 420 | 1000 | 105 | Y |  |  |  | H2 tank pressure 1260psi
2016-07-10 | 192 | 1319 | TG | offline | offline | offline | 1620 | 250 | 2 | Y | 420 | 1000 | 105 | Y |  |  |  | H2 tank pressure 1250psi
2016-07-11 | 193 | 1323 | EB | offline | offline | offline | 1620 | 250 | 2 | Y | 450 | 1010 | 105 | Y |  |  |  | H2 tank pressure 1210psi
2016-07-12 | 194 | 1228 | TG/EB | offline | offline | offline | 1620 | 250 | 2 | Y | 420 | 1000 | 105 | Y |  |  |  | H2 tank pressure 1220psi
2016-07-13 | 195 | 1137 | TG | offline | offline | offline | 1620 | 250 | 2 | Y | 420 | 1000 | 105 | Y |  |  |  | H2 tank pressure 1200psi
2016-07-14 | 196 | 1147 | TG/EB | offline | offline | offline | 1610 | 250 | 2 | Y | 440 | 1000 | 105 | Y |  |  |  | H2 tank pressure 1160psi
2016-07-15 | 197 | 1235 | TG/EB | offline | offline | offline | 1620 | 250 | 2 | Y | 410 | 1000 | 105 | Y |  |  |  | H2 tank pressure 1150psi
2016-07-16 | 198 | 1301 | TG/EB | offline | offline | offline | 1610 | 245 | 2 | Y | 440 | 1000 | 105 | Y |  |  |  | H2 tank pressure 1110psi
2016-07-17 | 199 | 1627 | EB | offline | offline | offline | 1610 | 240 | 2 | Y | 440 | 1000 | 105 | Y |  |  |  | H2 tank pressure 1070psi
2016-07-18 | 200 | 1207 | TG | offline | offline | offline | 1610 | 240 | 2 | Y | 410 | 1000 | 105 | Y |  |  |  | H2 tank pressure 1060psi in and out of N winds, but construction has continued on station power outage 2048-54z
2016-07-19 | 201 | 1237 | TG | offline | offline | offline | 1610 | 240 | 2 | Y | 420 | 1000 | 105 | Y |  |  |  | H2 tank pressure 1060psi
2016-07-20 | 202 | 1609 | TG/EB | offline | offline | offline | 1620 | 240 | 2 | Y | 420 | 1000 | 105 | Y |  |  |  | H2 tank pressure 1010 psi
2016-07-21 | 203 | 1224 | TG/EB | offline | offline | offline | 1610 | 240 | 2 | Y | 420 | 1000 | 105 | Y |  |  |  | H2 tank pressure 1000 psi
2016-07-22 | 204 | 1141 | TG | offline | offline | offline | 1620 | 230 | 2 | Y | 420 | 1000 | 105 | Y |  |  |  | H2 tank pressure 1000 psi
2016-07-23 | 205 | 1317 | TG | offline | offline | offline | 1610 | 230 | 2 | Y | 420 | 1000 | 105 | Y |  |  |  | H2 tank pressure 995 psi
2016-07-24 | 206 | 1747 | TG | offline | offline | offline | 1620 | 230 | 2 | Y | 420 | 1000 | 105 | Y |  |  |  | H2 tank pressure 990 psi
2016-07-25 | 207 | 1234 | TG/EB | offline | offline | offline | 1620 | 220 | 2 | Y | 420 | 1000 | 105 | Y |  |  |  | H2 tank pressure 920psi
2016-07-26 | 208 | 1214 | TG/EB | offline | offline | offline | 1610 | 220 | 2 | Y | 420 | 1000 | 105 | Y |  |  |  | H2 tank pressure 900psi
2016-07-27 | 209 | 0 | 0
2016-07-28 | 210 | 1313 | TG/EB | offline | offline | offline | 1610 | 220 | 2 | Y | 410 | 1000 | 105 | Y |  |  |  | H2 tank pressure 850psi
2016-07-29 | 211 | 1611 | TG/EB | offline | offline | offline | 1610 | 220 | 2 | Y | 420 | 1000 | 105 | Y
2016-07-30 | 212 | 1236 | TG/EB | offline | offline | offline | 1620 | 210 | 2 | Y | 410 | 1000 | 105 | Y |  |  |  | H2 tank pressure 800psi
2016-07-31 | 213 | 1521 | EB | offline | offline | offline | 1610 | 210 | 2 | Y | 420 | 1000 | 105 | Y |  |  |  | H2 tank pressure 790psi
2016-08-01 | 214 | 1254 | EB | offline | offline | offline | 1610 | 220 | 2 | Y | 410 | 1000 | 105 | Y |  |  |  | H2 tank pressure 760psi
2016-08-02 | 215 | 1259 | EB | offline | offline | offline | 1610 | 210 | 2 | Y | 420 | 1000 | 105 | Y |  |  |  | H2 tank pressure 740psi
2016-08-03 | 216 | 1121 | TG | offline | offline | offline | 1610 | 200 | 2 | Y | 410 | 1000 | 105 | Y |  |  |  | H2 tank pressure 710psi
2016-08-04 | 217 | 1400 | EB | offline | offline | offline | 1610 | 200 | 2 | Y | 410 | 1000 | 105 | Y |  |  |  | H2 tank pressure 680psi
2016-08-05 | 218 | 1353 | EB/TG | offline | offline | offline | 1610 | 200 | 2 | Y | 410 | 1000 | 105 | Y |  |  |  | H2 tank pressure 670psi
2016-08-06 | 219 | 1717 | EB | offline | offline | offline | 1610 | 200 | 2 | Y | 410 | 1000 | 105 | Y |  |  |  | H2 tank pressure 610psi
2016-08-07 | 220 | 1632 | TG | offline | offline | offline | 1620 | 200 | 2 | Y | 410 | 1000 | 105 | Y |  |  |  | H2: 600psi
2016-08-08 | 221 | 1541 | TG | offline | offline | offline | 1620 | 200 | 2 | Y | 410 | 1000 | 100 | Y |  |  |  | H2: 580psi
2016-08-09 | 222 | 1832 | TG | offline | offline | offline | 1620 | 190 | 2 | Y | 410 | 1000 | 105 | Y |  |  |  | H2: 560psi
2016-08-10 | 223 | 1400 | EB | offline | offline | offline | 1620 | 200 | 2 | Y | 410 | 1000 | 105 | Y |  |  |  | H2: 540psi
2016-08-11 | 224 | 1243 | EB | offline | offline | offline | 1600 | 200 | 2 | Y | 405 | 1000 | 105 | Y |  |  |  | H2: 510psi
2016-08-12 | 225 | 1534 | TG | offline | offline | offline | 1610 | 190 | 2 | Y | 410 | 1000 | 105 | Y |  |  |  | H2: 500psi Power outage 1609-1611z. 1642-1644z
2016-08-13 | 226 | 1755 | EB | offline | offline | offline | 1600 | 200 | 2 | Y | 410 | 1000 | 105 | Y |  |  |  | H2: 480psi
2016-08-14 | 227 | 1759 | EB | offline | offline | offline | 1600 | 200 | 2 | Y | 410 | 1000 | 105 | Y |  |  |  | H2:450psi
2016-08-15 | 228 | 1907 | TG/EB | 60 | * | 47 | 1600 | 190 | 2 | Y | 410 | 990 | 105 | Y |  |  |  | H2:450psi
2016-08-16 | 229 | 1810 | TG | 60 | * | 46 | 1530 | 1820 | 2 | Y | 400 | 970 | 150 | Y | * | * | * | Notified Jacque of this, and he reset to 105, jacque swtiched out methane std 2
2016-08-17 | 230 | 1652 | TG | 60 | * | 45 | 400 | 1800 | 2 | Y | 400 | 990 | 105 | Y | * | * | * | Looks like Jacque switched CH4 std 1 tank from the large tank to the small yellow tank that has much less pressure
2016-08-18 | 231 | 1731 | TG | 60 | * | 43 | 400 | 1800 | 2 | y | 400 | 990 | 105 | Y | * | * | *
2016-08-19 | 232 | 1612 | EB/TG | 60 | * | 37 | 400 | 1800 | 2 | Y | 400 | 990 | 105 | Y | * | * | *
2016-08-20 | 233 | 1222 | TG | 60 | * | 45 | 400 | 1790 | 2 | Y | 400 | 990 | 105 | Y | * | * | *
2016-08-21 | 234 | 1545 | EB | 60 | * | 47 | 400 | 1790 | 2 | Y | 400 | 990 | 95 | Y | * | * | *
2016-08-22 | 235 | 1634 | TG | 60 | * | 45 | 400 | 1790 | 2 | Y | 400 | 990 | 105 | Y | * | * | *
2016-08-23 | 236 | 1255 | EB | 60 | * | 44 | 400 | 1780 | 2 | Y | 410 | 1000 | 105 | Y | * | * | *
2016-08-24 | 237 | 1330 | TG | 60 | * | 46 | 400 | 1790 | 2 | Y | 400 | 990 | 105 | Y | * | * | *
2016-08-25 | 238 | 1223 | EB | 60 | * | 45 | 400 | 1780 | 2 | Y | 410 | 980 | 105 | Y | * | * | *
2016-08-26 | 239 | 1352 | TG | 60 | * | 43 | 400 | 1790 | 2 | Y | 410 | 970 | 105 | Y | * | * | *
2016-08-27 | 240 | 1744 | EB | 60 | * | 46 | 400 | 1790 | 2 | Y | 400 | 960 | 105 | Y | * | * | *
2016-08-28 | 241 | 1715 | TG | 60 | * | 44 | 390 | 1790 | 2 | Y | 400 | 960 | 105 | Y | * | * | * | Is NMHC dropping faster than it should? Didn't see anything with a snoop test
2016-08-29 | 242 | 1354 | EB | 60 | * | 41 | 400 | 1770 | 2 | Y | 400 | 960 | 90 | Y | * | * | *
2016-08-30 | 243 | 1300 | TG | 60 | * | 45 | 400 | 1780 | 2 | Y | 400 | 960 | 105 | Y | * | * | *
2016-08-31 | 244 | 1306 | EB | 60 | * | 41 | 400 | 1780 | 2 | Y | 400 | 955 | 105 | Y | * | * | *
2016-09-01 | 245 | 1316 | TG | 60 | * | 45 | 400 | 1780 | 2 | Y | 400 | 960 | 105 | Y | * | * | *
2016-09-02 | 246 | 1700 | EB | 60 | * | 43 | 400 | 1760 | 2 | Y | 400 | 950 | 100 | Y | * | * | *
2016-09-03 | 247 | 1323 | TG | 60 | * | 43 | 400 | 1760 | 2 | Y | 390 | 950 | 100 | Y | * | * | * | Added 1.5 L DI H2O
2016-09-04 | 248 | 1419 | EB | 60 | * | 47 | 390 | 1760 | 2 | Y | 390 | 950 | 105 | Y | * | * | *
2016-09-05 | 249 | 1250 | TG | 60 | * | 45 | 400 | 1780 | 2 | Y | 390 | 950 | 105 | Y | * | * | *
2016-09-06 | 250 | 1342 | EB | 60 | * | 45 | 400 | 1760 | 2 | Y | 400 | 950 | 105 | Y | * | * | *
2016-09-07 | 251 | 1243 | TG | 60 | * | 45 | 400 | 1780 | 2 | Y | 400 | 950 | 105 | y | * | * | *
2016-09-08 | 252 | 1237 | EB | 60 | * | 45 | 400 | 1760 | 2 | Y | 400 | 950 | 105 | Y | * | * | *
2016-09-09 | 253 | 1249 | TG | 60 | * | 45 | 400 | 1770 | 2 | Y | 400 | 940 | 105 | Y | * | * | *
2016-09-10 | 254 | 1233 | EB | 60 | * | 45 | 400 | 1750 | 2 | Y | 400 | 950 | 105 | Y | * | * | *
2016-09-11 | 255 | 1742 | TG | 60 | * | 48 | 400 | 1770 | 2 | Y | 400 | 950 | 105 | Y | * | * | *
2016-09-12 | 256 | 1237 | EB | 60 | * | 45 | 400 | 1660 | 2 | Y | 400 | 950 | 105 | Y | * | * | *
2016-09-13 | 257 | 1237 | TG | 60 | * | 45 | 400 | 1770 | 2 | Y | 400 | 950 | 105 | Y | * | * | *
2016-09-14 | 258 | 1246 | EB | 60 | * | 46 | 400 | 1750 | 2 | Y | 400 | 940 | 105 | Y | * | * | *
2016-09-15 | 259 | 1220 | TG | 60 | * | 46 | 400 | 1760 | 2 | Y | 400 | 940 | 105 | Y | * | * | *
2016-09-16 | 260 | 1340 | EB | 60 | * | 45 | 400 | 1750 | 2 | Y | 400 | 950 | 105 | Y | * | * | *
2016-09-17 | 261 | 1247 | TG | 60 | * | 45 | 400 | 1760 | 2 | Y | 400 | 940 | 105 | Y | * | * | *
2016-09-18 | 262 | 1457 | EB | 60 | * | 45 | 400 | 1750 | 2 | Y | 400 | 940 | 105 | Y | * | * | *
2016-09-19 | 263 | 1323 | TG | 60 | * | 45 | 400 | 1760 | 2 | Y | 400 | 940 | 105 | Y | * | * | *
2016-09-20 | 264 | 1324 | EB | 60 | * | 67 | 400 | 1740 | 2 | Y | 400 | 940 | 105 | Y | * | * | *
2016-09-21 | 265 | 1226 | TG | 60 | * | 45 | 400 | 1750 | 2 | Y | 400 | 940 | 105 | Y | * | * | *
2016-09-22 | 266 | 1318 | EB | 60 | * | 43 | 400 | 1750 | 2 | Y | 400 | 950 | 105 | Y | * | * | *
2016-09-23 | 267 | 1216 | TG | 60 | * | 46 | 400 | 1750 | 2 | Y | 400 | 940 | 105 | Y | * | * | *
2016-09-24 | 268 | 1409 | EB | 60 | * | 45 | 400 | 1740 | 2 | Y | 400 | 940 | 105 | Y | * | * | *
2016-09-25 | 269 | 1337 | TG | 60 | * | 50 | 400 | 1740 | 2 | y | 400 | 940 | 105 | Y | * | * | *
2016-09-26 | 270 | 1257 | EB | 60 | * | 45 | 400 | 1740 | 2 | Y | 400 | 940 | 105 | Y | * | * | * | Added 1.5 L DI H2O
2016-09-27 | 271 | 1207 | TG | 60 | * | 45 | 400 | 1740 | 2 | Y | 400 | 940 | 105 | Y | * | * | *
2016-09-28 | 272 | 1325 | EB | 60 | * | 45 | 400 | 1740 | 2 | Y | 400 | 940 | 105 | Y | * | * | *
2016-09-29 | 273 | 1247 | TG | 60 | * | 44 | 400 | 1740 | 2 | Y | 400 | 940 | 105 | Y | * | * | *
2016-09-30 | 274 | 1402 | EB | 60 | * | 44 | 400 | 1730 | 2 | Y | 400 | 940 | 105 | Y | * | * | *
2016-10-01 | 275 | 1323 | TG | 60 | * | 43 | 400 | 1740 | 2 | Y | 400 | 940 | 105 | Y | * | * | *
2016-10-02 | 276 | 1403 | EB | 60 | * | 43 | 400 | 1730 | 2 | Y | 400 | 940 | 105 | Y | * | * | *
2016-10-03 | 277 | 1510 | TG | 60 | * | 44 | 400 | 1730 | 2 | Y | 400 | 940 | 105 | Y | * | * | *
2016-10-04 | 278 | 1219 | EB | 60 | * | 44 | 400 | 1720 | 2 | Y | 400 | 940 | 105 | Y | * | * | *
2016-10-05 | 279 | 1408 | TG | 60 | * | 44 | 400 | 1720 | 2 | Y | 400 | 940 | 105 | Y | * | * | *
2016-10-06 | 280 | 1314 | EB | 60 | * | 42 | 400 | 1710 | 2 | Y | 400 | 920 | 105 | Y | * | * | *
2016-10-07 | 281 | 1303 | TG | 60 | * | 42 | 400 | 1730 | 2 | Y | 400 | 930 | 105 | Y | * | * | * | Added 800mL DI H2O
2016-10-08 | 282 | 0 | 0 | [14 empty] | High winds/low visibility, didn't make it out to TAWO
2016-10-09 | 283 | 1459 | TG | 60 | * | 44 | 400 | 1730 | 2 | Y | 400 | 930 | 105 | Y | * | * | *
2016-10-10 | 284 | 1644 | EB | 60 | * | 44 | 400 | 1710 | 2 | Y | 400 | 930 | 105 | Y | * | * | *
2016-10-11 | 285 | 1338 | TG | 60 | * | 44 | 400 | 1720 | 2 | Y | 400 | 930 | 105 | Y | * | * | *
2016-10-12 | 286 | 0 | 0
2016-10-13 | 287 | 1303 | TG | 60 | * | 43 | 400 | 1720 | 2 | Y | 400 | 930 | 105 | Y | * | * | *
2016-10-14 | 288 | 1716 | EB | 60 | * | 43 | 400 | 1700 | 2 | Y | 400 | 930 | 105 | Y | * | * | *
2016-10-15 | 289 | 1300 | TG | 60 | * | 44 | 400 | 1700 | 2 | Y | 400 | 930 | 105 | Y | * | * | *
2016-10-16 | 290 | 0 | 0
2016-10-17 | 291 | 1431 | TG | 60 | * | 43 | 400 | 1710 | 2 | Y | 400 | 920 | 105 | Y | * | * | *
2016-10-18 | 292 | 1259 | EB | 60 | * | 44 | 400 | 1700 | 2 | Y | 390 | 940 | 105 | Y | * | * | *
2016-10-19 | 293 | 1257 | TG | 60 | * | 43 | 400 | 1710 | 2 | Y | 390 | 940 | 105 | Y | * | * | * | climbed tower and removed blown snow from inlet
2016-10-20 | 294 | 0 | 0 | [14 empty] | Con 1. Didn't go to TAWO.
2016-10-21 | 295 | 1619 | TG | 60 | * | 35 | 400 | 1710 | 2 | Y | 390 | 930 | 105 | Y | * | * | *
2016-10-22 | 296 | 1213 | EB | 60 | * | 43 | 400 | 1690 | 2 | Y | 400 | 940 | 105 | Y | * | * | *
2016-10-23 | 297 | 1331 | TG | 60 | * | 44 | 400 | 1690 | 2 | Y | 400 | 930 | 105 | Y | * | * | *
2016-10-24 | 298 | 1544 | TG/MG/NB | 60 | * | 43 | 400 | 1690 | 2 | Y | 400 | 930 | 95 | Y | * | * | *
2016-10-25 | 299 | 1133 | MG/TG | 61 | * | 72 | 395 | 1690 | 2 | Y | 400 | 925 | 100 | Y | * | * | * | Cleaned inlet at 1142-1155z.
2016-10-26 | 300 | 1213 | TG | 61 | * | 43 | 400 | 1680 | 2 | Y | 400 | 920 | 105 | Y | * | * | * | Added 1 L of DI water to H2 Generator
2016-10-27 | 301 | 1128 | MG/NB | 61 | * | 35 | 400 | 1680 | 2 | Y | 400 | 920 | 100 | Y | * | * | *
2016-10-28 | 302 | 1135 | MG/NB | 61 | * | 34 | 400 | 1680 | 2 | Y | 390 | 910 | 100 | Y | * | * | *
2016-10-29 | 303 | 1150 | MG/NB | 61 | * | 35 | 390 | 1680 | 2 | Y | 390 | 910 | 105 | Y | * | * | *
2016-10-30 | 304 | 1133 | NB | 61 | * | 44 | 390 | 1680 | 2 | Y | 390 | 910 | 100 | Y | * | * | *
2016-10-31 | 305 | 1318 | NB | 61 | * | 44 | 390 | 1680 | 2 | Y | 390 | 910 | 105 | Y | * | * | *
2016-11-01 | 306 | 1253 | MG | 61 | * | 44 | 390 | 1670 | 2 | Y | 390 | 900 | 105 | Y | * | * | *
2016-11-02 | 307 | 1251 | NB | 61 | * | 44 | 390 | 1670 | 2 | Y | 390 | 910 | 105 | Y | * | * | *
2016-11-03 | 308 | 1245 | MG | 61 | * | 44 | 390 | 1670 | 2 | Y | 380 | 900 | 105 | Y | * | * | *
2016-11-04 | 309 | 1309 | MG | 61 | * | 64 | 390 | 1670 | 2 |  | 390 | 910 | 105 | Y | * | * | *
2016-11-05 | 310 | 1230 | MG | 61 | * | 44 | 390 | 1670 | 2 | Y | 390 | 910 | 105 | Y | * | * | *
2016-11-06 | 311 | 1325 | NB | 61 | * | 41 | 390 | 1660 | 2 | Y | 390 | 910 | 105 | Y | * | * | *
2016-11-07 | 312 | 1230 | MG | 61 | * | 43 | 390 | 1660 | 2 | Y | 390 | 910 | 100 | Y | * | * | *
2016-11-08 | 313 | 1225 | MG | 61 | * | 41 | 390 | 1660 | 2 | Y | 390 | 910 | 100 | Y | * | * | *
2016-11-09 | 314 | 1216 | MG | 61 | * | 44 | 390 | 1660 | 2 | Y | 390 | 910 | 100 | Y | * | * | *
2016-11-10 | 315 | 1242 | MG | 61 | * | 44 | 390 | 1660 | 2 | Y | 390 | 910 | 100 | Y | * | * | *
2016-11-11 | 316 | 0 | 0 | [14 empty] | Did not go to TAWO condition 1
2016-11-12 | 317 | 1240 | MG/NB | 61 | * | 44 | 390 | 1570 | 2 | Y | 350 | 870 | 100 | Y | * | * | * | Cold in TAWO today do to strong east wind (42F). Tank pressures low as a result.
2016-11-13 | 318 | 1250 | MG | 61 | * | 40 | 390 | 1660 | 2 | Y | 390 | 910 | 105 | Y | * | * | *
2016-11-14 | 319 | 1900 | NB | 61 | * | 43 | 390 | 1660 | 2 | Y | 380 | 910 | 100 | Y | * | * | *
2016-11-15 | 320 | 1400 | MG | 61 | * | 41 | 390 | 1650 | 2 | Y | 380 | 910 | 100 | Y | * | * | *
2016-11-16 | 321 | 1302 | MG | 61 | * | 43 | 390 | 1640 | 2 | Y | 380 | 900 | 105 | Y | * | * | *
2016-11-17 | 322 | 1218 | MG/NB | 61 | * | 35 | 390 | 1670 | 2 | Y | 390 | 900 | 105 | Y | * | * | * | Added 1 L of DI water to H2 Generator
2016-11-18 | 323 | 1323 | MG | 60 | * | 42 | 390 | 1640 | 2 | Y | 390 | 900 | 105 | Y | * | * | *
2016-11-19 | 324 | 1235 | NB | 61 | * | 44 | 380 | 1620 | 2 | Y | 390 | 900 | 105 | Y | * | * | * | MG climbed tower 1252 - 1336z and removed frost from sample inlet.
2016-11-20 | 325 | 1312 | NB | 61 | * | 43 | 390 | 1630 | 2 | Y | 390 | 900 | 105 | Y | * | * | *
2016-11-21 | 326 | 1305 | NB | 61 | * | 43 | 390 | 1630 | 2 | Y | 380 | 900 | 105 | Y | * | * | *
2016-11-22 | 327 | 1301 | MG/NB | 61 | * | 44 | 390 | 1630 | 2 | Y | 380 | 900 | 105 | Y | * | * | *
2016-11-23 | 328 | 1230 | NB | 61 | * | 44 | 390 | 1620 | 2 | Y | 380 | 900 | 105 | Y | * | * | *
2016-11-24 | 329 | 1250 | NB | 61 | * | 43 | 390 | 1620 | 2 | Y | 380 | 890 | 100 | Y | * | * | *
2016-11-25 | 330 | 1249 | MG | 61 | * | 44 | 390 | 1620 | 2 | Y | 380 | 900 | 105 | Y | * | * | *
2016-11-26 | 331 | 1335 | NB | 61 | * | 43 | 390 | 1620 | 2 | Y | 380 | 900 | 105 | Y | * | * | *
2016-11-27 | 332 | 1420 | MG | 61 | * | 43 | 390 | 1620 | 2 | Y | 380 | 900 | 105 | Y | * | * | *
2016-11-28 | 333 | 1235 | NB | 61 | * | 44 | 390 | 1630 | 2 | Y | 380 | 900 | 105 | Y | * | * | *
2016-11-29 | 334 | 1245 | NB | 61 | * | 43 | 390 | 1620 | 2 | Y | 380 | 900 | 105 | Y | * | * | *
2016-11-30 | 335 | 1325 | MG | 61 | * | 48 | 390 | 1620 | 2 | Y | 370 | 900 | 105 | Y | * | * | *
2016-12-01 | 336 | 1250 | NB | 61 | * | 44 | 390 | 1620 | 2 | Y | 370 | 900 | 105 | Y | * | * | * | MG climbed tower 1302-1329z and removed frost from sample inlet.
2016-12-02 | 337 | 1321 | MG | 61 | * | 44 | 390 | 1620 | 2 | Y | 360 | 890 | 105 | Y | * | * | *
2016-12-03 | 338 | 1235 | NB | 61 | * | 43 | 390 | 1620 | 2 | Y | 360 | 890 | 105 | Y | * | * | *
2016-12-04 | 339 | 1348 | NB | 61 | * | 40 | 390 | 1620 | 2 | Y | 360 | 900 | 105 | Y | * | * | *
2016-12-05 | 340 | 1242 | NB | 61 | * | 44 | 390 | 1620 | 2 | Y | 360 | 900 | 105 | Y | * | * | *
2016-12-06 | 341 | 1244 | MG | 61 | * | 44 | 390 | 1620 | 2 | Y | 360 | 890 | 105 | Y | * | * | * | NB climbed tower 1251-1320z and removed frost from sample inlet.
2016-12-07 | 342 | 1240 | MG | 61 | * | 43 | 390 | 1620 | 2 | Y | 360 | 890 | 105 | Y | * | * | *
2016-12-08 | 343 | 1234 | MG/NB | 61 | * | 44 | 390 | 1610 | 2 | Y | 360 | 900 | 105 | Y | * | * | * | Added 1 L of DI water to H2 Generator
2016-12-09 | 344 | 1328 | MG | 61 | * | 44 | 390 | 1610 | 2 | Y | 360 | 890 | 105 | Y | * | * | *
2016-12-10 | 345 | 1920 | MG | 61 | * | 44 | 390 | 1610 | 2 | Y | 360 | 890 | 105 | Y | * | * | *
2016-12-11 | 346 | 1340 | MG | 61 | * | 43 | 390 | 1610 | 2 | Y | 360 | 890 | 105 | Y | * | * | *
2016-12-12 | 347 | 1230 | NB | 61 | * | 43 | 390 | 1610 | 2 | Y | 360 | 890 | 105 | Y | * | * | *
2016-12-13 | 348 | 1225 | NB | 61 | * | 43 | 390 | 1600 | 2 | Y | 360 | 890 | 105 | Y | * | * | *
2016-12-14 | 349 | 1304 | MG | 61 | * | 43 | 390 | 1590 | 2 | Y | 350 | 880 | 105 | Y | * | * | *
2016-12-15 | 350 | 1415 | MG | 61 | * | 43 | 390 | 1600 | 2 | Y | 350 | 890 | 105 | Y | * | * | *
2016-12-16 | 351 | 1308 | MG | 61 | * | 41 | 390 | 1590 | 2 | Y | 350 | 880 | 105 | Y | * | * | *
2016-12-17 | 352 | 1252 | MG | 61 | * | 43 | 390 | 1590 | 2 | Y | 350 | 890 | 105 | Y | * | * | *
2016-12-18 | 353 | 1328 | NB | 61 | * | 41 | 390 | 1580 | 2 | Y | 360 | 890 | 105 | Y | * | * | *
2016-12-19 | 354 | 1220 | NB | 61 | * | 42 | 390 | 1580 | 2 | Y | 360 | 880 | 105 | Y | * | * | *
2016-12-20 | 355 | 1506 | MG | 61 | * | 43 | 390 | 1570 | 2 | Y | 350 | 880 | 100 | Y | * | * | *
2016-12-21 | 356 | 1325 | MG | 61 | * | 41 | 390 | 1570 | 2 | Y | 350 | 880 | 100 | Y | * | * | *
2016-12-22 | 357 | 1350 | MG | 61 | * | 42 | 390 | 1540 | 2 | Y | 350 | 880 | 100 | Y | * | * | *
2016-12-23 | 358 | 1244 | NB | 61 | * | 43 | 390 | 1570 | 2 | Y | 360 | 880 | 100 | Y | * | * | *
2016-12-24 | 359 | 1326 | NB | 61 | * | 43 | 390 | 1570 | 2 | Y | 360 | 880 | 100 | Y | * | * | *
2016-12-25 | 360 | 1456 | MG | 61 | * | 43 | 390 | 1560 | 2 | Y | 360 | 880 | 100 | Y | * | * | *
2016-12-26 | 361 | 1235 | NB | 61 | * | 43 | 390 | 1560 | 2 | Y | 360 | 880 | 100 | Y | * | * | * | Added 1 L of DI water to H2 Generator
2016-12-27 | 362 | 1034 | MG | 61 | * | 44 | 320 | 1510 | 2 | Y | 360 | 850 | 100 | Y | * | * | *
2016-12-28 | 363 | 1358 | MG | 61 | * | 41 | 390 | 1560 | 2 | Y | 350 | 860 | 90 | Y | * | * | *
2016-12-29 | 364 | 1308 | MG/NB | 61 | * | 43 | 390 | 1560 | 2 | Y | 350 | 880 | 95 | Y | * | * | *
2016-12-30 | 365 | 1314 | MG/NB | 61 | * | 41 | 390 | 1570 | 2 | Y | 350 | 880 | 105 | Y | * | * | *
2016-12-31 | 366 | 1332 | NB | 61 | * | 43 | 390 | 1560 | 2 | Y | 350 | 870 | 105 | Y | * | * | *

## Aethalometer AE16-AE33
Date | Julian Day | Time (UTC) | Observer's Initials | AE16 Sample inlet clear (AE33 inlet as of July 2016)? | Check date/time (GMT) | Panel Readout BC ng/m3 AE16/AE33 | Flow rate, lpm AE16/AE33 | Check Disk | Weekly Data Transfer | Weekly Tape Check | Comments
2016-01-01 | 1 | 1510 | JG | clr | * | -25/-533 | 6.2/4.0 | *
2016-01-02 | 2 | 1229 | NB | rm frost | * | -123/-268 | 6.3/4.0 | *
2016-01-03 | 3 | 1310 | JG | rm frost | * | -223/431 | 6.3/4.0 | *
2016-01-04 | 4 | 1228 | NB | rm frost | * | -159/-270 | 6.4/4.0 | * |  | *
2016-01-05 | 5 | 1230 | JG | clr | * | -374/-85 | 6.3/4.0 | * |  |  | AE33 annual PMs: Inspected inlet tubing; cleaned optical chamber; lubricated slides; clean air test finished at 1341z (passed); stability test; stability test finished at 1408z (passed); flow verification at 1420z.
2016-01-06 | 6 | 1307 | NB | rm frost | * | -206/-375 | 6.4/4.0 | * |  |  | AE33 leakage test finished at 1355z, reported 2% leakage.
2016-01-07 | 7 | 1235 | JG | rm frost | * | -200/-773 | 6.4/4.0 | *
2016-01-08 | 8 | 1347 | NB | clr | * | 10/201 | 6.5/4.0 | * |  |  | Adjusted AE33 leakage factor to .02 based on the results of the previous leakage test.
2016-01-09 | 9 | 1240 | JG | lt frost | * | -649/279 | 6.3/4.0 | *
2016-01-10 | 10 | 1504 | NB | rm frost | * | -229/55 | 6.4/4.0 | *
2016-01-11 | 11 | 1240 | JG | rm frost | * | -289/-402 | 6.3/4.0 | * | * | *
2016-01-12 | 12 | 1232 | NB | rm frost | * | -380/454 | 5.1/4.0 | * |  |  | AE16 showing low flow alarm. Shut down AE16 1303-1346z to backflush sample inlet. Flow reestablished at 6.6 LPM.
2016-01-13 | 13 | 1255 | JG | rm frost | * | -297/-883 | 6.4/4.0 | * |  |  | Planned power outage for e-gen test 1307-1309 and 1321-1324 UTC. AE16 off 1307-1328 UTC.
2016-01-14 | 14 | 1225 | NB | rm frost | * | 152/-65 | 6.3/4.0 | failure |  |  | AE16 showing disk problem alarm and display was not updating. Reformatted diskette and loaded it with old .CSV files. Restarted at 1236z. Restart unsuccessful. Installed new diskette and restarted again at 1242z.
2016-01-15 | 15 | 1255 | JG | rm frost | * | -270/89 | 6.2/4.0 | *
2016-01-16 | 16 | 1304 | NB | rm frost | * | -946/-200 | 6.4/4.0 | failure |  |  | AE16 showing disk problem alarm. Reformatted diskette at 1315z.
2016-01-17 | 17 | 1725 | JG | rm frost | * | -325/-991 | 6.4/4.0 | failure |  |  | AE16 disk error. Disk could not be reformatted, had to insert a new one.
2016-01-18 | 18 | 1229 | NB | clr | * | -532/352 | 6.3/4.0 | * |  | *
2016-01-19 | 19 | 1310 | JG | rm frost | * | -2093/-1234 | 6.2/4.0 | *
2016-01-20 | 20 | 1232 | NB | clr | * | 33/-264 | 6.4/4.0 | failure |  |  | AE16 display frozen showing disk problem error message. Shut instrument down, reformatted disk, and restarted at 1236z.
2016-01-21 | 21 | 1250 | JG | rm frost | * | -253/-310 | 6.3/4.0 | *
2016-01-22 | 22 | 1234 | NB | lt frost | * | -281/-800 | 6.2/4.0 | *
2016-01-23 | 23 | 1210 | JG | rm frost | * | -600/-323 | 6.2/4.0 | *
2016-01-24 | 24 | 1258 | NB | lt frost | * | -215/651 | 6.0/4.0 | *
2016-01-25 | 25 | 1230 | JG | rm frost | * | -375/179 | 6.1/4.0 | * |  | *
2016-01-26 | 26 | 1254 | NB | rm frost | * | -321/-45 | 5.7/4.0 | *
2016-01-27 | 27 | 1340 | JG | rm frost | * | -171/154 | 6.2/4.0 | * |  |  | End piece for AE33 inlet found on snow sfc. Brought in to thaw, reinstalled at 1400z.
2016-01-28 | 28 | 1234 | NB | lt frost | * | -203/553 | 6.1/4.0 | *
2016-01-29 | 29 | 1255 | JG | rm frost | * | -543/791 | 6.2/4.0 | *
2016-01-30 | 30 | 1225 | NB | rm frost | * | -539/20 | 6.3/4.0 | *
2016-01-31 | 31 | 1455 | JG | rm frost | * | -889/-597 | 6.3/4.0 | *
2016-02-01 | 32 | 1225 | NB | lt frost | * | -547/54 | 6.2/4.0 | * |  | *
2016-02-02 | 33 | 1240 | JG | rm frost | * | 5/498 | 6.3/4.0 | failure |  |  | Disk error, reformatted diskette and restarted instrument at 1238z. Disk error again at 1705z; reformatted and replaced diskette, did not have to restart AE16.
2016-02-03 | 34 | 1248 | NB | rm frost | * | -5/-434 | 5.9/4.0 | failure |  |  | AE16 frozen with disk error. Installed new diskette and restarted at 1251z.
2016-02-04 | 35 | 1220 | JG | rm frost | * | -556/-102 | 5.8/4.0 | failure |  |  | Disk error. Reformated diskette and installed at 1223z.
2016-02-05 | 36 | 1333 | NB | rm frost | * | 20/-168 | 6.3/4.0 | failure |  |  | AE16 frozen with disk error. Reformatted diskette and restarted at 1343z.
2016-02-06 | 37 | 1355 | JG | rm frost | * | 104/700 | 6.1/4.0 | failure |  |  | AE16 frozen with disk error. Reformatted diskette and restarted at 1500z.
2016-02-07 | 38 | 1455 | MB/JJ | rm frost | * | -91/578 | 6.3/4.0 | 0h |  |  | check light blinking disk at 0 hours left running as data is collecting.
2016-02-08 | 39 | 1240 | MB/JJ/SD | rm frost | * | -209/197 | 6.4/4.0 | failure |  | * | Disk error. Reformated diskette and installed at 1233z.
2016-02-09 | 40 | 1230 | MB/JJ/NB | rm frost | * | -367/320 | 6.5/4.0 | * |  |  | Shut down AE16 at 1258z for station GEN swap. Shut down AE33 at 1430z for flow verification test. Restarted AE33 at 1455z.
2016-02-10 | 41 | 1226 | SWD/JG/NB | rm frost | * | -41/-108 | 6.1/4.0 | *
2016-02-11 | 42 | 1340 | MB/JJ | rm frost | * | -258/236 | 6.2/3.4 | * |  |  | AE33 Tstd NIST chosen, 0degC displayed, flow displays 3.4 lpm.
2016-02-12 | 43 | 1625 | MB/JJ | rm frost | * | -606/-792 | 6.2/3.4 | *
2016-02-13 | 44 | 1220 | MB/JJ | rm frost | * | -153/-153 | 6.2/3.4 | failure |  |  | Disk error. Reformated diskette and installed at 1244z.
2016-02-14 | 45 | 1700 | MB/JJ | rm frost | * | -55/336 | 6.1/3.4 | failure |  |  | Disk error and incorrect time (date was correct). Disk would not reformat. Inserted new disk. Disk error persisted. Shut down AE16 and restarted instrument at 1755z.
2016-02-15 | 46 | 1641 | JJ | rm frost | * | -92/-266 | 5.7/3.4 | * |  | *
2016-02-16 | 47 | 1321 | jj/mb | rm frost | * | -485/137 | 5.9/3.4 | *
2016-02-17 | 48 | 1415 | MB/JJ | rm frost | * | 83/436 | 5.9/3.4 | failure |  |  | Disk error. Reformatted diskette and inserted disk. Time & date did not display and instrument would not 'stop and exit cleanly.' Had to shut down and restart at 1507z.
2016-02-18 | 49 | 1405 | MB/JJ | clr | * | -285/-221 | 5.9/3.4 | *
2016-02-19 | 50 | 1349 | JJ/MB | rm frost | * | -120/19 | 5.8/3.4 | failure |  |  | Disk error. Reformatted diskette and installed at 1420z.
2016-02-20 | 51 | 1311 | JJ/MB | rm frost | AE16 stuck at 1149 | 4/257 | 6.1/3.4 | failure |  |  | Disk error and incorrect time (date was correct). Disk would not reformat. Inserted new disk. Disk error persisted. Shut down AE16 and restarted instrument at 1415z.
2016-02-21 | 52 | 1510 | JJ/MB | rm frost | * | -756/-396 | 6.2/3.4 | *
2016-02-22 | 53 | 1244 | JJ/MB | rm frost | * | -636/404 | 6.2/3.5 | * |  | *
2016-02-23 | 54 | 1704 | MB | rm frost | * | -116/-29 | 6.5/3.5 | *
2016-02-24 | 55 | 1314 | JJ/MB | clr | * | -483/-74 | 6.5/3.5 | *
2016-02-25 | 56 | 1336 | jj/mb | clr | * | 46/803 | 6.4/3.5 | *
2016-02-26 | 57 | 1348 | JJ/MB | rm frost | * | -246/-123 | 6.4/3.5 | *
2016-02-27 | 58 | 1305 | MB | rm frost | * | -189/169 | 6.3/3.5 | *
2016-02-28 | 59 | 1708 | JJ | rm frost | * | -61/28 | 6.4/3.5 | *
2016-02-29 | 60 | 1339 | jj/mb | rm frost | * | -87/-92 | 6.3/3.4 | * |  | *
2016-03-01 | 61 | 1248 | JJ/MB | rm frost | * | -108/-700 | 6.3/3.4 | *
2016-03-02 | 62 | 1711 | MB/JJ | rm frost | * | -331/636 | 6.4/3.4 | *
2016-03-03 | 63 | 1338 | MB/JJ | rm frost | AE16 time incorrect | 206/102 | 6.5/3.5 | failure |  |  | Disk error and incorrect time (date was correct). Disk would not reformat. Inserted new disk. Disk error persisted. Shut down AE16 and restarted instrument at 1401z.
2016-03-04 | 64 | 1309 | jj/mb | rm frost | * | -2710/-106 | 6.5/3.5 | *
2016-03-05 | 65 | 1235 | jj/mb | rm frost | * | -91/-1111 | 6.4/3.5 | *
2016-03-06 | 66 | 1342 | MB | rm frost | * | -309/267 | 6.2/3.4 | *
2016-03-07 | 67 | 1321 | jj/mb | rm frost | * | -430/256 | 6.3/3.4 | * |  | *
2016-03-08 | 68 | 1331 | jj/mb | rm frost | * | -570/27 | 6.5/3.4 | *
2016-03-09 | 69 | 1235 | JJ | clr | * | -111/555 | 6.5/3.4 | *
2016-03-10 | 70 | 0 | 0 | [7 empty] | White out conditions (Condition 1) did not travel to TAWO this day.
2016-03-11 | 71 | 1352 | JJ | rm frost | * | -299/-300 | 6.0/3.3 | *
2016-03-12 | 72 | 1437 | MB | rm frost | * | -3426/-874 | 6.1/3.4 | *
2016-03-13 | 73 | 1346 | JJ | rm frost | * | 278/293 | 6.2/3.4 | failure |  |  | Unexpected power outage from 1316-1323z, AE16 had to be restared after this.
2016-03-14 | 74 | 1353 | jj/mb | clr | * | -132/39 | 6.1/3.4 | * |  | *
2016-03-15 | 75 | 1533 | jj/mb | rm frost | * | -13/-117 | 6.5/3.5 | *
2016-03-16 | 76 | 1319 | MB | clr | * | 64/1077 | 6.6/3.5 | *
2016-03-17 | 77 | 1338 | jj/mb | clr | * | -170/788 | 6.6/3.5 | *
2016-03-18 | 78 | 1305 | MB/JJ | clr | * | -80/653 | 6.5/3.5 | *
2016-03-19 | 79 | 1854 | MB/JJ | rm frost | * | -238/-21 | 6.4/3.5 | *
2016-03-20 | 80 | 1334 | MB | rm frost | * | -199/-594 | 6.4/3.4 | *
2016-03-21 | 81 | 1307 | jj/mb | rm frost | * | -366/-612 | 6.3/3.4 | * |  | *
2016-03-22 | 82 | 1320 | MB | rm frost | * | -136/1187 | 6.4/3.4 | *
2016-03-23 | 83 | 1744 | MB/JJ | rm frost | * | -187/-227 | 6.2/3.4 | *
2016-03-24 | 84 | 1329 | jj/mb | rm frost | * | -432/-50 | 6.2/3.4 | *
2016-03-25 | 85 | 1303 | JJ | rm frost | * | -183/256 | 6.2/3.4 | *
2016-03-26 | 86 | 1303 | MB | rm frost | * | -1229/-666 | 6.2/3.4 | *
2016-03-27 | 87 | 1715 | JJ | clr | * | 8/-695 | 6.5/3.4 | failure |  |  | AE 16 had to be retstarted due to disk failure.
2016-03-28 | 88 | 1327 | MB | clr | * | -65/849 | 6.5/3.4 | * |  | *
2016-03-29 | 89 | 1151 | JJ | rm frost | * | 478/298 | 6.5/3.4 | *
2016-03-30 | 90 | 1134 | MB | rm frost | * | -4454/477 | 6.5/3.4 | *
2016-03-31 | 91 | 1718 | jj/mb | rm frost | * | -131/789 | 6.5/3.4 | *
2016-04-01 | 92 | 1304 | JJ | clr | * | -32/-393 | 6.4/3.4 | *
2016-04-02 | 93 | 1213 | jj/mb | clr | * | 28/-194 | 6.5/3.4 | *
2016-04-03 | 94 | 1550 | MB | clr | * | -125/52 | 6.6/3.4 | *
2016-04-04 | 95 | 1613 | jj/mb | rm frost | * | 12/483 | 6.6/3.4 | failure |  | *
2016-04-05 | 96 | 1251 | MB/JJ | rm frost | * | -204/-343 | 6.5/3.4 | * |  |  | Updated AE33 software to version 1.1.8.0 firmware version 518
2016-04-06 | 97 | 1332 | JJ | rm frost | * | -104/131 | 6.5/3.5 | * |  |  | Planned power outage today AE16 was restarted after regular power was restored.
2016-04-07 | 98 | 1224 | MB/JJ | clr | * | -2/-73 | 6.3/3.5 | *
2016-04-08 | 99 | 1139 | JJ/MB | clr | * | 114/-23 | 6.4/3.5 | failure |  |  | AE16 had to be restarted due to disk failure, 1212z.
2016-04-09 | 100 | 1254 | JJ/MB | clr | * | 28/91 | 6.5/3.6 | *
2016-04-10 | 101 | 1707 | JJ | clr | AE16 time incorrect | 94/-50 | 6.6/3.6 | failure |  |  | AE16 had to be restarted due to disk failure.
2016-04-11 | 102 | 1321 | JJ/MB | clr | * | -43/648 | 6.6/3.6 | * |  | *
2016-04-12 | 103 | 1554 | JJ/MB | clr | * | -40/-3 | 6.4/3.5 | *
2016-04-13 | 104 | 1317 | JJ/MB | clr | * | -115/9 | 6.5/3.5 | *
2016-04-14 | 105 | 1329 | JJ/MB | rm frost | * | 8/430 | 6.5/3.5 | * |  |  | Missing curved intake tube for Neph, CLAP, and AE33.
2016-04-15 | 106 | 1711 | JJ/MB/SWD | clr | * | -68/-90 | 6.3/3.5 | * |  |  | Missing curved intake tube for Neph, CLAP, and AE33.
2016-04-16 | 107 | 1309 | JJ/MB | rm frost | * | -76/-626 | 6.2/3.4 | * |  |  | Missing curved intake tube for Neph, CLAP, and AE33.
2016-04-17 | 108 | 1628 | MB | rm frost | AE16 time incorrect | 65/-572 | 6.1/3.4 | failure |  |  | Missing curved intake tube for Neph, CLAP, and AE33. AE16 had to be restarted due to disk failure.
2016-04-18 | 109 | 1249 | MB | rm frost | * | 42/66 | 6.4/3.4 | failure |  | * | Disk error. Replaced disk and reinstalled at 1249z.
2016-04-19 | 110 | 1152 | JJ/MB | clr | AE16 time reset | -56/191 | 6.9/3.4 | * |  |  | Missing curved intake tube for Neph, CLAP, and AE33.
2016-04-20 | 111 | 1143 | JJ/MB | clr | * | -314/384 | 6.9/3.5 | * |  |  | Missing curved intake tube for Neph, CLAP, and AE33.
2016-04-21 | 112 | 1122 | JJ/MB/SWD | rm frost | * | -181/-80 | 6.9/3.5 | * |  |  | Missing curved intake tube for Neph, CLAP, and AE33.
2016-04-22 | 113 | 1838 | JJ/MB | clr | * | -17/387 | 7.0/3.6 | * |  |  | Missing curved intake tube for Neph, CLAP, and AE33.
2016-04-23 | 114 | 1536 | 0 | [7 empty] | White out conditions (Condition 1) did not travel to TAWO this day.
2016-04-24 | 115 | 1536 | JJ/MB | clr |  | -63/-24 | 7.0/3.6 | * |  |  | Missing curved intake tube for Neph, CLAP, and AE33.
2016-04-25 | 116 | 1207 | JJ/MB | clr | * | -88/-293 | 7.1/3.6 | * |  | * | Missing curved intake tube for Neph, CLAP, and AE33.
2016-04-26 | 117 | 1130 | JJ | clr | * | -87/34 | 7.1/3.6 | * |  |  | Missing curved intake tube for Neph, CLAP, and AE33.
2016-04-27 | 118 | 1549 | JJ/MB | clr | * | -58/-248 | 6.9/3.6 | * |  |  | Missing curved intake tube for Neph, CLAP, and AE33.
2016-04-28 | 119 | 1315 | JJ/MB | rm frost | * | -122/-61 | 6.9/3.6 | * |  |  | Missing curved intake tube for Neph, CLAP, and AE33.
2016-04-29 | 120 | 1300 | MB/JJ | rm frost | * | 157/-483 | 6.8/3.6 | * |  |  | Attached teflon tape and new curved intake tube made out of 1/2" electrical conduit to Neph/CLAP/AE33 intake cyclone.
2016-04-30 | 121 | 1336 | MB/JJ | rm frost | * | 39/13 | 6.8/3.6 | *
2016-05-01 | 122 | 1624 | MB/JJ | clr | * | -166/405 | 6.6/3.5 | *
2016-05-02 | 123 | 1322 | JJ/MB | rm frost | * | 45/432 | 6.6/3.5 | * |  | * | Planned power outages from 1635 to 1700z. 1710z restarted AE33 after power outages. 1722z reset AE33 time. AE33 pump unplugged for power outage and restarted at 1723z.
2016-05-03 | 124 | 1230 | MB/JJ | rm frost | * | 1/543 | 6.5/3.5 | *
2016-05-04 | 125 | 1451 | MB/JJ | clr | * | -96/975 | 6.4/3.5 | *
2016-05-05 | 126 | 1528 | MB | clr | * | 92/544 | 6.4/3.5 | *
2016-05-06 | 127 | 1845 | JJ/MB | rm frost | * | -171/671 | 6.5/3.5 | *
2016-05-07 | 128 | 1315 | JJ/MB | rm frost | * | -151/-339 | 6.5/3.5 | *
2016-05-08 | 129 | 1558 | JJ | rm frost | * | -4/117 | 6.6/3.6 | *
2016-05-09 | 130 | 1635 | JJ/MB | clr | * | -124/-263 | 6.5/3.5 | * |  | *
2016-05-10 | 131 | 1648 | JJ/MB | clr | * | -282/-434 | 6.5/3.5 | *
2016-05-11 | 132 | 1753 | MB/JJ | clr | * | -179/-5 | 6.6/3.6 | *
2016-05-12 | 133 | 1622 | MB/JJ | rm frost | * | -75/-3 | 6.6/3.6 | *
2016-05-13 | 134 | 1308 | MB/JJ | clr | * | 12/-401 | 6.6/3.6 | *
2016-05-14 | 135 | 1149 | JJ/MB | clr | * | 88/5 | 6.6/3.5 | *
2016-05-15 | 136 | 1730 | MB | clr | * | 28/465 | 6.5/3.5 | *
2016-05-16 | 137 | 1609 | JJ/MB | rm frost | AE16 time reset | 69/750 | 6.6/3.5 | * |  | * | AE16 clock reset at 1618z
2016-05-17 | 138 | 1722 | JJ/MB | rm frost | * | 328/169 | 6.5/3.5 | *
2016-05-18 | 139 | 1309 | JJ/MB | clr | * | -68/-492 | 6.4/3.5 | *
2016-05-19 | 140 | 1322 | JJ | clr | * | 3/80 | 6.3/3.4 | *
2016-05-20 | 141 | 1850 | JJ/MB | clr | * | 0/-333 | 6.3/3.4 | *
2016-05-21 | 142 | 1228 | JJ/MB | rm frost | * | -10/201 | 6.3/3.4 | *
2016-05-22 | 143 | 2327 | JJ | clr | * | 6/-674 | 6.2/3.4 | *
2016-05-23 | 144 | 1656 | JJ | clr | * | 9/347 | 6.2/3.6 | * |  | *
2016-05-24 | 145 | 1211 | JJ/MB | rm frost | * | -108/834 | 6.3/3.6 | *
2016-05-25 | 146 | 1726 | JJ | clr | * | -51/-68 | 6.1/3.5 | *
2016-05-26 | 147 | 1232 | JJ/MB | clr | * | -113/-151 | 6.2/3.5 | *
2016-05-27 | 148 | 1845 | JJ/mb | clr | * | -80/-685 | 6.1/3.5 | *
2016-05-28 | 149 | 1221 | MB | clr | * | 38/-199 | 6.3/3.5 | *
2016-05-29 | 150 | 1530 | MB | clr | * | 44/-87 | 6.2/3.5 | *
2016-05-30 | 151 | 1226 | JJ | clr | * | -44/8 | 6.2/3.5 | * |  | * | Planned power outage today from 2130 to 2230. AE16 was shutdown during the outage
2016-05-31 | 152 | 1731 | JJ | clr | * | 48/196 | 6.5/3.6 | *
2016-06-01 | 153 | 1840 | JJ | clr | * | -68/384 | 6.6/3.5 | *
2016-06-02 | 154 | 1831 | MB/JJ | clr | * | 44/13 | 6.6/3.6 | *
2016-06-03 | 155 | 1737 | MB/JJ/HJ | clr | * | -95/21 | 6.4/3.5 | *
2016-06-04 | 156 | 1156 | MB | clr | * | -49/-275 | 6.6/3.6 | *
2016-06-05 | 157 | 1403 | JJ/TG/EB | clr | * | -72/741 | 6.5/3.5 | *
2016-06-06 | 158 | 1613 | MB/JJ/TG/EB | clr | * | 55/77 | 6.5/3.5 | * |  | *
2016-06-07 | 159 | 1536 | TG/EB/JJ | clr | * | -72/153 | 6.8/3.6 | *
2016-06-08 | 160 | 1309 | TG/EB/JJ/MB | clr | * | 7/-356 | 6.7/3.5 | * |  |  | BC readouts taken just after restart following planned power outage. AE16 restart after second planned power outage 1650z.
2016-06-09 | 161 | 1553 | TG/EB | clr | * | 11/891 | 6.8/3.5 | *
2016-06-10 | 162 | 1239 | TG/EB | clr | * | 41/78 | 6.7/3.5 | * |  |  | Work on ventilation duct from 1105 to 1330, higher room temps.
2016-06-11 | 163 | 1203 | TG/EB | clr | * | 10/391 | 6.7/3.5 | *
2016-06-12 | 164 | 1509 | TG/EB | clr | * | -65/0 | 6.7/3.5 | *
2016-06-13 | 165 | 1637 | TG/EB | clr | * | -54/1104 | 6.7/3.5 | * |  | *
2016-06-14 | 166 | 1153 | TG/EB | clr | * | -40/259 | 6.6/3.5 | *
2016-06-15 | 167 | 1202 | TG/EB | clr | * | 15/102 | 6.6/3.5 | *
2016-06-16 | 168 | 1156 | TG | clr | * | 27/-137 | 6.7/3.5 | *** |  |  | disk error "try to replace disk", put in new reformatted disk. Didn't work. Took old disk and erased some files and then inserted. Didn't work. Shut down and restarted aethelometer and it worked. Disk says 7wks. Also reset date and time on start up
2016-06-17 | 169 | 1314 | TG | clr | * | -100/-967 | 6.6/3.5 | *
2016-06-18 | 170 | 1241 | TG/EB | clr | * | 50/327 | 6.6/3.5 | *
2016-06-19 | 171 | 1808 | EB | clr | * | -8/653 | 6.5/3.4 | *
2016-06-20 | 172 | 1812 | TG/EB | clr | * | 105/139 | 6.6/3.4 | * |  | *
2016-06-21 | 173 | 1204 | TG/EB | clr | * | -46/365 | 6.5/3.4 | *
2016-06-22 | 174 | 1332 | TG/EB | clr | * | -80/213 | 6.6/3.6 | *
2016-06-23 | 175 | 1313 | EB | rm frost | * | 5/594 | 6.6/3.6 | *
2016-06-24 | 176 | 1329 | TG | clr | * | -2/73 | 6.6/3.6 | *
2016-06-25 | 177 | 1717 | EB/TG | clr | * | 9/119 | 6.5/3.6 | *
2016-06-26 | 178 | 1828 | TG | clr | * | -31/-431 | 6.6/3.6 | *
2016-06-27 | 179 | 1603 | EB/TG | clr | * | 23/-539 | 6.6/3.6 | * |  | *
2016-06-28 | 180 | 1452 | EB/TG | clr | * | -106/207 | 6.6/3.6 | *
2016-06-29 | 181 | 1135 | EB/TG | clr | * | -9/521 | 6.6/3.6 | *
2016-06-30 | 182 | 1143 | TG/EB | clr | * | -21/-118 | 6.6/3.6 | *
2016-07-01 | 183 | 1202 | TG/EB | clr | * | -72/-1570 | 6.7/3.6 | * |  |  | 1231z turned off both aethelometers in prep for outside TAWO construction
2016-07-02 | 184 | 1239 | TG/EB | clr | - | - | - | - |  |  | Aethelometers remained off for TAWO construction 22:01 Turned both back on and agnes pump back on
2016-07-03 | 185 | 1835 | EB | clr | * | 42/102 | 6.8/3.6 | *
2016-07-04 | 186 | 1029 | TG | rm frost | * | -47/176 | 6.7/3.6 | * |  |  | Turned off at 1036z due to constructin on TAWO tower instruments and pumps turned back on 14:03z
2016-07-05 | 187 | 1207 | TG/EB | clr | * | 24/-92 | 6.3/3.6 | * |  | *
2016-07-06 | 188 | 0.47153 | TG | clr | * | -27/-47 | 6.3/3.6 | *
2016-07-07 | 189 | 1214 | TG/EB | clr | * | 29/-99 | 6.3/3.6 | *
2016-07-08 | 190 | 1604 | TG/EB | clr | - | - | - | - |  |  | Instrument has disappeared. We suspect Any and Christy in th e Library with a candlestick
2016-07-09 | 191 | 1342 | TG | clr | - | - | - | - |  |  | Instrument has disappeared. We suspect CMDR in the vestibule with a knot
2016-07-10 | 192 | 1319 | TG | - | - | 404 | 3.6 | - |  |  | AE16 is now officially offline. No more daily checks for it…
2016-07-11 | 193 | 1323 | EB | clr | - | 202 | 3.6 | -
2016-07-12 | 194 | 1228 | TG/EB | - | - | 957 | 3.6 | -
2016-07-13 | 195 | 1137 | TG | - | - | 709 | 3.6 | -
2016-07-14 | 196 | 1147 | TG/EB | - | - | -409 | 3.6 | -
2016-07-15 | 197 | 1235 | TG/EB | - | - | 351 | 3.6 | -
2016-07-16 | 198 | 1301 | TG/EB | - | - | 233 | 3.6 | - |  |  | HEPA filter installed afternoon for overnight filtered air test.
2016-07-17 | 199 | 1627 | EB | - | - | 165 | 3.6 | -
2016-07-18 | 200 | 1207 | TG | - | - | 558 | 3.6 | - |  |  | Instrument left to continue overnight filter test in and out of N winds, but construction has continued on station. HEPA filter removed and external air intake restored at 1844z. power outage 2048-54z
2016-07-19 | 201 | 1237 | TG | - | - | 345 | 3.6 | -
2016-07-20 | 202 | 1609 | TG/EB | - | - | 334 | 3.6 | -
2016-07-21 | 203 | 1224 | TG/EB | - | - | 355 | 3.6 | -
2016-07-22 | 204 | 1141 | TG | - | - | -630 | 3.6 | -
2016-07-23 | 205 | 1317 | TG | - | - | -632 | 3.6 | -
2016-07-24 | 206 | 1747 | TG | - | - | 633 | 3.6 | -
2016-07-25 | 207 | 1234 | TG/EB | - | - | 234 | 3.6 | -
2016-07-26 | 208 | 1214 | TG/EB | - | - | 1133 | 3.6 | -
2016-07-27 | 209 | 0 | 0 | - | - |  |  | -
2016-07-28 | 210 | 1313 | TG/EB | - | - | 40 | 3.6 | -
2016-07-29 | 211 | 1611 | TG/EB | - | - | 243 | 3.6 | -
2016-07-30 | 212 | 1236 | TG/EB | - | - | 822 | 3.6 | -
2016-07-31 | 213 | 1521 | EB | - | - | 2064 | 3.6 | -
2016-08-01 | 214 | 1254 | EB | - | - | 1737 | 3.5 | -
2016-08-02 | 215 | 1259 | EB | - | - | 589 | 3.5 | -
2016-08-03 | 216 | 1121 | TG | - | - | 1955 | 3.6 | -
2016-08-04 | 217 | 1400 | EB | - | - | 183 | 3.6 | -
2016-08-05 | 218 | 1353 | EB/TG | - | - | -353 | 3.6 | -
2016-08-06 | 219 | 1717 | EB | - | - | 688 | 3.6 | -
2016-08-07 | 220 | 1632 | TG | - | - | 603 | 3.6 | -
2016-08-08 | 221 | 1541 | TG | - | - | -689 | 3.6 | -
2016-08-09 | 222 | 1832 | TG | - | - | 1011 | 3.6 | -
2016-08-10 | 223 | 1400 | EB | - | - | -1080 | 3.6 | -
2016-08-11 | 224 | 1243 | EB | - | - | 184 | 3.6 | -
2016-08-12 | 225 | 1534 | TG | - | - | -1366 | 3.5 | - |  |  | Power outage 1609-1611z. 1642-1644z
2016-08-13 | 226 | 1755 | EB | - | - | 48 | 3.5 | -
2016-08-14 | 227 | 1759 | EB | - | - | 243 | 3.5 | -
2016-08-15 | 228 | 1907 | TG/EB | - | - | -378 | 3.5 | -
2016-08-16 | 229 | 1810 | TG | - | - | 11 | 3.5 | - |  | *
2016-08-17 | 230 | 1652 | TG | - | - | 11 | 3.5 | -
2016-08-18 | 231 | 1731 | TG | - | - | 920 | 3.5 | -
2016-08-19 | 232 | 1612 | EB/TG | - | - | 118 | 3.6 | -
2016-08-20 | 233 | 1222 | TG | - | - | 192 | 3.6 | -
2016-08-21 | 234 | 1545 | EB | - | - | 454 | 3.6 | -
2016-08-22 | 235 | 1634 | TG | - | - | 236 | 3.6 | - |  | *
2016-08-23 | 236 | 1255 | EB | - | - | 1148 | 3.6 | -
2016-08-24 | 237 | 1330 | TG | - | - | 508 | 3.6 | -
2016-08-25 | 238 | 1223 | EB | - | - | 57 | 3.6 | -
2016-08-26 | 239 | 1352 | TG | - | - | 301 | 3.6 | -
2016-08-27 | 240 | 1744 | EB | - | - | 1211 | 3.6 | -
2016-08-28 | 241 | 1715 | TG | - | - | 435 | 3.6 | -
2016-08-29 | 242 | 1354 | EB | - | - | -210 | 3.5 | -
2016-08-30 | 243 | 1300 | TG | - | - | -137 | 3.5 | -
2016-08-31 | 244 | 1306 | EB | - | - | 164 | 3.5 | -
2016-09-01 | 245 | 1316 | TG | - | - | -23 | 3.5 | - |  |  | status warning: 4 flow low/high by more than 0.5 lpm contacted Betsy
2016-09-02 | 246 | 1700 | EB | - | - | -749 | 3.5 | -
2016-09-03 | 247 | 1323 | TG | - | - | -875 | 3.5 | -
2016-09-04 | 248 | 1419 | EB | rm frost | - | -538 | 3.5 | -
2016-09-05 | 249 | 1250 | TG | rm frost | - | -477 | 3.5 | - |  | *
2016-09-06 | 250 | 1342 | EB | clr | - | 470 | 3.5 | -
2016-09-07 | 251 | 1243 | TG | rm frost | - | -346 | 3.5 | - |  |  | 1325z stopped AE33 adjusted flow from 4 lpm to 5lpm 1326z re-started Aeee with new settings
2016-09-08 | 252 | 1237 | EB | clr | - | -140 | 3.5 | -
2016-09-09 | 253 | 1249 | TG | rm hv frost | - | 1052 | 3.5 | -
2016-09-10 | 254 | 1233 | EB | clr | - | 88 | 3.5 | -
2016-09-11 | 255 | 1742 | TG | rm frost | - | -343 | 3.5 | -
2016-09-12 | 256 | 1237 | EB | clr | - | 186 | 3.5 | -
2016-09-13 | 257 | 1237 | TG | rm hv frost | - | -711 | 3.5 | -
2016-09-14 | 258 | 1246 | EB | rm lt frost | - | -1434 | 3.5 | -
2016-09-15 | 259 | 1220 | TG | rm frost | - | -201 | 3.5 | -
2016-09-16 | 260 | 1340 | EB | rm frost | - | -219 | 3.5 | -
2016-09-17 | 261 | 1247 | TG | rm frost | - | 191 | 3.4 | -
2016-09-18 | 262 | 1457 | EB | rm frost | - | -162 | 3.5 | -
2016-09-19 | 263 | 1323 | TG | rm frost | - | 6 | 3.5 | - |  | *
2016-09-20 | 264 | 1324 | EB | rm frost | - | 608 | 3.5 | -
2016-09-21 | 265 | 1226 | TG | rm frost | - | -600 | 3.5 | -
2016-09-22 | 266 | 1318 | EB | rm frost | - | 8 | 3.5 | -
2016-09-23 | 267 | 1216 | TG | rm lt frost | - | 363 | 3.5 | -
2016-09-24 | 268 | 1409 | EB | clr | - | -321 | 3.5 | -
2016-09-25 | 269 | 1337 | TG | clr | - | 151 | 3.5 | -
2016-09-26 | 270 | 1257 | EB | rm frost | - | 146 | 3.5 | -
2016-09-27 | 271 | 1207 | TG | rm frost | - | -15 | 3.5 | - |  | *
2016-09-28 | 272 | 1325 | EB | rm frost | - | 265 | 3.5 | -
2016-09-29 | 273 | 1247 | TG | rm lt frost | - | -337 | 3.5 | -
2016-09-30 | 274 | 1402 | EB | clr | - | 284 | 3.5 | -
2016-10-01 | 275 | 1323 | TG | clr | - | 464 | 3.4 | -
2016-10-02 | 276 | 1403 | EB | clr | - | -274 | 3.5 | -
2016-10-03 | 277 | 1510 | TG | clr | - | 734 | 3.5 | - |  | *
2016-10-04 | 278 | 1219 | EB | rm lt frost | - | 159 | 3.5 | -
2016-10-05 | 279 | 1408 | TG | rm hv frost | - | -470 | 3.5 | -
2016-10-06 | 280 | 1314 | EB | clr | - | 120 | 3.5 | -
2016-10-07 | 281 | 1303 | TG | rm windblown snow | - | -427 | 3.5 | -
2016-10-08 | 282 | 0 | 0 | 0 | - |  |  | - |  |  | High winds/low visibility, didn't make it out to TAWO
2016-10-09 | 283 | 1459 | TG | rm frost | - | -5 | 3.5 | -
2016-10-10 | 284 | 1644 | EB | lt frost | - | 85 | 3.5 | -
2016-10-11 | 285 | 1338 | TG | rm frost | - | 282 | 3.5 | - |  | *
2016-10-12 | 286 | 0 | 0 | 0 | - |  |  | -
2016-10-13 | 287 | 1303 | TG | clr | - | 464 | 3.4 | -
2016-10-14 | 288 | 1716 | EB | clr | - | 400 | 3.5 | -
2016-10-15 | 289 | 1300 | TG | clr | - | -222 | 3.5 | -
2016-10-16 | 290 | 0 | 0 | 0 | - |  |  | -
2016-10-17 | 291 | 1431 | TG | rm frost | - | 701 | 3.5 | - |  | *
2016-10-18 | 292 | 1259 | EB | rm frost | - | 753 | 3.5 | -
2016-10-19 | 293 | 1257 | TG | rm lt frost | - | 31 | 3.4 | -
2016-10-20 | 294 | 0 | 0 | 0 | - |  |  | - |  |  | Con 1. Didn't go to TAWO.
2016-10-21 | 295 | 1619 | TG | rm lt frost | - | -18 | 3.4 | -
2016-10-22 | 296 | 1213 | EB | rm frost | - | -169 | 3.5 | -
2016-10-23 | 297 | 1331 | TG | rm lt frost | - | -208 | 3.5 | -
2016-10-24 | 298 | 1544 | TG/MG/NB | rm frost | - | 202 | 3.5 | -
2016-10-25 | 299 | 1133 | MG/TG | lt frost | - | 707 | 3.4 | - |  | *
2016-10-26 | 300 | 1213 | TG | rm frost | - | 711 | 3.4 | - |  |  | AE33 stopped during computer change out. Restarted data collection at 1339z
2016-10-27 | 301 | 1128 | MG/NB | rm frost | - | -303 | 3.5 | - |  |  | AE33 stopped during computer reboot. (Computer found frozen, rebooted 1126z) Restarted at 1141z.
2016-10-28 | 302 | 1135 | MG/NB | rm frost | - | -500 | 3.5 | -
2016-10-29 | 303 | 1150 | MG/NB | clr | - | -425 | 3.5 | -
2016-10-30 | 304 | 1133 | NB | clr | - | -181 | 3.5 | -
2016-10-31 | 305 | 1318 | NB | rm frost | - | -701 | 3.4 | -
2016-11-01 | 306 | 1253 | MG | rm lt frost | - | 364 | 3.5 | -
2016-11-02 | 307 | 1251 | NB | clr | - | 325 | 3.5 | -
2016-11-03 | 308 | 1245 | MG | rm frost | - | 382 | 3.5 | - |  | *
2016-11-04 | 309 | 1309 | MG | rm frost | - | -83 | 3.4 | -
2016-11-05 | 310 | 1230 | MG | rm lt frost | - | -230 | 3.5 | -
2016-11-06 | 311 | 1325 | NB | clr | - | 716 | 3.4 | - |  |  | Main aerosol pump (Colgan) found shut down upon arrival at the TAWO today. Moved sample line over to spare pump (Dignan) and reestablished sample flow at 1338z.
2016-11-07 | 312 | 1230 | MG | rm frost | - | 777 | 3.3 | -
2016-11-08 | 313 | 1225 | MG | rm lt frost | - | 52 | 3.3 | -
2016-11-09 | 314 | 1216 | MG | clr | - | 183 | 3.4 | -
2016-11-10 | 315 | 1242 | MG | clr | - | 262 | 3.4 | -
2016-11-11 | 316 |  |  |  | - |  |  | - |  |  | Did not go to TAWO condition 1
2016-11-12 | 317 | 1240 | MG/NB | clr | - | 283 | 3.5 | -
2016-11-13 | 318 | 1250 | MG | rm lt frost | - | -80 | 3.2 | -
2016-11-14 | 319 | 1900 | NB | rm frost | - | 927 | 3.3 | -
2016-11-15 | 320 | 1400 | MG | rm lt frost | - | 155 | 3.2 | -
2016-11-16 | 321 | 1302 | MG | rm lt frost | - | -393 | 3.3 | - |  | * | Wiped old data, 153 days of space left
2016-11-17 | 322 | 1218 | MG/NB | rm lt frost | - | 262 | 3.4 | -
2016-11-18 | 323 | 1323 | MG | rm frost | - | -40 | 3.4 | -
2016-11-19 | 324 | 1235 | NB | rm lt frost | - | 139 | 3.3 | -
2016-11-20 | 325 | 1312 | NB | rm frost | - | -300 | 3.3 | -
2016-11-21 | 326 | 1305 | NB | rm frost | - | 262 | 3.3 | -
2016-11-22 | 327 | 1301 | MG/NB | rm frost | - | 444 | 3.3 | -
2016-11-23 | 328 | 1230 | NB | lt frost | - | -260 | 3.3 | -
2016-11-24 | 329 | 1250 | NB | rm frost | - | -100 | 3.3 | -
2016-11-25 | 330 | 1249 | MG | rm frost | - | 1214 | 3.3 | -
2016-11-26 | 331 | 1335 | NB | clr | - | 20 | 3.4 | -
2016-11-27 | 332 | 1420 | MG | clr | - | 1707 | 3.4 | -
2016-11-28 | 333 | 1235 | NB | rm frost | - | -553 | 3.4 | -
2016-11-29 | 334 | 1245 | NB | rm frost | - | 330 | 3.4 | -
2016-11-30 | 335 | 1325 | MG | rm lt frost | - | 616 | 3.4 | -
2016-12-01 | 336 | 1250 | NB | rm frost | - | 997 | 3.3 | -
2016-12-02 | 337 | 1321 | MG | rm frost | - | 2215 | 3.3 | - |  | *
2016-12-03 | 338 | 1235 | NB | rm frost | - | -161 | 3.2 | -
2016-12-04 | 339 | 1348 | NB | rm frost | - | 535 | 3.3 | -
2016-12-05 | 340 | 1242 | NB | clr | - | 601 | 3.4 | -
2016-12-06 | 341 | 1244 | MG | rm frost | - | 496 | 3.4 | -
2016-12-07 | 342 | 1240 | MG | clr | - | -115 | 3.3 | -
2016-12-08 | 343 | 1234 | MG/NB | clr | - | -289 | 3.4 | -
2016-12-09 | 344 | 1328 | MG | rm frost | - | -702 | 3.3 | -
2016-12-10 | 345 | 1920 | MG | rm frost | - | 307 | 3.2 | -
2016-12-11 | 346 | 1340 | MG | rm lt frost | - | 669 | 3.4 | -
2016-12-12 | 347 | 1230 | NB | rm lt frost | - | 540 | 3.3 | -
2016-12-13 | 348 | 1225 | NB | rm lt frost | - | -332 | 3.3 | -
2016-12-14 | 349 | 1304 | MG | rm lt frost | - | 642 | 3.4 | -
2016-12-15 | 350 | 1415 | MG | rm lt frost | - | 1215 | 3.4 | - |  | *
2016-12-16 | 351 | 1308 | MG | rm lt frost | - | -121 | 3.3 | -
2016-12-17 | 352 | 1252 | MG | rm lt frost | - | 1226 | 3.2 | -
2016-12-18 | 353 | 1328 | NB | clr | - | 475 | 3.3 | -
2016-12-19 | 354 | 1220 | NB | clr | - | -158 | 3.3 | -
2016-12-20 | 355 | 1506 | MG | rm lt frost | - | -570 | 3.2 | -
2016-12-21 | 356 | 1325 | MG | rm lt frost | - | 1174 | 3 | -
2016-12-22 | 357 | 1350 | MG | rm lt frost | - | -480 | 3.3 | -
2016-12-23 | 358 | 1244 | NB | clr | - | 237 | 3.4 | - |  |  | Power outages for e-gen swap: 1330-1333z and 1645-1648z.
2016-12-24 | 359 | 1326 | NB | clr | - | 591 | 3.4 | -
2016-12-25 | 360 | 1456 | MG | rm lt frost | - | -444 | 3.4 | -
2016-12-26 | 361 | 1235 | NB | rm lt frost | - | -233 | 3.4 | -
2016-12-27 | 362 | 1034 | MG | - | - | 148 | 3.4 | -
2016-12-28 | 363 | 1358 | MG | clr | - | -181 | 3.3 | - |  | *
2016-12-29 | 364 | 1308 | MG/NB | clr | - | 673 | 3.3 | -
2016-12-30 | 365 | 1314 | MG/NB | clr | - | 138 | 3.4 | -
2016-12-31 | 366 | 1332 | NB | clr | - | -69 | 3.5 | -

## MET Tower
Date | Julian Day | Time (UTC) | Observer's Initials | 2m Temp | 2m TD | 9m SONIC | CATS GC (inlet tower) | Sfc O3 (inlet tower) | CU GC | 10m Windbird | 10m temp | 16m Windbird | 16m temp | Bunting Pressure (+/-0.2) | Offset Zeno Time (reset if > 3sec) | CR1000 Connection | Comments (Climbing, removing instruments, etc)
2016-01-01 | 1 | 1510 | JG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 651.1/651.8 | 1 | *
2016-01-02 | 2 | 1229 | NB | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | 657.0/657.6 | 3 | * | NB climbed tower and removed frost 1404-1450z. Cleaned snow from 10m and 16m RTD fans and reset plugs into N outlet but 16m fan still not running.
2016-01-03 | 3 | 1310 | JG | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 664.7/665.4 | 4 | * | Reset zeno clock at 1318z.
2016-01-04 | 4 | 1228 | NB | rm frost | rm frost | clr | lt frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | 669.2/669.9 | 0 | * | JG climbed tower 1355-1417z and removed frost from instruments.
2016-01-05 | 5 | 1230 | JG | clr | clr | clr | lt frost | clr | clr | clr | clr | clr | clr | 670.5/671.2 | 1 | * | Network time program unable to connect to server.
2016-01-06 | 6 | 1307 | NB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 669.3/670.2 | 3 | * | Network time program has resumed time sync to server.
2016-01-07 | 7 | 1235 | JG | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 671.0/671.8 | 5 | * | Reset zeno clock at 1246z.
2016-01-08 | 8 | 1347 | NB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 667.1/668.0 | 0 | *
2016-01-09 | 9 | 1240 | JG | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 665.4/666.2 | 1 | *
2016-01-10 | 10 | 1504 | NB | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | 663.6/664.4 | 3 | * | JG climbed tower 1513-1530z and removed frost from instruments.
2016-01-11 | 11 | 1240 | JG | rm frost | rm frost | clr | lt frost | lt frost | lt frost | clr | lt frost | clr | lt frost | 668.7/669.5 | 4 | * | Reset zeno clock at 1249z.
2016-01-12 | 12 | 1232 | NB | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | 681.3/682.1 | 0 | * | JG climbed tower 1243-1307z and removed frost from instruments.
2016-01-13 | 13 | 1255 | JG | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 673.8/674.3 | 2 | * | Sonic reading lower than 10m RM Young - more blocked by bldg in N winds. Planned power outage for e-gen test 1307-1309 and 1321-1324 UTC.
2016-01-14 | 14 | 1225 | NB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 671.6/672.2 | 3 | *
2016-01-15 | 15 | 1255 | JG | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | 671.5/671.7 | 5 | * | Reset zeno clock at 1302z. JG climbed tower 1327-1347 and cleaned frost from instruments.
2016-01-16 | 16 | 1304 | NB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 675.2/675.4 | 1 | *
2016-01-17 | 17 | 1725 | JG | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 672.2/672.4 | 2 | *
2016-01-18 | 18 | 1229 | NB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 670.1/670.2 | 4 | * | Reset zeno clock at 1239z.
2016-01-19 | 19 | 1310 | JG | clr | lt frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 667.5/667.8 | 0 | *
2016-01-20 | 20 | 1232 | NB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 662.9/663.0 | 1 | *
2016-01-21 | 21 | 1250 | JG | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 660.8/661.1 | 3 | *
2016-01-22 | 22 | 1234 | NB | clr | clr | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 654.8/654.8 | 5 | * | Reset zeno clock at 1241z.
2016-01-23 | 23 | 1210 | JG | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | 648.4/648.7 | 0 | * | NB climbed tower and removed frost from instruments 1217-1234z.
2016-01-24 | 24 | 1258 | NB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 649.6/649.7 | 2 | *
2016-01-25 | 25 | 1230 | JG | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 642.1/642.4 | 4 | * | Reset zeno clock at 1238z.
2016-01-26 | 26 | 1254 | NB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 643.2/643.5 | 1 | *
2016-01-27 | 27 | 1340 | JG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 647.1/647.3 | 3 | *
2016-01-28 | 28 | 1234 | NB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 648.9/649.2 | 5 | * | Reset zeno clock at 1306z.
2016-01-29 | 29 | 1255 | JG | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 648.7/648.9 | 0 | *
2016-01-30 | 30 | 1225 | NB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 650.4/650.6 | 2 | *
2016-01-31 | 31 | 1455 | JG | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 653.3/653.4 | 4 | * | Reset zeno clock at 1502z.
2016-02-01 | 32 | 1225 | NB | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | 659.4/659.6 | 0 | * | JG climbed tower 1310-1327z and removed frost from instruments.
2016-02-02 | 33 | 1240 | JG | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 664.4/664.6 | 1 | *
2016-02-03 | 34 | 1248 | NB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 663.4/663.6 | 4 | * | Reset zeno clock at 1302z.
2016-02-04 | 35 | 1220 | JG | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | 664.4/664.6 | 0 | * | NB climbed tower 1224-1239z and removed frost from instruments. Bunting was restarted around 1645z.
2016-02-05 | 36 | 1333 | NB | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | 659.9/660.0 | 2 | * | JG climbed tower 1342-1359z and removed frost from instruments.
2016-02-06 | 37 | 1355 | JG | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | 654.3/654.5 | 4 | * | NB climbed tower 1407-1421z and removed frost from instruments. Reset zeno clock at 1442z.
2016-02-07 | 38 | 1455 | MB/JJ | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 653.4/653.5 | 0 | *
2016-02-08 | 39 | 1240 | MB/JJ/SD | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 655.0/655.1 | 2 | *
2016-02-09 | 40 | 1230 | MB/JJ/NB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 654.6/654.7 | 3 | *
2016-02-10 | 41 | 1226 | SWD/JG/NB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 652.0/652.1 | 5 | * | Reset zeno clock at 1239z.
2016-02-11 | 42 | 1340 | MB/JJ | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 653.7/653.8 | 3 | *
2016-02-12 | 43 | 1625 | MB/JJ | rm frost | rm frost | clr | lt frost | lt frost | lt frost | frost | lt frost | lt frost | lt frost | 656.6/656.7 | 4 | * | 10m windbird reading 2.7m/s slower than sonic. Light frost persists on the windbirds slowing them down. Conditions not suitable for climbing. Reset Zeno Clock at 16:41UTC
2016-02-13 | 44 | 1220 | MB/JJ | rm frost | rm frost | clr | lt frost | lt frost | lt frost | frost | lt frost | lt frost | lt frost | 661.3/661.6 | 0 | * | 10m windbird reading 3.0m/s slower than sonic. Light frost persists on the windbirds slowing them down. Conditions not suitable for climbing.
2016-02-14 | 45 | 1700 | MB/JJ | rm frost | rm frost | clr | lt frost | lt frost | lt frost | frost | lt frost | frost | lt frost | 662.6/662.8 | 2 | * | 10 m windbird reading 4.5 m/s slower than sonic.
2016-02-15 | 46 | 1641 | JJ | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 656.5/656.7 | 4 | * | 10 m windbird reading 1.4 m/s slower than sonic. Reset Zeno Clock at 16:55 Z.
2016-02-16 | 47 | 1321 | jj/mb | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 650.4/650.6 | -4 | *
2016-02-17 | 48 | 1415 | MB/JJ | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | 656.8/656.9 | -2 | * | JJ climbed tower 1428-1455z and removed frost from instruments. 1848-1904z disconnected 2m temp sensor #235 and replaced with sensor #281. Incorrect temp values persist.
2016-02-18 | 49 | 1405 | MB/JJ | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 649.7/650.2 | -1 | * | Troubleshot Zeno 2m RTD wiring 1454-1615z.
2016-02-19 | 50 | 1349 | JJ/MB | rm frost | rm frost | clr | clr | clr | clr | clr | clr | clr | clr | 647.9/648.4 | -16 | * | Reset zeno clock at 1407z. Troubleshot Zeno 2m RTD 1650-0045z, used 10m RTD wiring as part of troubleshooting process.
2016-02-20 | 51 | 1311 | JJ/MB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 652.1/655.1 | -1 | * | JJ climbed tower 1332-1358z and 1457-1526z to troubleshoot 2m RTD cable. Troubleshot Zeno 2m RTD wiring 1534-1543z.
2016-02-21 | 52 | 1510 | JJ/MB | rm frost | rm frost | clr | frost | frost | frost | frost | frost | frost | frost | 661.4/661.6 | 1 | *
2016-02-22 | 53 | 1244 | JJ/MB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 667.1/667.3 | 2 | *
2016-02-23 | 54 | 1704 | MB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 672.7/673.0 | 4 | * | Reset zeno clock at 1714z.
2016-02-24 | 55 | 1314 | JJ/MB | clr | clr | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 671.3/671.5 | 0 | *
2016-02-25 | 56 | 1336 | jj/mb | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 670.7/671.0 | 2 | *
2016-02-26 | 57 | 1348 | JJ/MB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 672.0/672.3 | 3 | *
2016-02-27 | 58 | 1305 | MB | rm frost | rm frost | clr | hv frost | hv frost | hv frost | hv frost | hv frost | hv frost | hv frost | 669.3/669.5 | 5 | * | Reset zeno clock at 1314z. Windbirds siezed with heavy frost. Sonic reading 4.1 m/s. SWD climbed tower to clean instruments and troubleshoot 2m-RTD cable from 1444-1542z.
2016-02-28 | 59 | 1708 | JJ | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 663.5/663.7 | 2 | *
2016-02-29 | 60 | 1339 | jj/mb | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 658.7/658.7 | 3 | * | Removed 2m RTD cable leads from Zeno to check resistances 1442-1456z.
2016-03-01 | 61 | 1248 | JJ/MB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 656.5/658.0 | 5 | * | Reset zeno clock at 1255z.
2016-03-02 | 62 | 1711 | MB/JJ | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | 659.3/660.9 | 0 | * | JJ climbed tower to clean instruments and wrap 2m-RTD cable from 1737-1833z.
2016-03-03 | 63 | 1338 | MB/JJ | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 663.3/664.9 | 0 | *
2016-03-04 | 64 | 1309 | jj/mb | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 667.5/668.4 | 3 | *
2016-03-05 | 65 | 1235 | jj/mb | rm frost | rm frost | clr | lt frost | lt frost | lt frost | clr | lt frost | clr | lt frost | 663.1/663.7 | -2 | * | reset zeno clock around 0030 UTC
2016-03-06 | 66 | 1342 | MB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | clr | lt frost | clr | lt frost | 654.9/655.5 | -1 | *
2016-03-07 | 67 | 1321 | jj/mb | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 647.4/648.5 | -2 | *
2016-03-08 | 68 | 1331 | jj/mb | rm frost | rm frost | clr | lt frost | lt frost | lt frost | clr | lt frost | clr | lt frost | 655.0/655.3 | -3 | *
2016-03-09 | 69 | 1235 | JJ | clr | clr | clr | lt frost | lt frost | lt frost | clr | lt frost | clr | lt frost | 660.7/661.1 | -2 | *
2016-03-10 | 70 | 0 | 0 | [13 empty] | White out conditions (Condition 1) did not travel to TAWO this day.
2016-03-11 | 71 | 1352 | JJ | clr | clr | clr | clr | clr | clr | clr | clr | clr | lt frost | 636.7/636.9 | -2 | *
2016-03-12 | 72 | 1437 | MB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | clr | lt frost | clr | lt frost | 647.9/648.0 | 10 | * | Reset Zeno clock at 1445z.
2016-03-13 | 73 | 1346 | JJ | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 652.5/652.8 | -2 | * | unexpected power outage today from 13:16 to 13:23 UTC
2016-03-14 | 74 | 1353 | jj/mb | clr | clr | clr | lt frost | lt frost | lt frost | clr | clr | clr | lt frost | 644.2/644.7 | -2 | *
2016-03-15 | 75 | 1533 | jj/mb | clr | clr | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 665.0/665.3 | -1 | *
2016-03-16 | 76 | 1319 | MB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 670.7/670.9 | 4 | * | Reset Zeno clock at 1330z.
2016-03-17 | 77 | 1338 | jj/mb | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 675.6/675.9 | 1 | *
2016-03-18 | 78 | 1305 | MB/JJ | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 673.5/673.8 | 2 | *
2016-03-19 | 79 | 1854 | MB/JJ | rm frost | rm frost | clr | lt frost | lt frost | lt frost | clr | lt frost | lt frost | lt frost | 666.8/667.0 | 4 | * | Reset Zeno clock at 1859z.
2016-03-20 | 80 | 1334 | MB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 664.0/664.2 | -1 | *
2016-03-21 | 81 | 1307 | jj/mb | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 663.9/664.2 | 1 | *
2016-03-22 | 82 | 1320 | MB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 665.2/665.5 | 2 | *
2016-03-23 | 83 | 1744 | MB/JJ | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 659.0/659.0 | 5 | * | Reset Zeno clock at 1752z.
2016-03-24 | 84 | 1329 | jj/mb | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 653.6/653.9 | 1 | *
2016-03-25 | 85 | 1303 | JJ | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 654.2/654.4 | -1 | *
2016-03-26 | 86 | 1303 | MB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 657.6/657.9 | 2 | *
2016-03-27 | 87 | 1715 | JJ | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 656.8/656.9 | 5 | * | Reset Zeno clock at 1741z.
2016-03-28 | 88 | 1327 | MB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 656.2/656.4 | -1 | *
2016-03-29 | 89 | 1151 | JJ | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 658.3/658.7 | 0 | *
2016-03-30 | 90 | 1134 | MB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 658.7/659.0 | 2 | *
2016-03-31 | 91 | 1718 | jj/mb | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 659.1/659.0 | 4 | * | Reset Zeno clock at 1743z.
2016-04-01 | 92 | 1304 | JJ | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 660.1/660.4 | 0 | * | Reset Zeno clock at 2330z
2016-04-02 | 93 | 1213 | jj/mb | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 663.3/663.6 | -2 | *
2016-04-03 | 94 | 1550 | MB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | clr | lt frost | clr | lt frost | 668.9/669.1 | 0 | *
2016-04-04 | 95 | 1613 | jj/mb | rm frost | rm frost | clr | lt frost | lt frost | lt frost | clr | lt frost | clr | lt frost | 670.9/671.1 | 2 | *
2016-04-05 | 96 | 1251 | MB/JJ | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 665.7/665.8 | 3 | *
2016-04-06 | 97 | 1332 | JJ | rm frost | rm frost | clr | lt frost | lt frost | lt frost | clr | lt frost | clr | lt frost | 667.8/667.9 | 5 | * | Reset Zeno clock at 1345z Planned power outage today
2016-04-07 | 98 | 1224 | MB/JJ | rm frost | rm frost | clr | clr | clr | clr | clr | clr | clr | clr | 668.6/668.9 | -1 | *
2016-04-08 | 99 | 1139 | JJ/MB | rm frost | rm frost | clr | clr | lt frost | lt frost | clr | lt frost | clr | lt frost | 672.5/672.9 | 1 | * | SWD climbed the tower from 1715-1800z to measure 10m and 16m windbird alignment.
2016-04-09 | 100 | 1254 | JJ/MB | clr | clr | clr | clr | clr | lt frost | clr | clr | clr | clr | 681.1/681.3 | 2 | *
2016-04-10 | 101 | 1707 | JJ | clr | clr | clr | clr | clr | lt frost | clr | clr | clr | clr | 682.6/682.9 | 4 | * | Reset Zeno clock at 1726z
2016-04-11 | 102 | 1321 | JJ/MB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 684.3/684.4 | -3 | *
2016-04-12 | 103 | 1554 | JJ/MB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 687.6/687.5 | -1 | *
2016-04-13 | 104 | 1317 | JJ/MB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 680.2/680.4 | 1 | *
2016-04-14 | 105 | 1329 | JJ/MB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 678.6/679.0 | 2 | * | JJ climbed tower from 1356-1422z to work on surface ozone inlet and clean instruments.
2016-04-15 | 106 | 1711 | JJ/MB/SWD | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 669.0/669.3 | 5 | failure | Reset Zeno clock at 1733z. CR1000 communication failure from 1630-1740z.Communications restored after unplugging and replugging the USB-Serial connector from Met computer.
2016-04-16 | 107 | 1309 | JJ/MB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 662.1/662.3 | -1 | *
2016-04-17 | 108 | 1628 | MB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 663.1/663.4 | 2 | *
2016-04-18 | 109 | 1249 | MB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 661.4/661.9 | 3 | *
2016-04-19 | 110 | 1152 | JJ/MB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 661.1/661.4 | 5 | * | Reset Zeno clock at 1203z
2016-04-20 | 111 | 1143 | JJ/MB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 666.0/666.3 | -2 | *
2016-04-21 | 112 | 1122 | JJ/MB/SWD | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 672.7/673.1 | 0 | * | 1426-1458z and 1513-1527z SWD climbed tower to clean instruments and realign 10m windbird.
2016-04-22 | 113 | 1838 | JJ/MB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 677.2/677.4 | 1 | failure | CR1000 communication failure from 1300-1830z.Communications restored after unplugging and replugging the USB-Serial connector from Met computer.
2016-04-23 | 114 | 1536 | 0 | [13 empty] | White out conditions (Condition 1) did not travel to TAWO this day.
2016-04-24 | 115 | 1536 | JJ/MB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 681.9/682.0 | 5 | * | Reset Zeno clock at 1559z
2016-04-25 | 116 | 1207 | JJ/MB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 687.6/687.6 | -2 | *
2016-04-26 | 117 | 1130 | JJ | rm frost | rm frost | clr | lt frost | lt frost | lt frost | clr | lt frost | clr | lt frost | 691.6/691.7 | 0 | *
2016-04-27 | 118 | 1549 | JJ/MB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | clr | lt frost | clr | lt frost | 692.3/692.3 | 2 | *
2016-04-28 | 119 | 1315 | JJ/MB | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | rm frost | clr | rm frost | 689.1/689.1 | 3 | * | 1255-1305z JJ climbed the tower to clean instruments.
2016-04-29 | 120 | 1300 | MB/JJ | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 680.3/680.4 | 5 | * | Reset Zeno clock at 1310z.
2016-04-30 | 121 | 1336 | MB/JJ | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 667.1/667.4 | 0 | *
2016-05-01 | 122 | 1624 | MB/JJ | clr | clr | clr | clr | rm frost | rm frost | clr | clr | clr | clr | 670.6/670.8 | 1 | * | 1743-1815z JJ and MB climbed tower to replace filter on CUGC sample inlet.
2016-05-02 | 123 | 1322 | JJ/MB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 671.5/671.7 | 3 | * | Planned power outages starting from 1635 to 1700z
2016-05-03 | 124 | 1230 | MB/JJ | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | 673.5/673.7 | 5 | * | 1153-1217z JJ climbed tower to clean instruments and attach ties to CUGC intake hose. Reset Zeno clock at 1233z.
2016-05-04 | 125 | 1451 | MB/JJ | rm frost | rm frost | clr | clr | clr | clr | clr | clr | clr | clr | 674.6/674.7 | 0 | *
2016-05-05 | 126 | 1528 | MB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 674.6/675.0 | 2 | *
2016-05-06 | 127 | 1845 | JJ/MB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 673.7/674.0 | 4 | * | Reset Zeno clock at 1856z.
2016-05-07 | 128 | 1315 | JJ/MB | rm frost | rm frost | clr | clr | clr | clr | clr | clr | clr | clr | 627.8/673.1 | -2 | *
2016-05-08 | 129 | 1558 | JJ | rm frost | rm frost | clr | lt frost | lt frost | lt frost | clr | lt frost | clr | lt frost | 676.6/677.0 | 0 | *
2016-05-09 | 130 | 1635 | JJ/MB | clr | rm frost | clr | clr | clr | clr | clr | clr | clr | clr | 678.2/678.3 | 1 | *
2016-05-10 | 131 | 1648 | JJ/MB | clr | rm frost | clr | clr | clr | clr | clr | clr | clr | clr | 668.2/668.5 | 3 | *
2016-05-11 | 132 | 1753 | MB/JJ | rm frost | rm frost | clr | clr | clr | clr | clr | lt frost | clr | lt frost | 681.6/681.8 | 4 | * | Reset Zeno clock at 1759z.
2016-05-12 | 133 | 1622 | MB/JJ | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 683.4/683.6 | 0 | *
2016-05-13 | 134 | 1308 | MB/JJ | clr | clr | clr | clr | clr | clr | clr | lt frost | clr | lt frost | 681.2/681.4 | 2 | *
2016-05-14 | 135 | 1149 | JJ/MB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 677.7/678.0 | 3 | *
2016-05-15 | 136 | 1730 | MB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 676.8/677.2 | 5 | failure | CR1000 communication failure from last night to 1730z today.Communications restored after unplugging and replugging the USB-Serial connector from Met computer. Reset Zeno clock at1743z.
2016-05-16 | 137 | 1609 | JJ/MB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | clr | lt frost | clr | lt frost | 678.6/678.9 | 0 | *
2016-05-17 | 138 | 1722 | JJ/MB | clr | rm frost | clr | clr | lt frost | lt frost | clr | lt frost | clr | lt frost | 671.8/672.0 | 2 | *
2016-05-18 | 139 | 1309 | JJ/MB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 666.7/667.1 | 3 | *
2016-05-19 | 140 | 1322 | JJ | rm frost | clr | clr | clr | clr | clr | clr | lt frost | clr | lt frost | 666.3/666.6 | 4 | * | Reset Zeno clock at 1338z.
2016-05-20 | 141 | 1850 | JJ/MB | clr | clr | clr | clr | clr | clr | clr | lt frost | clr | lt frost | 668.3/668.8 | -2 | *
2016-05-21 | 142 | 1228 | JJ/MB | rm frost | rm frost | clr | lt frost | lt frost | lt frost | clr | lt frost | lt frost | lt frost | 668.1/668.5 | 0 | *
2016-05-22 | 143 | 2327 | JJ | rm frost | rm frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 671.6/672.0 | 2 | *
2016-05-23 | 144 | 1656 | JJ | rm frost | rm frost | clr | clr | clr | clr | clr | clr | clr | clr | 673.0/673.1 | 4 | * | Reset Zeno clock at 1709z.
2016-05-24 | 145 | 1211 | JJ/MB | rm frost | rm frost | clr | clr | lt frost | lt frost | clr | lt frost | clr | lt frost | 673.2/673.6 | -2 | *
2016-05-25 | 146 | 1726 | JJ | clr | clr | clr | clr | clr | clr | lt frost | clr | lt frost | clr | 662.8/663.1 | 0 | * | Restarted Intrcept to get system time to match sample time.
2016-05-26 | 147 | 1232 | JJ/MB | rm frost | rm frost | clr | clr | clr | clr | clr | clr | clr | clr | 660.2/660.6 | 1 | *
2016-05-27 | 148 | 1845 | JJ/mb | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 663.7/663.9 | 3 | *
2016-05-28 | 149 | 1221 | MB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 670.6/670.9 | 6 | * | Reset Zeno clock at 1236z.
2016-05-29 | 150 | 1530 | MB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 677.4/677.8 | -1 | *
2016-05-30 | 151 | 1226 | JJ | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 677.4/677.7 | 2 | * | Planned power outage today from 2130 to 2230
2016-05-31 | 152 | 1731 | JJ | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 680.0/680.4 | 3 | *
2016-06-01 | 153 | 1840 | JJ | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 679.6/679.9 | 5 | * | Reset Zeno clock at 1849z.
2016-06-02 | 154 | 1831 | MB/JJ | rm frost | rm frost | clr | clr | clr | clr | clr | clr | clr | clr | 684.8/684.9 | -2 | *
2016-06-03 | 155 | 1737 | MB/JJ/HJ | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 685.6/685.8 | 0 | *
2016-06-04 | 156 | 1156 | MB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 682.4/682.5 | 1 | *
2016-06-05 | 157 | 1403 | JJ/TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 679.0/679.3 | 4 | * | Reset Zeno clock at 1646z.
2016-06-06 | 158 | 1613 | MB/JJ/TG/EB | rm frost | rm frost | clr | clr | clr | clr | clr | clr | clr | clr | 685.4/685.6 | -2 | *
2016-06-07 | 159 | 1536 | TG/EB/JJ | rm frost | rm frost | clr | clr | clr | clr | clr | clr | clr | clr | 687.01/687.1 | -1 | *
2016-06-08 | 160 | 1309 | TG/EB/JJ/MB | rm frost | rm frost | clr | clr | clr | clr | clr | clr | clr | clr | 686.9/687.1 | 1 | *
2016-06-09 | 161 | 1553 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 685.3/685.5 | 3 | *
2016-06-10 | 162 | 1239 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 678.0/678.2 | 4 | * | Work on ventilation duct from 1105 to 1330, higher room temps.
2016-06-11 | 163 | 1203 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 681.1/681.3 | 5 | * | Reset Zeno clock at 1309z; Closed Intercept and had to restart Bunting in order to reopen (Application Error: integer divide by 0); Bunting restarted, Intercept restored at 1309z; Installed new version of RJ Clockwise (old version lost upon restart).
2016-06-12 | 164 | 1509 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 680.7/680.9 | -2 | *
2016-06-13 | 165 | 1637 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 679.4/679.9 | 0 | *
2016-06-14 | 166 | 1153 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 678.7/670.1 | 1 | *
2016-06-15 | 167 | 1202 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 675.2/675.7 | 3 | *
2016-06-16 | 168 | 1156 | TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 677/677.2 | 4 | *
2016-06-17 | 169 | 1314 | TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 678/678.3 | 6 | * | Reset Zeno clock at 1326z
2016-06-18 | 170 | 1241 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 679.7/679.9 | -3 | *
2016-06-19 | 171 | 1808 | EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 675.1/675.3 | 0 | *
2016-06-20 | 172 | 1812 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 673.4/673.9 | 1 | *
2016-06-21 | 173 | 1204 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 674.7/675.1 | 3 | *
2016-06-22 | 174 | 1332 | TG/EB | rm frost | rm frost | rm frost | rm frost | clr | clr | rm frost | rm frost | rm frost | rm frost | 674.1/674.4 | 4 | *
2016-06-23 | 175 | 1313 | EB | rm frost | rm frost | clr | clr | clr | clr | clr | rm frost | clr | clr | 674.8/675.0 | 6 | * | Reset Zeno clock at 1331z
2016-06-24 | 176 | 1329 | TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 673.3/673.7 | -2 | *
2016-06-25 | 177 | 1717 | EB/TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 674.4/674.7 | 0 | *
2016-06-26 | 178 | 1828 | TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 677.4/677.8 | 1 | *
2016-06-27 | 179 | 1603 | EB/TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 678.6/679.0 | 3 | *
2016-06-28 | 180 | 1452 | EB/TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 679.9/686.7 | 5 |  | Unable to do the CR1000 connection. Perhaps someone is working on it right now as we had trouble getting into Bunting. Will check tomorrow
2016-06-29 | 181 | 1135 | EB/TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 677.8/692.3 | -2 |  | Unable to make CR1000 connection.
2016-06-30 | 182 | 1143 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 678.3/678.4 | 0 | * | CR1000 connection ocurred after a reboot of MET computer
2016-07-01 | 183 | 1202 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 680.9/680.9 | 2 | * | 2142 - Swapped HMP155 for testing with Swiss tower HMP155. 2219 - Returned original HMP155 to tower.
2016-07-02 | 184 | 1239 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 680.5/688.8 | 3 | *
2016-07-03 | 185 | 1835 | EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 681.9/688.7 | 5 | *
2016-07-04 | 186 | 1029 | TG | rm frost | rm frost | rm frost | clr | clr | clr | rm frost | rm frost | rm frost | rm frost | 681.4/682.3 | 6 | * | 2242 - Swapped HMP155 for testing with Swiss tower HMP155. Reset Zeno clock at 1108z
2016-07-05 | 187 | 1207 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 683.0/683.8 | -1 | *
2016-07-06 | 188 | 0.47153 | TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 683.5/684.1 | 0 | *
2016-07-07 | 189 | 1214 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 683.4/684.3 | 1 | *
2016-07-08 | 190 | 1604 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 683.3/683.1 | -20 | * | ResetZeno Clock at 1616z
2016-07-09 | 191 | 1342 | TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 685.8/685.5 | 0 | *
2016-07-10 | 192 | 1319 | TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 686.4/686.2 | 1 | *
2016-07-11 | 193 | 1323 | EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 687.4/687.6 | 3 |  | TAWO MET down for troubleshooting
2016-07-12 | 194 | 1228 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 686.5/686.7 | 5 | *
2016-07-13 | 195 | 1137 | TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 684.1/684.1 | 6 | * | Reset Zeno clock at 1247z
2016-07-14 | 196 | 1147 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 681.3/681.4 | 0 | *
2016-07-15 | 197 | 1235 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 679.8/679.8 | 2 | *
2016-07-16 | 198 | 1301 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 680.0/680.1 | 3 | *
2016-07-17 | 199 | 1627 | EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 679.4/679.8 | 5 | *
2016-07-18 | 200 | 1207 | TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 680.0/680.4 | 7 | * | restet zeno clock at 1220z in and out of N winds, but construction has continued on station power outage 2048-54z
2016-07-19 | 201 | 1237 | TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 682.2/682.5 | -2 | *
2016-07-20 | 202 | 1609 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 683.5/683.8 | 1 | *
2016-07-21 | 203 | 1224 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 686.1/686.3 | 2 | *
2016-07-22 | 204 | 1141 | TG | rm frost | rm frost | rm frost | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | 687.7/687.9 | 3 | *
2016-07-23 | 205 | 1317 | TG | rm frost | rm frost | clr | clr | clr | clr | clr | clr | clr | clr | 688.0/688.1 | -23 | * | reset zeno clock at 1330z. We were playing around on intercept and shut it down at **** for data recovery
2016-07-24 | 206 | 1747 | TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 689.6/689.8 | -2 | *
2016-07-25 | 207 | 1234 | TG/EB | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | clr | clr | 689.2/689.5 | -1 | *
2016-07-26 | 208 | 1214 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 689.6/689.9 | 1 | *
2016-07-27 | 209 | 0 | 0
2016-07-28 | 210 | 1313 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 688.4/388.7 | 4 | *
2016-07-29 | 211 | 1611 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 682.6/682.5 | 6 | *
2016-07-30 | 212 | 1236 | TG/EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 686.4/686.7 | 8 | * | reset zeno clock at 1244z
2016-07-31 | 213 | 1521 | EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 684.3/684.7 | -2 | *
2016-08-01 | 214 | 1254 | EB | rm frost | rm frost | clr | clr | clr | rm frost | clr | rm frost | clr | rm frost | 683.0/683.3 | 0 | *
2016-08-02 | 215 | 1259 | EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 682.4/682.7 | 2 | *
2016-08-03 | 216 | 1121 | TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 682.4/682.7 | 3 | *
2016-08-04 | 217 | 1400 | EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 682.9/683.1 | 5 | *
2016-08-05 | 218 | 1353 | EB/TG | rm frost | rm frost | clr | clr | rm frost | clr | clr | rm frost | clr | rm frost | 682.7/682.8 | 6 | * | reset zeno clock at 1401z
2016-08-06 | 219 | 1717 | EB | rm frost | rm frost | clr | clr | rm frost | clr | clr | rm frost | clr | rm frost | 681.6/682.0 | -2 | *
2016-08-07 | 220 | 1632 | TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 678.2/678.7 | 0 | *
2016-08-08 | 221 | 1541 | TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 681.5/681.9 | 1 | *
2016-08-09 | 222 | 1832 | TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 680.1/680.4 | 3 | *
2016-08-10 | 223 | 1400 | EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 676.7/677.0 | 5 | *
2016-08-11 | 224 | 1243 | EB | rm frost | rm frost | clr | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | 674.3/674.7 | 6 | * | Reset zeno clock at 1238z.
2016-08-12 | 225 | 1534 | TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 673.7/674.2 | 14 | * | Reset zeno clock at 1546z Power outage 1609-1611z. 1642-1644z
2016-08-13 | 226 | 1755 | EB | rm frost | rm frost | clr | rm frost | rm frost | clr | clr | rm frost | clr | rm frost | 671.6/672.0 | 0 | *
2016-08-14 | 227 | 1759 | EB | clr | clr | clr | clr | clr | clr | clr | rm frost | clr | clr | 672.0/672.6 | 1 | *
2016-08-15 | 228 | 1907 | TG/EB | rm frost | rm frost | clr | rm frost | rm frost | clr | clr | rm frost | clr | clr | 673.2/673.6 | 3 | *
2016-08-16 | 229 | 1810 | TG | clr | clr | clr | clr | rm frost | clr | clr | rm frost | clr | rm frost | 673.3/673.8 | 4 | * | bunting computer was *very* slow. VWS was out of sync. A restart fixed the clock issue and sped up the computer
2016-08-17 | 230 | 1652 | TG | clr | clr | clr | clr | rm frost | clr | clr | clr | clr | clr | 673.8/674.2 | 6 | * | reset zeno clock at 1711z
2016-08-18 | 231 | 1731 | TG | clr | clr | clr | clr | rm frost | lt frost | clr | lt frost | clr | lt frost | 673.5/674.0 | -4 | * | computer is acting very slow again
2016-08-19 | 232 | 1612 | EB/TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 679.5/679.9 | -2 | *
2016-08-20 | 233 | 1222 | TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 681.6/682.0 | -1 | *
2016-08-21 | 234 | 1545 | EB | rm frost | rm frost | clr | rm frost | rm frost | lt frost | clr | lt frost | clr | lt frost | 683.7/684.1 | 1 | *
2016-08-22 | 235 | 1634 | TG | rm frost | clr | clr | clr | rm frost | clr | clr | lt frost | clr | lt frost | 686.3/686.6 | 3 | *
2016-08-23 | 236 | 1255 | EB | rm frost | rm frost | clr | rm frost | rm frost | rm frost | clr | lt frost | clr | lt frost | 687.3/687.5 | 3 | *
2016-08-24 | 237 | 1330 | TG | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | frost | clr | frost | 686.0/686.3 | 7 | * | reset zeno at 1349z
2016-08-25 | 238 | 1223 | EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 682.4/682.8 | -3 | *
2016-08-26 | 239 | 1352 | TG | rm frost | rm frost | clr | clr | rm frost | rm frost | clr | lt frost | clr | lt frost | 682.7/683.1 | -1 | *
2016-08-27 | 240 | 1744 | EB | rm frost | rm frost | clr | rm frost | rm frost | clr | clr | rm frost | clr | rm frost | 683.7/684.3 | 0 | *
2016-08-28 | 241 | 1715 | TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 681.2/681.5 | 2 | *
2016-08-29 | 242 | 1354 | EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 671.8/672.3 | 5 | *
2016-08-30 | 243 | 1300 | TG | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | frost | rm frost | frost | 673.2/673.6 | 8 | * | reset zeno at 1415z
2016-08-31 | 244 | 1306 | EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 670.1/670.4 | -4 | *
2016-09-01 | 245 | 1316 | TG | rm frost | rm frost | clr | rm frost | rm frost | rm frost | clr | lt frost | clr | lt frost | 665.9/666.4 | -3 | *
2016-09-02 | 246 | 1700 | EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 665.2/665.6 | 0 | *
2016-09-03 | 247 | 1323 | TG | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | hv frost | rm frost | hv frost | 666.5/666.9 | 1 | *
2016-09-04 | 248 | 1419 | EB | rm frost | rm frost | clr | rm frost | rm frost | clr | clr | lt frost | lt frost | lt frost | 668.9/669.3 | 3 | *
2016-09-05 | 249 | 1250 | TG | rm frost | rm frost | clr | rm frost | rm frost | frost | clr | frost | clr | frost | 667.4/667.9 | 4 | *
2016-09-06 | 250 | 1342 | EB | clr | clr | clr | clr | clr | clr | clr | lt frost | clr | lt frost | 664.5/664.9 | 6 | * | reset zeno at 1415z
2016-09-07 | 251 | 1243 | TG | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | frost | rm frost | frost | 669.3/669.6 | -3 | *
2016-09-08 | 252 | 1237 | EB | clr | clr | clr | clr | clr | clr | clr | lt frost | clr | lt frost | 667.8/668.3 | -2 | *
2016-09-09 | 253 | 1249 | TG | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | frost | rm frost | frost | 668.9/669.3 | 0 | *
2016-09-10 | 254 | 1233 | EB | clr | clr | clr | clr | clr | clr | clr | lt frost | clr | clr | 669.9/670.3 | 4 | *
2016-09-11 | 255 | 1742 | TG | rm frost | rm frost | clr | clr | clr | rm frost | clr | lt frost | clr | lt frost | 669.3/669.7 | 4 | *
2016-09-12 | 256 | 1237 | EB | clr | clr | clr | clr | clr | clr | clr | lt frost | clr | clr | 671.3/671.7 | 5 | *
2016-09-13 | 257 | 1237 | TG | rm frost | rm frost | clr | rm frost | rm frost | hv frost/partial removal | rm frost | hv frost | rm frost | hv frost | 667.6/668.1 | 7 | * | reset zeno clock at 1255z.
2016-09-14 | 258 | 1246 | EB | clr | clr | clr | clr | rm frost | lt frost | lt frost | lt frost | lt frost | lt frost | 669.7/670.1 | -3 | *
2016-09-15 | 259 | 1220 | TG | rm lt frost | rm lt frost | clr | rm frost | rm frost | hv frost/partial removal | rm frost | frost | rm frost | frost | 672.2/672.6 | -2 | *
2016-09-16 | 260 | 1340 | EB | rm lt frost | rm lt frost | clr | rm frost | rm frost | rm frost | clr | frost | clr | frost | 665.8/666.3 | 0 | *
2016-09-17 | 261 | 1247 | TG | rm lt frost | rm lt frost | clr | rm frost | rm frost | rm frost | clr | hv frost | clr | hv frost | 657.9/658.3 | 1 | *
2016-09-18 | 262 | 1457 | EB | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | frost | rm frost | frost | 660.8/661.2 | 2 | *
2016-09-19 | 263 | 1323 | TG | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | 664.1/664.5 | 4 | * | Reset Zeno clock at 1338z. Computing computer incredibly slow
2016-09-20 | 264 | 1324 | EB | rm frost | rm frost | clr | rm frost | rm frost | rm frost | lt frost | frost | lt frost | frost | 668.5/669.0 | -3 | *
2016-09-21 | 265 | 1226 | TG | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | frost | rm frost | frost | 673.2/673.7 | -1 | * | Andy sent an email that he did some work on intercept
2016-09-22 | 266 | 1318 | EB | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | frost | rm frost | frost | 673.3/673.7 | 0 | *
2016-09-23 | 267 | 1216 | TG | clr | clr | clr | clr | rm frost | clr | clr | frost | clr | frost | 671.1/671.5 | 2 | *
2016-09-24 | 268 | 1409 | EB | clr | clr | clr | clr | clr | clr | clr | lt frost | clr | lt frost | 671.8/672.2 | 4 | *
2016-09-25 | 269 | 1337 | TG | clr | clr | clr | clr | clr | clr | clr | lt frost | clr | lt frost | 676.4/676.9 | 8 | * | Reset Zeno clock at 1349z.
2016-09-26 | 270 | 1257 | EB | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | 676.8/677.2 | -3 | * | Climbed TAWO tower at 1330z.
2016-09-27 | 271 | 1207 | TG | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | frost | rm frost | frost | 672.7/673.1 | -1 | *
2016-09-28 | 272 | 1325 | EB | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | frost | rm frost | frost | 669.6/669.9 | -1 | *
2016-09-29 | 273 | 1247 | TG | clr | clr | clr | rm lt frost | rm lt frost | clr | lt frost | lt frost | lt frost | lt frost | 667.1/667.4 | 1 | *
2016-09-30 | 274 | 1402 | EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 667.1/662.0 | 3 | *
2016-10-01 | 275 | 1323 | TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 658.7/658.5 | 5 | * | Reset Zeno clock at 1332z.
2016-10-02 | 276 | 1403 | EB | clr | clr | clr | clr | clr | clr | clr | lt frost | clr | lt frost | 666.3/666.6 | -4 | *
2016-10-03 | 277 | 1510 | TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 670.6/671.0 | -2 | *
2016-10-04 | 278 | 1219 | EB | clr | clr | clr | clr | clr | clr | clr | lt frost | clr | lt frost | 671.2/671.5 | 0 | *
2016-10-05 | 279 | 1408 | TG | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | 675.1/675.6 | 1 | * | Climbed tower 1400z. TAWO weather station in big house not working. Looked at bunting laptop, saw that it had installed updates and there was a message from Andy that he had tried installing updates overnight but it got too late, and he would fix problems in the morning. Restarted bunting computer to install updates at 1344z. Elissa on TAWO tower cleaning ~1420-1440z
2016-10-06 | 280 | 1314 | EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 672.3/672.8 | 3 | *
2016-10-07 | 281 | 1303 | TG | rm snow | rm snow | clr | clr | clr | clr | clr | clr | clr | clr | 681.7/682.1 | 5 | * | reset Zeno clock at 1324z
2016-10-08 | 282 | 0 | 0 | [13 empty] | High winds/low visibility, didn't make it out to TAWO
2016-10-09 | 283 | 1459 | TG | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | clr | clr | clr | clr | clr | 676.8/677.2 | -1 | *
2016-10-10 | 284 | 1644 | EB | clr | clr | clr | clr | clr | clr | clr | lt frost | clr | lt frost | 673.3/673.6 | 1 | *
2016-10-11 | 285 | 1338 | TG | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | frost | rm frost | frost | 674.4/675.0 | 3 | * | Our GMD web data for Summit wasn't updating so I logged into Bunting and > found the script hung trying to connect to the ftp server. > I killed that dos window and restarted the scheduler and it appears to be > working again. > Andy
2016-10-12 | 286 | 0 | 0
2016-10-13 | 287 | 1303 | TG | clr | clr | clr | clr | clr | frost | clr | clr | clr | clr | 666.6/667.1 | 6 | * | reset Zeno clock at 1317z
2016-10-14 | 288 | 1716 | EB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 669.8/670.2 | -1 | *
2016-10-15 | 289 | 1300 | TG | clr | clr | clr | rm lt frost | clr | frost | clr | clr | clr | clr | 672.2/672.7 | 0 | *
2016-10-16 | 290 | 0 | 0
2016-10-17 | 291 | 1431 | TG | rm frost | rm frost | clr | rm frost | rm frost | frost | frost | frost | frost | frost | 667.1/667.7 | 5 | * | reset Zeno clock at 1641z
2016-10-18 | 292 | 1259 | EB | rm frost | rm frost | clr | rm frost | rm frost | lt frost | lt frost | lt frost | lt frost | lt frost | 663.8/664.2 | -3 | *
2016-10-19 | 293 | 1257 | TG | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | 658.7/659.1 | -1 | * | TG climbed tower at ~1310-1345z to clean instruments. bunting computer *very* slow
2016-10-20 | 294 | 0 | 0 | [13 empty] | Con 1. Didn't go to TAWO.
2016-10-21 | 295 | 1619 | TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 665.1/665.5 | 5 | * | reset Zen clock 1631z
2016-10-22 | 296 | 1213 | EB | clr | clr | clr | rm lt frost | rm lt frost | lt frost | clr | clr | clr | clr | 670.6/671.0 | -2 | *
2016-10-23 | 297 | 1331 | TG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 671.3/671.7 | -1 | *
2016-10-24 | 298 | 1544 | TG/MG/NB | rm frost | rm frost | clr | rm frost | rm frost | frost | lt frost | frost | lt frost | frost | 667.3/667.7 | 1 | * | reset Zeno clock at 1610
2016-10-25 | 299 | 1133 | MG/TG | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | rm frost | rm frost | rm frost | rm frost | rm frost | 660.7/661.2 | -7 | * | NB on tower 1142z *1155z to remove frost
2016-10-26 | 300 | 1213 | TG | rm lt frost | rm lt frost | clr | rm lt frost | rm frost | clr | clr | lt frost | clr | lt frost | 656.6/657.0 | -6 | *
2016-10-27 | 301 | 1128 | MG/NB | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 659.7/660.2 | -4 | *
2016-10-28 | 302 | 1135 | MG/NB | rm frost | rm frost | clr | rm frost | rm frost | lt frost | lt frost | lt frost | lt frost | lt frost | 666.5/666.9 | -3 | *
2016-10-29 | 303 | 1150 | MG/NB | rm lt frost | rm lt frost | clr | clr | clr | lt frost | lt frost | lt frost | lt frost | lt frost | 675.7/676.1 | -1 | *
2016-10-30 | 304 | 1133 | NB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 674.0/674.5 | 1 | *
2016-10-31 | 305 | 1318 | NB | rm frost | rm frost | clr | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | rm frost | 679.0/679.3 | 3 | * | MG climbed tower 1455-1520z and removed frost from instruments.
2016-11-01 | 306 | 1253 | MG | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 675.3/675.8 | 4 | *
2016-11-02 | 307 | 1251 | NB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 674.6/675.2 | 7 | * | Reset Zeno clock at 1300z.
2016-11-03 | 308 | 1245 | MG | rm lt frost | rm lt frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 677.2/677.7 | 0 | *
2016-11-04 | 309 | 1309 | MG | rm lt frost | rm lt frost | clr | clr | clr | clr | clr | clr | clr | clr | 681.2/681.6 | 1 | * | NB on tower 1314z *1327z to remove frost
2016-11-05 | 310 | 1230 | MG | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 674.0/674.6 | -2 | * | 10 m wind bird reading the wrong wind direction. Bird is pointed in the right direction and the encoder is also facing the correct direction.
2016-11-06 | 311 | 1325 | NB | clr | clr | clr | clr | clr | clr | clr | clr | clr | clr | 660.5/661.0 | 0 | *
2016-11-07 | 312 | 1230 | MG | rm lt frost | rm lt frost | clr | clr | clr | clr | clr | clr | clr | clr | 650.0/650.4 | 1 | * | NB on tower 1300z to 1310z and 1329 to 1356z
2016-11-08 | 313 | 1225 | MG | rm lt frost | rm lt frost | clr | lt frost | lt frost | clr | lt frost | lt frost | lt frost | lt frost | 648.9/648.3 | 3 | * | reset Zeno clock at 1317z
2016-11-09 | 314 | 1216 | MG | rm lt frost | rm lt frost | clr | clr | clr | lt frost | clr | lt frost | lt frost | lt frost | 657.2/657.7 | 0 | *
2016-11-10 | 315 | 1242 | MG | rm lt frost | rm lt frost | clr | clr | clr | clr | clr | lt frost | clr | lt frost | 657.7/658.1 | 1 | *
2016-11-11 | 316 | 1249 | MG | [10 empty] | 654.3/654.6 | 3 |  | cannot connect to 50m tower, Did not go to TAWO condition 1
2016-11-12 | 317 | 1240 | MG/NB | rm lt frost | rm lt frost | clr | clr | clr | clr | clr | clr | clr | clr | 650.6/651.0 | -2 | * | reset Zeno clock at 1255 UTC
2016-11-13 | 318 |  |  | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | clr | clr | lt frost | clr | lt frost | 643.8/644.4 | 0 |  | cannot connect to 50m tower
2016-11-14 | 319 | 1900 | NB | rm frost | rm frost | clr | rm frost | rm frost | lt frost | lt frost | lt frost | lt frost | lt frost | 644.0/644.5 | 3 |  | Cannot connect to 50m tower. Wx data for 11/14/16 lost in archive due to file overwrite associated with wrong date entered on 11/15/16. NB
2016-11-15 | 320 | 1400 | MG | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 645.5/645.8 | -2 | * | reset Zeno clock at 1255 UTC
2016-11-16 | 321 | 1302 | MG | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 654.7/655.1 | 0 | * | Reset Zeno clock at 1748z. Corrected date (was set to yesterday). Performed data dump of Zeno data logger and sent data to GMD for processing in order to recover lost data from 11/14/16. During dump Intercept was shut down from 1919z - 2030z. NB
2016-11-17 | 322 | 1218 | MG/NB | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 667.1/667.4 | 0 | *
2016-11-18 | 323 | 1323 | MG | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | frost | lt frost | frost | lt frost | frost | 669.3/669.6 | 1 | * | Closed taskeng.exe at 0955z. NB, tower frosty, too windy to climb
2016-11-19 | 324 | 1235 | NB | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | rm frost | rm frost | rm frost | rm frost | rm frost | 666.0/666.4 | 3 | * | MG climbed tower 1252 - 1336z and removed frost from instruments. 10m windbird is having encoder issues again. Currently roughly 13 degrees out from 16m.
2016-11-20 | 325 | 1312 | NB | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 670.1/670.6 | 5 | * | Reset Zeno clock at 1327.nb
2016-11-21 | 326 | 1305 | NB | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | rm frost | rm frost | rm frost | rm frost | rm frost | 667.9/668.3 | 0 | * | NB on TAWO tower 1406-1421 UTC and removed frost from instruments. 10m wind bird still has encoder issues, roughly 60 degrees off 16m
2016-11-22 | 327 | 1301 | MG/NB | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 661.7/662.0 | 1 | * | 10m windbird still reading wrong direction.
2016-11-23 | 328 | 1230 | NB | clr | clr | clr | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | lt frost | 655.3/655.5 | 3 | * | 10m windbird still reading wrong direction. NB climbed tower 1920-1950z and removed both windbirds. NB climbed tower again 2020-2038z and installed the windbird from 16m on the 10m boom. The former 10m windbird is in TAWO awaiting service. The 16m measurement level currently has no anemometer.
2016-11-24 | 329 | 1250 | NB | rm lt frost | rm lt frost | clr | lt frost | lt frost | lt frost | lt frost | lt frost | N/A | lt frost | 657.8/658.2 | 5 | * | Closed taskeng.exe at 1300z. Reset zeno clock at 1304z. Closed taskeng.exe again at 2304z.
2016-11-25 | 330 | 1249 | MG | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | lt frost | clr | lt frost | N/A | lt frost | 661.7/662.0 | 1 | * | Closed taskeng.exe at 1031z.
2016-11-26 | 331 | 1335 | NB | clr | clr | clr | clr | clr | lt frost | clr | lt frost | N/A | lt frost | 668.9/669.3 | 2 | *
2016-11-27 | 332 | 1420 | MG | clr | clr | clr | clr | clr | clr | clr | clr | N/A | clr | 661.5/662.0 | 4 | * | reset Zeno clock at 1429 UTC
2016-11-28 | 333 | 1235 | NB | clr | clr | clr | clr | clr | clr | clr | clr | N/A | clr | 661.5/661.9 | 0 | * | Closed taskeng.exe at 1010z.
2016-11-29 | 334 | 1245 | NB | lt frost | lt frost | clr | rm lt frost | rm lt frost | lt frost | clr | clr | N/A | lt frost | 663.9/664.2 | 2 | * | Closed taskeng.exe at 1311z.
2016-11-30 | 335 | 1325 | MG | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | lt frost | clr | clr | N/A | lt frost | 668.3/668.8 | 3 | * | Closed taskeng.exe at 1001z.
2016-12-01 | 336 | 1250 | NB | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | rm frost | rm frost | rm frost | N/A | rm frost | 669.2/669.7 | 5 | * | MG climbed tower 1302-1329z and removed frost from instruments. Reset zeno clock at 1306z.
2016-12-02 | 337 | 1321 | MG | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | lt frost | lt frost | lt frost | N/A | lt frost | 665.0/665.4 | 0 | *
2016-12-03 | 338 | 1235 | NB | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | lt frost | lt frost | lt frost | N/A | lt frost | 660.9/661.2 | 1 | * | Closed taskeng.exe at 1244z.
2016-12-04 | 339 | 1348 | NB | clr | clr | clr | rm lt frost | rm lt frost | lt frost | lt frost | lt frost | N/A | lt frost | 666.4/666.9 | 4 | * | Shook tower at 1348z - some frost fell of 10m windbird. Closed taskeng.exe at 1401z. Reset Zeno clock at 1403z.
2016-12-05 | 340 | 1242 | NB | clr | clr | clr | clr | clr | lt frost | lt frost | lt frost | N/A | lt frost | 667.3/667.7 | 0 | * | Closed taskeng.exe at 1012z.
2016-12-06 | 341 | 1244 | MG | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | rm frost | rm frost | rm frost | N/A | rm frost | 664.0/664.3 | 1 | * | NB on tower 1251 UTC - 1320 UTC Removed frost from insturments and tower.
2016-12-07 | 342 | 1240 | MG | clr | clr | clr | clr | clr | clr | clr | clr | N/A | clr | 658.0/658.3 | 3 | *
2016-12-08 | 343 | 1234 | MG/NB | clr | clr | clr | clr | clr | clr | clr | clr | N/A | clr | 661.2/661.5 | -2 | * | Reset Zeno Clock 1311
2016-12-09 | 344 | 1328 | MG | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | clr | clr | clr | N/A | clr | 663.9/664.2 | -1 | * | Shook tower at 1325z - some frost fell of 10m windbird.
2016-12-10 | 345 | 1920 | MG | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | lt frost | lt frost | lt frost | N/A | lt frost | 665.9/666.3 | 2 | * | Closed taskeng.exe at 1008z and 1923z. Shook tower at 1915, some frost fell from wind bird
2016-12-11 | 346 | 1340 | MG | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | lt frost | lt frost | lt frost | N/A | lt frost | 665.0/665.4 | 3 | *
2016-12-12 | 347 | 1230 | NB | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | lt frost | lt frost | lt frost | N/A | lt frost | 657.9/658.3 | 4 | * | Closed taskeng.exe at 1247z. Reset Zeno clock at 1251z.
2016-12-13 | 348 | 1225 | NB | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | lt frost | lt frost | lt frost | N/A | lt frost | 656.4/656.9 | 0 | * | Shook tower at 1220z - frost fell off 10m windbird. Closed taskeng.exe at 1235z.
2016-12-14 | 349 | 1304 | MG | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | lt frost | lt frost | lt frost | N/A | lt frost | 654.5/654.7 | 2 | * | Shook Tower at 1300 UTC
2016-12-15 | 350 | 1415 | MG | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | lt frost | lt frost | lt frost | N/A | lt frost | 652.4/652.8 | 4 | * | Closed taskeng.exe at 1416z. Reset Zeno clock at 1420z
2016-12-16 | 351 | 1308 | MG | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | lt frost | lt frost | lt frost | N/A | lt frost | 651.5/651.7 | 0 | * | Closed taskeng.exe at 1309z
2016-12-17 | 352 | 1252 | MG | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | lt frost | lt frost | lt frost | N/A | lt frost | 655.2/655.5 | 2 | * | Closed taskeng.exe at 1253z
2016-12-18 | 353 | 1328 | NB | clr | clr | clr | clr | clr | lt frost | lt frost | lt frost | N/A | lt frost | 648.1/648.5 | 3 | * | Shook tower 1328z. Closed taskeng.exe at 1343z.
2016-12-19 | 354 | 1220 | NB | clr | clr | clr | clr | clr | lt frost | lt frost | lt frost | N/A | lt frost | 636.5/636.9 | 5 | * | Shook tower at 1220z. Closed taskeng.exe at 1258z. Reset Zeno clock at 1307z.
2016-12-20 | 355 | 1506 | MG | clr | clr | clr | rm lt frost | rm lt frost | lt frost | lt frost | lt frost | N/A | lt frost | 633.7/634.2 | 0 | * | Shook tower 1800 UTC. Closed taskeng at 1816 UTC.
2016-12-21 | 356 | 1325 | MG | rm lt frost | rm lt frost | clr | rm lt frost | rm lt frost | lt frost | lt frost | lt frost | N/A | lt frost | 633.1/633.6 | 3 | * | Shook tower at 1320 UTC. Close Taskeng at 1332 UTC.
2016-12-22 | 357 | 1350 | MG | clr | clr | clr | clr | clr | lt frost | lt frost | lt frost | N/A | lt frost | 635.6/636.1 | 3 | * | Shook tower at 1345 UTC. Close Taskeng at 1358 UTC.
2016-12-23 | 358 | 1244 | NB | clr | clr | clr | clr | clr | lt frost | lt frost | lt frost | N/A | lt frost | 640.8/641.2 | 5 | * | reset zeno clock at 1255z.
2016-12-24 | 359 | 1326 | NB | clr | clr | clr | clr | clr | lt frost | lt frost | lt frost | N/A | lt frost | 649.2/649.6 | 0 | *
2016-12-25 | 360 | 1456 | MG | clr | clr | clr | rm lt frost | rm lt frost | lt frost | lt frost | lt frost | N/A | lt frost | 649.9/650.4 | 1 | * | Shook tower at 1450 UTC. Close Taskeng at 1500 UTC.
2016-12-26 | 361 | 1235 | NB | clr | clr | clr | rm lt frost | rm lt frost | lt frost | lt frost | lt frost | N/A | lt frost | 654.0/654.4 | 4 | * | Closed taskeng.exe at 1305z. Reset Zeno clock at 1307z.
2016-12-27 | 362 | 1034 | MG | clr | clr | clr | clr | clr | clr | clr | clr | N/A | clr | 645.8/646.3 | 0 | *
2016-12-28 | 363 | 1358 | MG | clr | clr | clr | clr | clr | clr | clr | clr | N/A | clr | 633.0/633.4 | 2 | * | No 50 meter tower connection from this date forward
2016-12-29 | 364 | 1308 | MG/NB | clr | clr | clr | clr | clr | clr | clr | clr | N/A | clr | 642.0/642.6 | 2 | *
2016-12-30 | 365 | 1314 | MG/NB | clr | clr | clr | clr | clr | clr | clr | clr | N/A | clr | 656.9/657.3 | 4 | * | reset zeno clock at 1226 UTC
2016-12-31 | 366 | 1332 | NB | clr | clr | clr | clr | clr | clr | clr | clr | N/A | clr | 671.8/672.3 | 7 | * | Reset zeno clock at 1244z.

## Surface TEI 03
Date | Julian Day | Time (UTC) | Observer's Initials | Sample Inlet | Panel Display (30-70 ppb) | Bench Temp (<45 C) | Pressure (>400 mb) | Cell A Flow rate (>0.35 lpm) | Cell B Flow rate (>0.35 lpm) | Intensity Cell A (> 45x10^3) | Intensity Cell B (> 45x10^3) | Comments
2016-01-01 | 1 | 1510 | JG | clr | 40.5 | 36.3 | 484.6 | 0.435 | 0.446 | 77 | 78
2016-01-02 | 2 | 1229 | NB | rm frost | 40.1 | 36.5 | 490.1 | 0.439 | 0.449 | 77 | 78
2016-01-03 | 3 | 1310 | JG | lt frost | 41.7 | 37.6 | 496.2 | 0.443 | 0.454 | 77 | 78
2016-01-04 | 4 | 1228 | NB | rm frost | 41.3 | 36.5 | 499.4 | 0.445 | 0.455 | 77 | 78 | Zero: -0.3, Level 1: 39.5, Level 2: 77.1. JG climbed tower 1355-1417z and removed frost from inlet.
2016-01-05 | 5 | 1230 | JG | clr | 42 | 37.5 | 500.2 | 0.446 | 0.457 | 77 | 78
2016-01-06 | 6 | 1307 | NB | clr | 35.6 | 36.2 | 498.8 | 0.445 | 0.454 | 77 | 78
2016-01-07 | 7 | 1235 | JG | lt frost | 41.2 | 35 | 500.1 | 0.446 | 0.452 | 77 | 78
2016-01-08 | 8 | 1347 | NB | clr | 39.2 | 35.9 | 497.4 | 0.444 | 0.453 | 77 | 78
2016-01-09 | 9 | 1240 | JG | lt frost | 45.2 | 36.2 | 496.4 | 0.444 | 0.453 | 77 | 78
2016-01-10 | 10 | 1504 | NB | rm frost | 41.4 | 36.6 | 495.2 | 0.443 | 0.452 | 77 | 78 | JG climbed tower 1513-1530z and removed frost from inlet.
2016-01-11 | 11 | 1240 | JG | lt frost | 43 | 37 | 498.9 | 0.445 | 0.455 | 77 | 78 | Zero: 0.3, Level 1: 39.0, Level 2: 78.0.
2016-01-12 | 12 | 1232 | NB | rm frost | 41 | 37.2 | 508.5 | 0.451 | 0.462 | 77 | 78 | JG climbed tower 1243-1307z and removed frost from inlet.
2016-01-13 | 13 | 1255 | JG | lt frost | 41.3 | 34.9 | 502.9 | 0.45 | 0.456 | 77 | 78 | Planned power outage for e-gen test 1307-1309 and 1321-1324 UTC.
2016-01-14 | 14 | 1225 | NB | clr | 46.4 | 37.2 | 500.8 | 0.446 | 0.457 | 77 | 78
2016-01-15 | 15 | 1255 | JG | rm frost | 45.1 | 36.9 | 500.8 | 0.447 | 0.457 | 77 | 78 | JG climbed tower 1327-1347 and cleaned frost from inlet.
2016-01-16 | 16 | 1304 | NB | lt frost | 47.2 | 34.8 | 504.3 | 0.448 | 0.456 | 77 | 78
2016-01-17 | 17 | 1725 | JG | lt frost | 48 | 34.4 | 501.5 | 0.447 | 0.453 | 77 | 78
2016-01-18 | 18 | 1229 | NB | lt frost | 48.9 | 35.1 | 500.2 | 0.446 | 0.453 | 77 | 78 | Zero: -0.3, Level 1: 38.8, Level 2: 75.7
2016-01-19 | 19 | 1310 | JG | lt frost | 46.6 | 36.8 | 498.1 | 0.445 | 0.454 | 77 | 78
2016-01-20 | 20 | 1232 | NB | clr | 41.2 | 36.5 | 494.5 | 0.442 | 0.452 | 77 | 78
2016-01-21 | 21 | 1250 | JG | lt frost | 46.6 | 37.3 | 492.6 | 0.441 | 0.452 | 77 | 78
2016-01-22 | 22 | 1234 | NB | lt frost | 46.2 | 37.4 | 488.4 | 0.438 | 0.449 | 77 | 77
2016-01-23 | 23 | 1210 | JG | rm frost | 47.9 | 37.3 | 483.8 | 0.435 | 0.446 | 77 | 77 | NB climbed tower and removed frost from inlets 1217-1234z.
2016-01-24 | 24 | 1258 | NB | clr | 46.2 | 37.5 | 484.3 | 0.435 | 0.447 | 77 | 77
2016-01-25 | 25 | 1230 | JG | lt frost | 44.7 | 36.4 | 478.4 | 0.431 | 0.441 | 77 | 77 | Zero: 0.3, Level 1: 39.1, Level 2: 76.3
2016-01-26 | 26 | 1254 | NB | lt frost | 43.6 | 36.9 | 479.8 | 0.432 | 0.443 | 77 | 77
2016-01-27 | 27 | 1340 | JG | clr | 40 | 36 | 482.1 | 0.433 | 0.444 | 77 | 77
2016-01-28 | 28 | 1234 | NB | lt frost | 40.6 | 36.7 | 483.7 | 0.434 | 0.445 | 77 | 77
2016-01-29 | 29 | 1255 | JG | lt frost | 45.9 | 37.4 | 484.1 | 0.435 | 0.446 | 77 | 77
2016-01-30 | 30 | 1225 | NB | lt frost | 48.3 | 35.5 | 485.6 | 0.436 | 0.444 | 77 | 77
2016-01-31 | 31 | 1455 | JG | lt frost | 37 | 36.4 | 487 | 0.437 | 0.447 | 77 | 77
2016-02-01 | 32 | 1225 | NB | rm frost | 38.6 | 35.7 | 492 | 0.44 | 0.449 | 77 | 77 | Zero: 0.0, Level 1: 39.0, Level 2: 76.5. JG climbed tower 1310-1327z and removed frost from inlet.
2016-02-02 | 33 | 1240 | JG | lt frost | 43.2 | 36.5 | 495.4 | 0.442 | 0.453 | 77 | 77
2016-02-03 | 34 | 1248 | NB | lt frost | 43.6 | 36.3 | 494.8 | 0.442 | 0.452 | 77 | 77
2016-02-04 | 35 | 1220 | JG | rm frost | 44.3 | 37.4 | 495.7 | 0.442 | 0.454 | 77 | 77 | NB climbed tower 1224-1239z and removed frost from inlet.
2016-02-05 | 36 | 1333 | NB | rm frost | 43.6 | 36.6 | 492.2 | 0.44 | 0.45 | 77 | 77 | JG climbed tower 1342-1359z and cleaned frost from inlet.
2016-02-06 | 37 | 1355 | JG | rm frost | 45.9 | 36 | 487.9 | 0.437 | 0.447 | 77 | 77 | NB climbed tower 1407-1421z and cleaned frost from inlet.
2016-02-07 | 38 | 1455 | MB/JJ | lt frost | 41 | 36.1 | 487.2 | 0.437 | 0.446 | 77 | 77
2016-02-08 | 39 | 1240 | MB/JJ/SD | lt frost | 41 | 36.6 | 488.9 | 0.438 | 0.448 | 77 | 77 | Zero: 0.0, Level 1: 39.1, Level 2: 76.2
2016-02-09 | 40 | 1230 | MB/JJ/NB | lt frost | 42 | 36.6 | 488.2 | 0.438 | 0.448 | 77 | 77
2016-02-10 | 41 | 1226 | SWD/JG/NB | lt frost | 41.8 | 36.2 | 486.9 | 0.436 | 0.446 | 77 | 77
2016-02-11 | 42 | 1340 | MB/JJ | lt frost | 42.7 | 36.4 | 487.5 | 0.437 | 0.447 | 77 | 77
2016-02-12 | 43 | 1625 | MB/JJ | lt frost | 47.4 | 36.1 | 489.8 | 0.438 | 0.448 | 77 | 77
2016-02-13 | 44 | 1220 | MB/JJ | lt frost | 46.4 | 36.6 | 493.3 | 0.441 | 0.451 | 77 | 77
2016-02-14 | 45 | 1700 | MB/JJ | lt frost | 44.7 | 36 | 494 | 0.441 | 0.451 | 77 | 77
2016-02-15 | 46 | 1641 | JJ | lt frost | 43.3 | 37.7 | 489.2 | 0.438 | 0.44 | 77 | 77 | Zero: 0.0, Level 1: 40.1, Level 2: 77.0
2016-02-16 | 47 | 1321 | jj/mb | lt frost | 43.1 | 37.6 | 485 | 0.435 | 0.447 | 77 | 77
2016-02-17 | 48 | 1415 | MB/JJ | rm frost | 44 | 37.6 | 489.9 | 0.438 | 0.45 | 77 | 77 | JJ climbed tower 1428-1455z and removed frost from instruments.
2016-02-18 | 49 | 1405 | MB/JJ | clr | 44.1 | 36.5 | 484.4 | 0.435 | 0.444 | 77 | 77
2016-02-19 | 50 | 1349 | JJ/MB | clr | 43.5 | 37.1 | 483.4 | 0.434 | 0.445 | 77 | 77
2016-02-20 | 51 | 1311 | JJ/MB | lt frost | 43.3 | 37.2 | 486.1 | 0.436 | 0.447 | 77 | 77 | JJ climbed tower 1332-1358z and 1457-1526z to troubleshoot 2m RTD cable.
2016-02-21 | 52 | 1510 | JJ/MB | frost | 43.7 | 37.8 | 493.1 | 0.441 | 0.453 | 77 | 77
2016-02-22 | 53 | 1244 | JJ/MB | lt frost | 47.6 | 38.3 | 497.3 | 0.444 | 0.436 | 77 | 77 | Zero: 0.6, Level 1: 40.2, Level 2: 78.5
2016-02-23 | 54 | 1704 | MB | lt frost | 46.7 | 36.3 | 501.6 | 0.447 | 0.456 | 76 | 77
2016-02-24 | 55 | 1314 | JJ/MB | lt frost | 46.8 | 36.1 | 500.7 | 0.446 | 0.455 | 76 | 77
2016-02-25 | 56 | 1336 | jj/mb | clr | 46.1 | 37.3 | 499.5 | 0.446 | 0.456 | 77 | 77
2016-02-26 | 57 | 1348 | JJ/MB | lt frost | 45 | 37.5 | 501.3 | 0.446 | 0.457 | 77 | 77
2016-02-27 | 58 | 1305 | MB | hv frost | 45.5 | 38.2 | 498.8 | 0.445 | 0.457 | 77 | 77 | SWD climbed tower to clean instruments and troubleshoot 2m-RTD cable from 1444-1542z.
2016-02-28 | 59 | 1708 | JJ | lt frost | 42.8 | 36.4 | 494.6 | 0.442 | 0.452 | 76 | 77
2016-02-29 | 60 | 1339 | jj/mb | lt frost | 45.3 | 38.1 | 490.9 | 0.44 | 0.452 | 77 | 77 | Zero: 0.0, Level 1: 40.1, Level 2: 77.7
2016-03-01 | 61 | 1248 | JJ/MB | lt frost | 48.4 | 38.1 | 489.4 | 0.439 | 0.451 | 77 | 77
2016-03-02 | 62 | 1711 | MB/JJ | rm frost | 44.9 | 38 | 491.4 | 0.44 | 0.452 | 77 | 77 | JJ climbed tower to clean instruments from 1737-1833z.
2016-03-03 | 63 | 1338 | MB/JJ | lt frost | 44.1 | 38 | 494.8 | 0.442 | 0.454 | 76 | 77
2016-03-04 | 64 | 1309 | jj/mb | lt frost | 43.8 | 38.1 | 498.1 | 0.444 | 0.456 | 76 | 77
2016-03-05 | 65 | 1235 | jj/mb | lt frost | 50.3 | 38.2 | 494.2 | 0.442 | 0.454 | 76 | 77
2016-03-06 | 66 | 1342 | MB | lt frost | 50.6 | 38.3 | 488.1 | 0.437 | 0.45 | 76 | 77
2016-03-07 | 67 | 1321 | jj/mb | lt frost | 45.9 | 38.8 | 483.1 | 0.434 | 0.448 | 76 | 77 | Zero: 0.3, Level 1: 40.0, Level 2: 78.0
2016-03-08 | 68 | 1331 | jj/mb | lt frost | 41.3 | 36.8 | 487.9 | 0.437 | 0.448 | 76 | 77
2016-03-09 | 69 | 1235 | JJ | lt frost | 43.3 | 38.2 | 492.3 | 0.44 | 0.453 | 76 | 77
2016-03-10 | 70 | 0 | 0 | 0 | [7 empty] | White out conditions (Condition 1) did not travel to TAWO this day.
2016-03-11 | 71 | 1352 | JJ | clr | 42.6 | 38.3 | 474.5 | 0.428 | 0.441 | 76 | 77
2016-03-12 | 72 | 1437 | MB | lt frost | 41.5 | 38.5 | 482.4 | 0.434 | 0.447 | 76 | 77
2016-03-13 | 73 | 1346 | JJ | clr | 43.9 | 38.7 | 486.2 | 0.46 | 0.449 | 76 | 77 | unexpected power outage today from 13:16 to 13:23 UTC
2016-03-14 | 74 | 1353 | jj/mb | lt frost | 42.8 | 38.8 | 479.9 | 0.432 | 0.446 | 76 | 77 | Zero: 0.0, Level 1: 39.5, Level 2: 77.7
2016-03-15 | 75 | 1533 | jj/mb | lt frost | 39.1 | 37.6 | 495.6 | 0.442 | 0.454 | 76 | 77
2016-03-16 | 76 | 1319 | MB | lt frost | 52.6 | 37.5 | 500.1 | 0.445 | 0.456 | 76 | 77
2016-03-17 | 77 | 1338 | jj/mb | clr | 45.5 | 37.9 | 504.7 | 0.449 | 0.46 | 76 | 77
2016-03-18 | 78 | 1305 | MB/JJ | clr | 50.1 | 38 | 503.6 | 0.448 | 0.459 | 76 | 77
2016-03-19 | 79 | 1854 | MB/JJ | lt frost | 43.8 | 37.9 | 498.2 | 0.445 | 0.456 | 76 | 77
2016-03-20 | 80 | 1334 | MB | lt frost | 44.7 | 37.9 | 495.3 | 0.442 | 0.454 | 76 | 77
2016-03-21 | 81 | 1307 | jj/mb | lt frost | 44.8 | 38.2 | 495.3 | 0.442 | 0.454 | 76 | 77 | Zero: 0.0, Level 1: 39.5, Level 2: 78.5
2016-03-22 | 82 | 1320 | MB | lt frost | 45.3 | 37.9 | 496.9 | 0.444 | 0.455 | 76 | 77
2016-03-23 | 83 | 1744 | MB/JJ | lt frost | 45.8 | 38.3 | 492 | 0.44 | 0.453 | 76 | 77
2016-03-24 | 84 | 1329 | jj/mb | lt frost | 47 | 38.2 | 485.4 | 0.435 | 0.448 | 76 | 77
2016-03-25 | 85 | 1303 | JJ | lt frost | 47.5 | 37.8 | 487.8 | 0.437 | 0.449 | 76 | 77
2016-03-26 | 86 | 1303 | MB | lt frost | 54.7 | 37.9 | 491.4 | 0.439 | 0.452 | 76 | 77
2016-03-27 | 87 | 1715 | JJ | clr | 55.1 | 38.1 | 491.1 | 0.44 | 0.452 | 76 | 77
2016-03-28 | 88 | 1327 | MB | clr | [7 empty] | Zero: -0.3, Level 1: 39.8, Level 2: 77.7. Daily diagnostics did not save, values were normal.
2016-03-29 | 89 | 1151 | JJ | lt frost | 56 | 38.2 | 492.3 | 0.44 | 0.452 | 76 | 77
2016-03-30 | 90 | 1134 | MB | lt frost | 53 | 37.7 | 492 | 0.439 | 0.452 | 76 | 77
2016-03-31 | 91 | 1718 | jj/mb | clr | 49.4 | 38.3 | 492.2 | 0.44 | 0.453 | 76 | 77
2016-04-01 | 92 | 1304 | JJ | clr | 50.7 | 38.7 | 493.4 | 0.441 | 0.454 | 76 | 77
2016-04-02 | 93 | 1213 | jj/mb | clr | 47.5 | 38 | 496 | 0.443 | 0.454 | 76 | 77
2016-04-03 | 94 | 1550 | MB | lt frost | 44.2 | 38 | 499.3 | 0.445 | 0.457 | 76 | 77
2016-04-04 | 95 | 1613 | jj/mb | lt frost | 42.1 | 38.2 | 501.4 | 0.446 | 0.458 | 76 | 77 | Zero: 0.7, Level 1: 41.0, Level 2: 79.1
2016-04-05 | 96 | 1251 | MB/JJ | lt frost | 45.7 | 37.6 | 497.2 | 0.443 | 0.454 | 76 | 77
2016-04-06 | 97 | 1332 | JJ | lt frost | 46.5 | 37.8 | 498.3 | 0.444 | 0.456 | 76 | 77 | Planned power outage today
2016-04-07 | 98 | 1224 | MB/JJ | clr | 64.7 | 38.2 | 499.7 | 0.445 | 0.457 | 76 | 77
2016-04-08 | 99 | 1139 | JJ/MB | lt frost | 50.8 | 38.6 | 502.4 | 0.447 | 0.46 | 76 | 77 | SWD climbed the tower from 1715-1800z to measure 10m and 16m windbird alignment.
2016-04-09 | 100 | 1254 | JJ/MB | clr | 60.7 | 38 | 508 | 0.452 | 0.463 | 76 | 77 | Noticed orange filter cap at inlet on tower is missing.
2016-04-10 | 101 | 1707 | JJ | clr | 49.3 | 37.7 | 510.5 | 0.453 | 0.463 | 76 | 76
2016-04-11 | 102 | 1321 | JJ/MB | clr | 55.3 | 38 | 511.4 | 0.453 | 0.464 | 76 | 76 | Zero: 0.0, Level 1: 40.5, Level 2: 78.8
2016-04-12 | 103 | 1554 | JJ/MB | clr | 53.6 | 41.2 | 514.1 | 0.455 | 0.471 | 76 | 77 | Windstorm distrupting ventilation system, room temperature increased. Instrument alarm, elevated bench temp.
2016-04-13 | 104 | 1317 | JJ/MB | clr | 42.5 | 37.6 | 508.3 | 0.451 | 0.461 | 76 | 76
2016-04-14 | 105 | 1329 | JJ/MB | lt frost | 57.3 | 37.7 | 507.5 | 0.451 | 0.461 | 76 | 76 | JJ climbed tower from 1356-1422z and replaced Orange filter cap at the inlet.
2016-04-15 | 106 | 1711 | JJ/MB/SWD | clr | 39.7 | 37.8 | 499.1 | 0.445 | 0.456 | 76 | 76
2016-04-16 | 107 | 1309 | JJ/MB | lt frost | 40.7 | 37.8 | 492.9 | 0.441 | 0.453 | 76 | 77
2016-04-17 | 108 | 1628 | MB | clr | 40.8 | 38.1 | 494.6 | 0.441 | 0.454 | 76 | 77
2016-04-18 | 109 | 1249 | MB | lt frost | 42.5 | 38.5 | 493.5 | 0.441 | 0.454 | 76 | 77 | Zero: 0.0, Level 1: 40.3, Level 2: 78.3
2016-04-19 | 110 | 1152 | JJ/MB | clr | 42.1 | 38.1 | 493.3 | 0.441 | 0.453 | 76 | 77
2016-04-20 | 111 | 1143 | JJ/MB | clr | 42.6 | 38.3 | 496.5 | 0.443 | 0.455 | 76 | 77
2016-04-21 | 112 | 1122 | JJ/MB/SWD | lt frost | 39.2 | 37.9 | 501.6 | 0.446 | 0.458 | 76 | 77 | 1426-1458z and 1513-1527z SWD climbed tower to clean instruments and realign 10m windbird. Also tied sample inlet line of TEI to the tower at the inlet. Filter cup at inlet was removed in between these climbs.
2016-04-22 | 113 | 1838 | JJ/MB | clr | 45.4 | 38 | 504.6 | 0.448 | 0.46 | 76 | 76
2016-04-23 | 114 | 1536 | 0 | 0 | [7 empty] | White out conditions (Condition 1) did not travel to TAWO this day.
2016-04-24 | 115 | 1536 | JJ/MB | clr | 52 | 37.7 | 509.5 | 0.452 | 0.462 | 76 | 76 | Found inlet filter cup on the ground below the tower. It has fallen off due to high winds.
2016-04-25 | 116 | 1207 | JJ/MB | clr | 51.9 | 37.4 | 514.2 | 0.455 | 0.465 | 76 | 76 | Zero: 0.3, Level 1: 40.3, Level 2: 79.4
2016-04-26 | 117 | 1130 | JJ | lt frost | 56.3 | 37.5 | 517 | 0.456 | 0.467 | 76 | 76
2016-04-27 | 118 | 1549 | JJ/MB | lt frost | 50.1 | 39 | 517.4 | 0.457 | 0.469 | 76 | 76
2016-04-28 | 119 | 1315 | JJ/MB | rm frost | 49.8 | 38.9 | 514 | 0.455 | 0.467 | 76 | 76 | 1255-1305z JJ climbed the tower to clean instruments and reattach TEI sample intake bell/filter.
2016-04-29 | 120 | 1300 | MB/JJ | lt frost | 53.7 | 38.4 | 507 | 0.45 | 0.462 | 76 | 76
2016-04-30 | 121 | 1336 | MB/JJ | lt frost | 48.4 | 38.1 | 504.6 | 0.448 | 0.46 | 76 | 76
2016-05-01 | 122 | 1624 | MB/JJ | rm frost | 60.5 | 39.7 | 499.7 | 0.446 | 0.46 | 76 | 77 | 1743-1815z JJ and MB climbed tower to replace filter on CUGC sample inlet.
2016-05-02 | 123 | 1322 | JJ/MB | lt frost | 54.1 | 39 | 500.4 | 0.446 | 0.459 | 76 | 76 | Zero: 0.6, Level 1: 40.8, Level 2: 79.7. Planned power outages starting from 1635 to 1700z
2016-05-03 | 124 | 1230 | MB/JJ | rm frost | 55.9 | 39.1 | 502.4 | 0.447 | 0.46 | 76 | 77 | 1153-1217z JJ climbed tower to clean instruments and attach ties to CUGC intake hose.
2016-05-04 | 125 | 1451 | MB/JJ | clr | 58.5 | 40.1 | 502.7 | 0.448 | 0.463 | 76 | 77
2016-05-05 | 126 | 1528 | MB | clr | 59.1 | 38.7 | 502.8 | 0.447 | 0.46 | 76 | 76
2016-05-06 | 127 | 1845 | JJ/MB | lt frost | 51.5 | 37.6 | 502.7 | 0.447 | 0.458 | 76 | 76
2016-05-07 | 128 | 1315 | JJ/MB | clr | 49.9 | 37.7 | 501.6 | 0.446 | 0.458 | 76 | 76
2016-05-08 | 129 | 1558 | JJ | lt frost | 51 | 37.8 | 504.9 | 0.448 | 0.46 | 76 | 76
2016-05-09 | 130 | 1635 | JJ/MB | clr | 47.2 | 39.8 | 505.3 | 0.449 | 0.463 | 76 | 76 | Zero: 0.3, Level 1: 41.5, Level 2: 79.4
2016-05-10 | 131 | 1648 | JJ/MB | clr | 60.1 | 36.6 | 498.3 | 0.444 | 0.453 | 75 | 76
2016-05-11 | 132 | 1753 | MB/JJ | clr | 52.2 | 37.5 | 508.2 | 0.451 | 0.462 | 76 | 76
2016-05-12 | 133 | 1622 | MB/JJ | lt frost | 42.4 | 36.9 | 509.8 | 0.452 | 0.461 | 76 | 76
2016-05-13 | 134 | 1308 | MB/JJ | clr | 59.9 | 37.1 | 508.3 | 0.451 | 0.461 | 76 | 76
2016-05-14 | 135 | 1149 | JJ/MB | clr | 59.1 | 36.1 | 505.4 | 0.449 | 0.458 | 75 | 76
2016-05-15 | 136 | 1730 | MB | clr | 48.6 | 37.1 | 505.1 | 0.449 | 0.458 | 76 | 76
2016-05-16 | 137 | 1609 | JJ/MB | lt frost | 50.1 | 37.3 | 506 | 0.45 | 0.46 | 76 | 76 | Zero: 0.0, Level 1: 40.3, Level 2: 78.4
2016-05-17 | 138 | 1722 | JJ/MB | lt frost | 59.4 | 36.4 | 500.9 | 0.446 | 0.455 | 75 | 76
2016-05-18 | 139 | 1309 | JJ/MB | clr | 62.5 | 36.9 | 497.4 | 0.443 | 0.454 | 76 | 76
2016-05-19 | 140 | 1322 | JJ | clr | 68.7 | 36.9 | 497.1 | 0.443 | 0.453 | 76 | 76
2016-05-20 | 141 | 1850 | JJ/MB | clr | 59.5 | 37.6 | 498.5 | 0.444 | 0.455 | 76 | 76
2016-05-21 | 142 | 1228 | JJ/MB | lt frost | 57.9 | 37.5 | 498.4 | 0.443 | 0.455 | 76 | 76
2016-05-22 | 143 | 2327 | JJ | lt frost | 55.4 | 37.9 | 500.9 | 0.446 | 0.456 | 76 | 76
2016-05-23 | 144 | 1656 | JJ | clr | 55 | 38 | 502.1 | 0.447 | 0.458 | 76 | 76 | Zero: 0.3, Level 1: 40.1, Level 2: 78.8
2016-05-24 | 145 | 1211 | JJ/MB | lt frost | 56.5 | 37.7 | 502.2 | 0.447 | 0.458 | 76 | 76
2016-05-25 | 146 | 1726 | JJ | clr | 45.8 | 37.5 | 494.5 | 0.441 | 0.452 | 76 | 76
2016-05-26 | 147 | 1232 | JJ/MB | clr | 45 | 36.7 | 492.3 | 0.44 | 0.45 | 75 | 76
2016-05-27 | 148 | 1845 | JJ/mb | clr | 43.2 | 37.8 | 494.5 | 0.441 | 0.453 | 76 | 76
2016-05-28 | 149 | 1221 | MB | clr | 45.6 | 37.3 | 500 | 0.445 | 0.456 | 75 | 76
2016-05-29 | 150 | 1530 | MB | clr | 52.1 | 37.5 | 504.5 | 0.448 | 0.459 | 75 | 76
2016-05-30 | 151 | 1226 | JJ | clr | 48.6 | 38.5 | 505.5 | 0.449 | 0.461 | 75 | 76 | Zero: 0.0, Level 1: 40.2, Level 2: 79.1, Planned power outage today from 2130 to 2230
2016-05-31 | 152 | 1731 | JJ | clr | 40.9 | 37.6 | 506.8 | 0.45 | 0.461 | 75 | 76
2016-06-01 | 153 | 1840 | JJ | clr | 43.1 | 36.6 | 506.7 | 0.45 | 0.459 | 75 | 76
2016-06-02 | 154 | 1831 | MB/JJ | clr | 51.8 | 37.4 | 510.6 | 0.452 | 0.462 | 75 | 76
2016-06-03 | 155 | 1737 | MB/JJ/HJ | clr | 41.8 | 41.2 | 510.8 | 0.453 | 0.469 | 76 | 76 | Instrument in alarm due to high bench/room temps.
2016-06-04 | 156 | 1156 | MB | clr | 53.5 | 36.3 | 508.8 | 0.451 | 0.46 | 75 | 76
2016-06-05 | 157 | 1403 | JJ/TG/EB | clr | 53.8 | 36.9 | 506.2 | 0.45 | 0.459 | 75 | 76
2016-06-06 | 158 | 1613 | MB/JJ/TG/EB | clr | 55.9 | 37 | 511.2 | 0.453 | 0.463 | 75 | 76 | Zero: 0.0, Level 1: 40.6, Level 2: 79.0
2016-06-07 | 159 | 1536 | TG/EB/JJ | clr | 52.5 | 37 | 512.1 | 0.454 | 0.465 | 75 | 76
2016-06-08 | 160 | 1309 | TG/EB/JJ/MB | clr | 41.7 | 36.9 | 512.5 | 0.453 | 0.462 | 75 | 76
2016-06-09 | 161 | 1553 | TG/EB | clr | 48.2 | 37.1 | 511.3 | 0.452 | 0.462 | 75 | 76
2016-06-10 | 162 | 1239 | TG/EB | clr | 46.9 | 38.1 | 505.6 | 0.449 | 0.461 | 76 | 76 | Work on ventilation duct from 1105 to 1330, higher room temps.
2016-06-11 | 163 | 1203 | TG/EB | clr | 46 | 37.2 | 508.1 | 0.451 | 0.46 | 75 | 76
2016-06-12 | 164 | 1509 | TG/EB | clr | 43.1 | 37.2 | 507.7 | 0.45 | 0.459 | 75 | 76
2016-06-13 | 165 | 1637 | TG/EB | clr | 44.5 | 37.3 | 506.8 | 0.45 | 0.459 | 75 | 76 | Calibration performed and normal but values not recorded.
2016-06-14 | 166 | 1153 | TG/EB | clr | 44 | 37.8 | 505.9 | 0.45 | 0.46 | 75 | 76
2016-06-15 | 167 | 1202 | TG/EB | clr | 42.8 | 36.8 | 503.7 | 0.447 | 0.457 | 75 | 76
2016-06-16 | 168 | 1156 | TG | clr | 39.1 | 36.2 | 504.6 | 0.449 | 0.457 | 75 | 76
2016-06-17 | 169 | 1314 | TG | clr | 44.7 | 37.8 | 505.3 | 0.448 | 0.46 | 75 | 76
2016-06-18 | 170 | 1241 | TG/EB | clr | 44 | 37.8 | 507 | 0.45 | 0.46 | 75 | 76
2016-06-19 | 171 | 1808 | EB | clr | 62.8 | 38.1 | 503.1 | 0.447 | 0.458 | 75 | 76
2016-06-20 | 172 | 1812 | TG/EB | clr | 52.9 | 37.7 | 502.4 | 0.446 | 0.457 | 75 | 76 | Calibration performed and normal but values not recorded.
2016-06-21 | 173 | 1204 | TG/EB | clr | 48.5 | 37.8 | 502.9 | 0.447 | 0.458 | 75 | 76
2016-06-22 | 174 | 1332 | TG/EB | clr | 51.2 | 37.8 | 502.4 | 0.447 | 0.458 | 75 | 76
2016-06-23 | 175 | 1313 | EB | clr | 44.7 | 37.9 | 502.9 | 0.447 | 0.458 | 75 | 76
2016-06-24 | 176 | 1329 | TG | clr | 39.8 | 37.6 | 501.8 | 0.446 | 0.457 | 75 | 76
2016-06-25 | 177 | 1717 | EB/TG | clr | 40.1 | 38.6 | 503.1 | 0.447 | 0.459 | 75 | 76
2016-06-26 | 178 | 1828 | TG | clr | 36.4 | 37.6 | 505.3 | 0.449 | 0.458 | 75 | 75
2016-06-27 | 179 | 1603 | EB/TG | clr | 41.5 | 37.7 | 506.1 | 0.449 | 0.459 | 75 | 76 | Zero: 0.0, Level 1: 40.0, Level 2: 78.1
2016-06-28 | 180 | 1452 | EB/TG | clr | 41.3 | 36.9 | 507 | 0.45 | 0.459 | 75 | 76
2016-06-29 | 181 | 1135 | EB/TG | clr | 47.2 | 37.9 | 505.1 | 0.449 | 0.46 | 75 | 75
2016-06-30 | 182 | 1143 | TG/EB | clr | 45.9 | 38.2 | 506 | 0.449 | 0.46 | 75 | 75
2016-07-01 | 183 | 1202 | TG/EB | clr | 41.7 | 37.9 | 507.9 | 0.45 | 0.461 | 75 | 76
2016-07-02 | 184 | 1239 | TG/EB | clr | 40.6 | 37.8 | 507.6 | 0.451 | 0.46 | 75 | 75
2016-07-03 | 185 | 1835 | EB | clr | 41.5 | 37.7 | 508.2 | 0.451 | 0.461 | 75 | 75
2016-07-04 | 186 | 1029 | TG | clr | 39.9 | 38 | 508.4 | 0.451 | 0.461 | 75 | 75 | Zero: 0.4, Level1: 39.3 Level2: 78.3
2016-07-05 | 187 | 1207 | TG/EB | clr | 40.8 | 37.6 | 509.4 | 0.451 | 0.461 | 75 | 75
2016-07-06 | 188 | 0.47153 | TG | clr | 46.2 | 37.6 | 509.8 | 0.452 | 0.461 | 75 | 75
2016-07-07 | 189 | 1214 | TG/EB | clr | 54.4 | 37.9 | 509.4 | 0.452 | 0.462 | 75 | 75
2016-07-08 | 190 | 1604 | TG/EB | clr | - | - | - | - | - | - | - | Instrument has disappeared. We suspect Hannah in Study with poison
2016-07-09 | 191 | 1342 | TG | clr | - | - | - | - | - | - | - | Instrument has disappeared. We suspect K-dub in the greenhouse with a beer bottle
2016-07-10 | 192 | 1319 | TG | clr | 45.1 | 35.9 | 512.9 | 0.454 | 0.461 | 75 | 76
2016-07-11 | 193 | 1323 | EB | clr | 44 | 35.7 | 513.4 | 0.456 | 0.462 | 75 | 76 | Zero: -0.3, Level1: 39.1, Level2: 75.6
2016-07-12 | 194 | 1228 | TG/EB | clr | 45.7 | 36.2 | 512.6 | 0.45 | 0.462 | 75 | 76
2016-07-13 | 195 | 1137 | TG | clr | 45.2 | 36.3 | 510.8 | 0.453 | 0.461 | 75 | 76
2016-07-14 | 196 | 1147 | TG/EB | clr | 45.9 | 36.4 | 509.2 | 0.452 | 0.46 | 76 | 76
2016-07-15 | 197 | 1235 | TG/EB | clr | 47.6 | 35.4 | 507.7 | 0.45 | 0.458 | 75 | 76
2016-07-16 | 198 | 1301 | TG/EB | clr | 46.7 | 35.2 | 507.9 | 0.45 | 0.458 | 75 | 76
2016-07-17 | 199 | 1627 | EB | clr | 43.7 | 35 | 507.9 | 0.45 | 0.457 | 76 | 76
2016-07-18 | 200 | 1207 | TG | clr | 46.4 | 35 | 508.4 | 0.451 | 0.458 | 75 | 75 | Zero: 0.0, Level1: 38.3, Leve2: 75.6 in and out of N winds, but construction has continued on station power outage 2048-54z
2016-07-19 | 201 | 1237 | TG | clr | 44.3 | 34.5 | 509.6 | 0.451 | 0.458 | 75 | 75
2016-07-20 | 202 | 1609 | TG/EB | clr | 34.7 | 34.6 | 510.5 | 0.452 | 0.459 | 75 | 75
2016-07-21 | 203 | 1224 | TG/EB | clr | 41.4 | 34.4 | 512.5 | 0.454 | 0.459 | 75 | 75
2016-07-22 | 204 | 1141 | TG | rm frost | 40.8 | 34.7 | 514.1 | 0.454 | 0.46 | 75 | 75
2016-07-23 | 205 | 1317 | TG | clr | 41.7 | 34.3 | 514.3 | 0.454 | 0.461 | 75 | 75
2016-07-24 | 206 | 1747 | TG | clr | 46 | 34.4 | 515.6 | 0.455 | 0.461 | 75 | 75
2016-07-25 | 207 | 1234 | TG/EB | rm frost | 31.8 | 34.5 | 514.9 | 0.455 | 0.462 | 75 | 75 | Zero: 0.0, Level1: 39.3, Level2: 76.7
2016-07-26 | 208 | 1214 | TG/EB | clr | 43.2 | 34.7 | 515.1 | 0.455 | 0.462 | 75 | 75
2016-07-27 | 209 | 0 | 0 | 0
2016-07-28 | 210 | 1313 | TG/EB | clr | 61.4 | 34.6 | 514.2 | 0.454 | 0.462 | 76 | 76
2016-07-29 | 211 | 1611 | TG/EB | clr | 35.7 | 34.2 | 513 | 0.454 | 0.459 | 75 | 75
2016-07-30 | 212 | 1236 | TG/EB | clr | 26.2 | 34.5 | 513.1 | 0.454 | 0.46 | 76 | 76
2016-07-31 | 213 | 1521 | EB | clr | 35.3 | 34.4 | 511.3 | 0.453 | 0.459 | 75 | 75
2016-08-01 | 214 | 1254 | EB | clr | 38 | 34.6 | 510.2 | 0.452 | 0.458 | 75 | 75 | Zero: -0.3, Level1: 38.8, Level2: 75.6
2016-08-02 | 215 | 1259 | EB | clr | 33.9 | 34.5 | 510.1 | 0.452 | 0.458 | 75 | 75
2016-08-03 | 216 | 1121 | TG | clr | 36.1 | 35 | 510.2 | 0.452 | 0.459 | 75 | 75
2016-08-04 | 217 | 1400 | EB | clr | 39 | 35.1 | 510.6 | 0.452 | 0.459 | 75 | 75
2016-08-05 | 218 | 1353 | EB/TG | rm frost | 39 | 34.7 | 509.3 | 0.452 | 0.458 | 75 | 75
2016-08-06 | 219 | 1717 | EB | rm frost | 38.4 | 34.2 | 509.5 | 0.451 | 0.457 | 75 | 75
2016-08-07 | 220 | 1632 | TG | clr | 34.9 | 35 | 506.5 | 0.45 | 0.457 | 75 | 75
2016-08-08 | 221 | 1541 | TG | clr | 35 | 34.9 | 509.4 | 0.451 | 0.458 | 75 | 75 | Zero: 0.6, Level1: 39.0, Level2: 75.4
2016-08-09 | 222 | 1832 | TG | clr | 34.6 | 35.3 | 508.3 | 0.451 | 0.458 | 75 | 75
2016-08-10 | 223 | 1400 | EB | clr | 49 | 35.4 | 505.8 | 0.449 | 0.456 | 75 | 75
2016-08-11 | 224 | 1243 | EB | rm frost | 44.9 | 34.5 | 503.8 | 0.448 | 0.454 | 75 | 75
2016-08-12 | 225 | 1534 | TG | clr | 47.7 | 34 | 503.6 | 0.448 | 0.454 | 75 | 75 | Power outage 1609-1611z. 1642-1644z
2016-08-13 | 226 | 1755 | EB | rm frost | 35.5 | 34.2 | 502.1 | 0.447 | 0.452 | 75 | 75
2016-08-14 | 227 | 1759 | EB | clr | 49 | 35.6 | 501.9 | 0.446 | 0.455 | 75 | 75
2016-08-15 | 228 | 1907 | TG/EB | rm frost | 38.5 | 36.4 | 502.8 | 0.458 | 0.456 | 75 | 75
2016-08-16 | 229 | 1810 | TG | rm frost | 33.7 | 37 | 503.1 | 0.447 | 0.458 | 75 | 75 | Zero: 0.0, Level1: 40.6, Level2: 77.0
2016-08-17 | 230 | 1652 | TG | rm frost | 31.2 | 34.7 | 503.7 | 0.448 | 0.454 | 75 | 75
2016-08-18 | 231 | 1731 | TG | rm frost | 30.3 | 34.7 | 503 | 0.447 | 0.454 | 75 | 75
2016-08-19 | 232 | 1612 | EB/TG | clr | 37.9 | 33.9 | 507.6 | 0.454 | 0.456 | 75 | 75
2016-08-20 | 233 | 1222 | TG | clr | 33 | 33.9 | 509.2 | 0.451 | 0.457 | 75 | 75
2016-08-21 | 234 | 1545 | EB | rm frost | 32.7 | 34.9 | 511.2 | 0.452 | 0.459 | 75 | 75
2016-08-22 | 235 | 1634 | TG | rm frost | [7 empty] | Zero: 0.3, Level1: 39.1, Level2: 76.3
2016-08-23 | 236 | 1255 | EB | rm frost | 51.1 | 35.2 | 513.4 | 0.455 | 0.461 | 75 | 75
2016-08-24 | 237 | 1330 | TG | rm frost | 49.9 | 35 | 512.9 | 0.454 | 0.46 | 75 | 75
2016-08-25 | 238 | 1223 | EB | clr | 45 | 34.5 | 510.1 | 0.452 | 0.458 | 75 | 75
2016-08-26 | 239 | 1352 | TG | rm frost | 50.4 | 35.1 | 510 | 0.452 | 0.459 | 75 | 75
2016-08-27 | 240 | 1744 | EB | rm frost | 44 | 35.8 | 510.8 | 0.453 | 0.461 | 75 | 75
2016-08-28 | 241 | 1715 | TG | clr | 46.7 | 34.7 | 508.9 | 0.451 | 0.458 | 75 | 75
2016-08-29 | 242 | 1354 | EB | clr | [7 empty] | Zero: -0.3, Level1: 39.0, Level2: 75.4
2016-08-30 | 243 | 1300 | TG | rm frost | 38.2 | 35.4 | 503 | 0.448 | 0.455 | 75 | 75
2016-08-31 | 244 | 1306 | EB | clr | 41.6 | 35.1 | 500.6 | 0.446 | 0.453 | 75 | 75
2016-09-01 | 245 | 1316 | TG | rm frost | 38.9 | 35.8 | 497.5 | 0.444 | 0.451 | 75 | 75
2016-09-02 | 246 | 1700 | EB | clr | 34.2 | 33.7 | 497.1 | 0.443 | 0.449 | 75 | 74
2016-09-03 | 247 | 1323 | TG | rm frost | 35.6 | 35.1 | 497.8 | 0.444 | 0.452 | 75 | 75
2016-09-04 | 248 | 1419 | EB | rm frost | 44.3 | 34.4 | 500.3 | 0.445 | 0.452 | 75 | 74
2016-09-05 | 249 | 1250 | TG | rm frost | 38.5 | 34.8 | 498.9 | 0.444 | 0.451 | 75 | 75 | zero: 0.0, Level1: 38.7, Level2: 74.6
2016-09-06 | 250 | 1342 | EB | clr | 40 | 35.3 | 496.8 | 0.442 | 0.451 | 75 | 75
2016-09-07 | 251 | 1243 | TG | rm frost | 36.9 | 36 | 500.3 | 0.446 | 0.454 | 75 | 75
2016-09-08 | 252 | 1237 | EB | clr | 38.7 | 35.7 | 498.8 | 0.444 | 0.452 | 75 | 75
2016-09-09 | 253 | 1249 | TG | rm frost | 40.1 | 35.1 | 499.4 | 0.445 | 0.453 | 75 | 75
2016-09-10 | 254 | 1233 | EB | clr | 37.1 | 35.9 | 500.8 | 0.445 | 0.453 | 75 | 75
2016-09-11 | 255 | 1742 | TG | clr | 37.2 | 35.9 | 499.8 | 0.444 | 0.453 | 75 | 75
2016-09-12 | 256 | 1237 | EB | clr | 35.4 | 35.5 | 501.5 | 0.446 | 0.454 | 75 | 75 | Zero: -0.3, Level1: 39.5, Level2: 75.6
2016-09-13 | 257 | 1237 | TG | rm frost | 39.2 | 35.8 | 498.5 | 0.443 | 0.453 | 75 | 75
2016-09-14 | 258 | 1246 | EB | rm frost | 35.5 | 34.5 | 500.4 | 0.445 | 0.451 | 75 | 75
2016-09-15 | 259 | 1220 | TG | rm frost | 33.7 | 35.7 | 502.2 | 0.446 | 0.445 | 75 | 74
2016-09-16 | 260 | 1340 | EB | rm frost | 40.3 | 36.3 | 497.2 | 0.443 | 0.452 | 75 | 75
2016-09-17 | 261 | 1247 | TG | rm frost | 43.2 | 35.7 | 491.5 | 0.438 | 0.448 | 75 | 75
2016-09-18 | 262 | 1457 | EB | rm frost | 31.2 | 34.9 | 493.8 | 0.44 | 0.447 | 75 | 75
2016-09-19 | 263 | 1323 | TG | rm frost | 41.7 | 35.7 | 496 | 0.441 | 0.451 | 75 | 75 | Zero: 0.0, Level1: 38.9, Level2: 75.2
2016-09-20 | 264 | 1324 | EB | rm frost | 39.4 | 34.6 | 499.2 | 0.443 | 0.45 | 75 | 75
2016-09-21 | 265 | 1226 | TG | rm frost | 38.5 | 35.8 | 502.8 | 0.445 | 0.454 | 75 | 75
2016-09-22 | 266 | 1318 | EB | rm frost | 38.2 | 35.5 | 502.9 | 0.445 | 0.454 | 75 | 75
2016-09-23 | 267 | 1216 | TG | rm frost | 37.7 | 35.4 | 501.4 | 0.444 | 0.452 | 75 | 75
2016-09-24 | 268 | 1409 | EB | clr | 40 | 35.8 | 502.2 | 0.444 | 0.453 | 75 | 75
2016-09-25 | 269 | 1337 | TG | clr | 35.8 | 35.5 | 505.1 | 0.445 | 0.454 | 75 | 75
2016-09-26 | 270 | 1257 | EB | rm frost | 36.5 | 36.4 | 505.2 | 0.446 | 0.455 | 75 | 75 | Zero: 0.3, Level1: 38.8, Level2: 76.4
2016-09-27 | 271 | 1207 | TG | rm frost | 41.8 | 35.1 | 502.7 | 0.443 | 0.45 | 74 | 74
2016-09-28 | 272 | 1325 | EB | rm frost | 43.9 | 35.9 | 499.9 | 0.441 | 0.451 | 75 | 74
2016-09-29 | 273 | 1247 | TG | rm lt frost | 44.2 | 35.5 | 498.6 | 0.439 | 0.448 | 74 | 74
2016-09-30 | 274 | 1402 | EB | clr | 36.7 | 35.1 | 494 | 0.436 | 0.444 | 75 | 74
2016-10-01 | 275 | 1323 | TG | clr | 33.7 | 35.3 | 492.1 | 0.435 | 0.443 | 75 | 75
2016-10-02 | 276 | 1403 | EB | clr | 33.8 | 34.3 | 498.1 | 0.438 | 0.445 | 75 | 75
2016-10-03 | 277 | 1510 | TG | clr | 35.3 | 35.5 | 500.9 | 0.44 | 0.449 | 75 | 75 | Zero: 0.0, Level1: 38.8, Level2: 75.9
2016-10-04 | 278 | 1219 | EB | clr | 37 | 35.7 | 501.2 | 0.439 | 0.448 | 75 | 75
2016-10-05 | 279 | 1408 | TG | rm frost | 36.3 | 34.7 | 504.1 | 0.442 | 0.448 | 75 | 75
2016-10-06 | 280 | 1314 | EB | clr | 43.2 | 36.2 | 502 | 0.441 | 0.45 | 75 | 75
2016-10-07 | 281 | 1303 | TG | clr | 37.1 | 35.4 | 509.6 | 0.445 | 0.452 | 75 | 74
2016-10-08 | 282 | 0 | 0 | 0 | [7 empty] | High winds/low visibility, didn't make it out to TAWO
2016-10-09 | 283 | 1459 | TG | rm lt frost | 35.2 | 35.9 | 505.5 | 0.437 | 0.446 | 75 | 75
2016-10-10 | 284 | 1644 | EB | clr | 35 | 34.7 | 503.3 | 0.435 | 0.442 | 75 | 74 | Zero: 0.0, Level1: 38.7, Level2: 76.2
2016-10-11 | 285 | 1338 | TG | rm frost | 36.2 | 35.1 | 503.8 | 0.435 | 0.444 | 75 | 74
2016-10-12 | 286 | 0 | 0 | 0
2016-10-13 | 287 | 1303 | TG | clr | 37.7 | 35.8 | 497.1 | 0.43 | 0.44 | 75 | 75
2016-10-14 | 288 | 1716 | EB | clr | 37.1 | 34.5 | 500.6 | 0.432 | 0.439 | 75 | 74
2016-10-15 | 289 | 1300 | TG | clr | 44.5 | 35.4 | 501.9 | 0.433 | 0.442 | 75 | 75
2016-10-16 | 290 | 0 | 0 | 0
2016-10-17 | 291 | 1431 | TG | rm frost | 50.3 | 35.9 | 498.3 | 0.432 | 0.441 | 75 | 74 | Zero: 0.3, Level 1: 38.9, Level2: 75.8
2016-10-18 | 292 | 1259 | EB | rm frost | 48.9 | 35.4 | 495.5 | 0.429 | 0.437 | 75 | 74
2016-10-19 | 293 | 1257 | TG | rm frost | 43.7 | 36.3 | 491.3 | 0.426 | 0.436 | 75 | 74
2016-10-20 | 294 | 0 | 0 | 0 | [7 empty] | Con 1. Didn't go to TAWO.
2016-10-21 | 295 | 1619 | TG | clr | 45.3 | 36 | 496.8 | 0.43 | 0.439 | 74 | 74
2016-10-22 | 296 | 1213 | EB | rm lt frost | 44.6 | 36 | 500.4 | 0.432 | 0.442 | 74 | 74
2016-10-23 | 297 | 1331 | TG | clr | 42.9 | 35.5 | 501.1 | 0.432 | 0.441 | 74 | 74
2016-10-24 | 298 | 1544 | TG/MG/NB | rm frost | 45.7 | 34.9 | 498.1 | 0.429 | 0.437 | 74 | 74
2016-10-25 | 299 | 1133 | MG/TG | rm lt frost | 41.5 | 35.3 | 493.2 | 0.425 | 0.434 | 74 | 74 | Zero: 0.1, Level 1: 38.3, Level2: 75.3
2016-10-26 | 300 | 1213 | TG | rm frost | 39.4 | 35 | 490.7 | 0.419 | 0.427 | 75 | 75
2016-10-27 | 301 | 1128 | MG/NB | rm lt frost | 36.8 | 34.6 | 493.7 | 0.302 | 0.31 | 74 | 74
2016-10-28 | 302 | 1135 | MG/NB | rm frost | 38.9 | 34.7 | 499.4 | 0.216 | 0.224 | 74 | 74 | Low flow alarm, climbed tower, cleaned inlet, breifly disconnected sample line, did not resolve issue
2016-10-29 | 303 | 1150 | MG/NB | clr | 35.6 | 35 | 505.6 | 0.206 | 0.215 | 73 | 73 | low flow alarm, waiting on response from PIs.
2016-10-30 | 304 | 1133 | NB | clr | 38.7 | 35.1 | 504.4 | 0.196 | 0.207 | 74 | 73 | Low flow alarm. Waiting on response from PIs.
2016-10-31 | 305 | 1318 | NB | rm frost | 39.4 | 34.8 | 508.3 | 0.186 | 0.195 | 74 | 73 | Low flow alarm. Waiting on response from PIs. Did not do weekly calibration due to low flow alarm status.
2016-11-01 | 306 | 1253 | MG | rm lt frost | 47 | 35.9 | 505.4 | 0.182 | 0.193 | 74 | 73 | Low flow alarm. Waiting on response from PIs.
2016-11-02 | 307 | 1251 | NB | clr | 35.2 | 33.3 | 505.4 | 0.172 | 0.182 | 73 | 73 | Low flow alarm.
2016-11-03 | 308 | 1245 | MG | lt frost | 35.5 | 34.7 | 507.3 | 0.167 | 0.18 | 73 | 73 | Low Flow alarm
2016-11-04 | 309 | 1309 | MG | clr | 40 | 35.6 | 510.3 | 0.168 | 0.174 | 73 | 73 | Low Flow alarm
2016-11-05 | 310 | 1230 | MG | clr | 32.9 | 34.2 | 505.2 | 0.444 | 0.45 | 73 | 73 | Replaced pump yesterday evening
2016-11-06 | 311 | 1325 | NB | clr | [7 empty] | Did daily checks, info some how lost
2016-11-07 | 312 | 1230 | MG | clr | 37.67 | 35.1 | 493.3 | 0.437 | 0.445 | 73 | 73 | Zero: -0.3, Level 1: 36.6, Level2: 71.1
2016-11-08 | 313 | 1225 | MG | lt frost | 34.8 | 34.2 | 484.3 | 0.431 | 0.437 | 74 | 43
2016-11-09 | 314 | 1216 | MG | clr | 36.6 | 33.6 | 490.6 | 0.435 | 0.44 | 74 | 73
2016-11-10 | 315 | 1242 | MG | clr | 37.86 | 33.6 | 490.9 | 0.435 | 0.441 | 74 | 73
2016-11-11 | 316 | [10 empty] | Did not go to TAWO condition 1
2016-11-12 | 317 | 1240 | MG/NB | clr | 37.9 | 25.5 | 485.9 | 0.43 | 0.423 | 73 | 72
2016-11-13 | 318 | 1250 | MG | clr | 37.2 | 36.5 | 480.6 | 0.428 | 0.427 | 75 | 74
2016-11-14 | 319 | 1900 | NB | rm frost | 38.7 | 35.1 | 480.9 | 0.428 | 0.435 | 75 | 74
2016-11-15 | 320 | 1400 | MG | rm lt frost | 37.9 | 35.4 | 482 | 0.428 | 0.437 | 75 | 74 | Zero: 0.0, Level 1: 37.8, Level2: 74.0
2016-11-16 | 321 | 1302 | MG | rm lt frost | 37.8 | 35 | 489.3 | 0.434 | 0.441 | 75 | 74
2016-11-17 | 322 | 1218 | MG/NB | rm lt frost | 37.9 | 32.1 | 498.5 | 0.44 | 0.443 | 74 | 73
2016-11-18 | 323 | 1323 | MG | rm lt frost | 36.1 | 34.9 | 500.1 | 0.441 | 0.449 | 74 | 74
2016-11-19 | 324 | 1235 | NB | rm lt frost | 35.2 | 35.4 | 497.2 | 0.44 | 0.448 | 74 | 74
2016-11-20 | 325 | 1312 | NB | rm lt frost | 34.4 | 35 | 500.9 | 0.441 | 0.449 | 74 | 74
2016-11-21 | 326 | 1305 | NB | rm lt frost | 31.96 | 35.8 | 498.7 | 0.441 | 0.449 | 74 | 74 | Zero: -0.0, Level 1: 38.4, Level2: 75.2
2016-11-22 | 327 | 1301 | MG/NB | rm lt frost | 33.57 | 34.8 | 494.6 | 0.438 | 0.444 | 74 | 73
2016-11-23 | 328 | 1230 | NB | lt frost | 37.2 | 33.2 | 489.7 | 0.434 | 0.438 | 74 | 73
2016-11-24 | 329 | 1250 | NB | lt frost | 36.5 | 27.8 | 491.8 | 0.435 | 0.431 | 74 | 73
2016-11-25 | 330 | 1249 | MG | rm lt frost | 33.73 | 34.5 | 494.6 | 0.437 | 0.444 | 74 | 73
2016-11-26 | 331 | 1335 | NB | clr | 40.4 | 35.1 | 499.7 | 0.44 | 0.447 | 74 | 74
2016-11-27 | 332 | 1420 | MG | clr | 38.2 | 34.6 | 493.8 | 0.437 | 0.444 | 74 | 74
2016-11-28 | 333 | 1235 | NB | clr | 37.7 | 32.6 | 494.3 | 0.436 | 0.441 | 74 | 73 | Zero: -0.3, Level 1: 36.9, Level 2: 72.1
2016-11-29 | 334 | 1245 | NB | rm lt frost | 40.8 | 31.1 | 496.7 | 0.439 | 0.44 | 74 | 73
2016-11-30 | 335 | 1325 | MG | rm lt frost | 37.1 | 34.1 | 499.2 | 0.44 | 0.446 | 74 | 73
2016-12-01 | 336 | 1250 | NB | rm lt frost | 35.39 | 34 | 499.8 | 0.441 | 0.447 | 74 | 73
2016-12-02 | 337 | 1321 | MG | rm lt frost | 41 | 34.6 | 497 | 0.438 | 0.445 | 74 | 74
2016-12-03 | 338 | 1235 | NB | rm lt frost | 39.6 | 34.9 | 493.6 | 0.436 | 0.444 | 74 | 73
2016-12-04 | 339 | 1348 | NB | rm lt frost | 39 | 34.4 | 498.1 | 0.439 | 0.445 | 74 | 73
2016-12-05 | 340 | 1242 | NB | clr | 42.8 | 32.9 | 498.2 | 0.44 | 0.444 | 74 | 73 | Zero: -0.3, Level 1: 37.3, Level 2: 73.3
2016-12-06 | 341 | 1244 | MG | rm lt frost | 41.1 | 34 | 495.9 | 0.438 | 0.444 | 74 | 73
2016-12-07 | 342 | 1240 | MG | clr | 44.5 | 33 | 491.3 | 0.434 | 0.439 | 74 | 73
2016-12-08 | 343 | 1234 | MG/NB | clr | 37.6 | 32.8 | 493.8 | 0.436 | 0.44 | 74 | 73
2016-12-09 | 344 | 1328 | MG | rm lt frost | 36.3 | 35.4 | 496 | 0.438 | 0.473 | 74 | 73
2016-12-10 | 345 | 1920 | MG | rm lt frost | 38.9 | 35.7 | 497.4 | 0.439 | 0.447 | 75 | 74
2016-12-11 | 346 | 1340 | MG | rm lt frost | 38.7 | 36.1 | 496.4 | 0.438 | 0.448 | 74 | 73
2016-12-12 | 347 | 1230 | NB | rm lt frost | 38.3 | 35.7 | 491.7 | 0.434 | 0.443 | 74 | 73 | Zero: -0.7, Level 1: 37.8, Level 2: 75.4
2016-12-13 | 348 | 1225 | NB | rm lt frost | 36.4 | 35.7 | 490.2 | 0.434 | 0.443 | 74 | 73
2016-12-14 | 349 | 1304 | MG | rm lt frost | 39.8 | 34.9 | 489.1 | 0.433 | 0.44 | 74 | 73
2016-12-15 | 350 | 1415 | MG | rm lt frost | 39.7 | 36 | 487.2 | 0.432 | 0.44 | 74 | 73
2016-12-16 | 351 | 1308 | MG | rm lt frost | 38.7 | 35.1 | 486.8 | 0.431 | 0.439 | 74 | 73
2016-12-17 | 352 | 1252 | MG | rm lt frost | 38.2 | 35.4 | 489.7 | 0.433 | 0.442 | 74 | 73
2016-12-18 | 353 | 1328 | NB | clr | 41.7 | 35 | 484.2 | 0.429 | 0.437 | 74 | 73
2016-12-19 | 354 | 1220 | NB | clr | 43.8 | 31.5 | 476 | 0.424 | 0.425 | 74 | 73 | Zero: 0.3, Level 1: 36.0, Level 2: 69.9
2016-12-20 | 355 | 1506 | MG | rm lt frost | 40 | 35.1 | 473.7 | 0.422 | 0.429 | 74 | 74
2016-12-21 | 356 | 1325 | MG | rm lt frost | 40.1 | 35.3 | 473.3 | 0.422 | 0.43 | 74 | 73
2016-12-22 | 357 | 1350 | MG | clr | 39 | 35.2 | 474.7 | 0.423 | 0.43 | 74 | 73
2016-12-23 | 358 | 1244 | NB | clr | 38.6 | 34.5 | 478.9 | 0.425 | 0.432 | 74 | 73 | Power outages for e-gen swap: 1330-1333z and 1645-1648z.
2016-12-24 | 359 | 1326 | NB | clr | 38.4 | 34.2 | 484.9 | 0.43 | 0.437 | 74 | 73
2016-12-25 | 360 | 1456 | MG | rm lt frost | 39.2 | 34.4 | 485.3 | 0.43 | 0.437 | 74 | 73
2016-12-26 | 361 | 1235 | NB | rm lt frost | 40 | 34.7 | 488.8 | 0.433 | 0.439 | 74 | 73 | Zero: -0.4, Level 1: 37.7, Level 2: 74.1
2016-12-27 | 362 | 1034 | MG | clr | 39.6 | 26.2 | 482.4 | 0.428 | 0.421 | 73 | 72
2016-12-28 | 363 | 1358 | MG | clr | 35.6 | 34.8 | 472.6 | 0.422 | 0.429 | 75 | 74
2016-12-29 | 364 | 1308 | MG/NB | clr | 34.7 | 34.9 | 479.6 | 0.427 | 0.433 | 75 | 74
2016-12-30 | 365 | 1314 | MG/NB | clr | 38 | 25.4 | 490.6 | 0.434 | 0.441 | 75 | 74
2016-12-31 | 366 | 1332 | NB | clr | 37.9 | 32.1 | 502.1 | 0.443 | 0.445 | 75 | 73

## DRUM Sampler
Date | Julian Day | Time (UTC) | Observer's Initials | Inlet Clear? | Pump running/temp OK? | (Weekly) Vacuum (in HG) | (Weekly) "Flow" (WC) | Comments
2016-01-01 | 1 | 1510 | JG | rm frost | *
2016-01-02 | 2 | 1229 | NB | clr | *
2016-01-03 | 3 | 1310 | JG | clr | *
2016-01-04 | 4 | 1228 | NB | rm frost | * | 11 | 5.5
2016-01-05 | 5 | 1230 | JG | clr | *
2016-01-06 | 6 | 1307 | NB | rm frost | *
2016-01-07 | 7 | 1235 | JG | rm frost | *
2016-01-08 | 8 | 1347 | NB | clr | *
2016-01-09 | 9 | 1240 | JG | clr | *
2016-01-10 | 10 | 1504 | NB | clr | *
2016-01-11 | 11 | 1240 | JG | rm frost | * | 11 | 5.5
2016-01-12 | 12 | 1232 | NB | clr | *
2016-01-13 | 13 | 1255 | JG | rm frost | *
2016-01-14 | 14 | 1225 | NB | clr | *
2016-01-15 | 15 | 1255 | JG | rm frost | *
2016-01-16 | 16 | 1304 | NB | rm frost | *
2016-01-17 | 17 | 1725 | JG | rm frost | *
2016-01-18 | 18 | 1229 | NB | clr | * | 10 | 5.5
2016-01-19 | 19 | 1310 | JG | clr | *
2016-01-20 | 20 | 1232 | NB | clr | *
2016-01-21 | 21 | 1250 | JG | rm frost | *
2016-01-22 | 22 | 1234 | NB | clr | *
2016-01-23 | 23 | 1210 | JG | rm frost | *
2016-01-24 | 24 | 1258 | NB | clr | *
2016-01-25 | 25 | 1230 | JG | clr | * | 10 | 5.5
2016-01-26 | 26 | 1254 | NB | rm frost | *
2016-01-27 | 27 | 1340 | JG | rm frost | *
2016-01-28 | 28 | 1234 | NB | rm frost | *
2016-01-29 | 29 | 1255 | JG | clr | *
2016-01-30 | 30 | 1225 | NB | rm frost | *
2016-01-31 | 31 | 1455 | JG | rm frost | *
2016-02-01 | 32 | 1225 | NB | clr | * | 10 | 5.5
2016-02-02 | 33 | 1240 | JG | clr | *
2016-02-03 | 34 | 1248 | NB | clr | *
2016-02-04 | 35 | 1220 | JG | rm frost | *
2016-02-05 | 36 | 1333 | NB | rm frost | *
2016-02-06 | 37 | 1355 | JG | rm frost | *
2016-02-07 | 38 | 1455 | MB/JJ | rm frost | *
2016-02-08 | 39 | 1240 | MB/JJ/SD | rm frost | * | 9.8 | 5.5
2016-02-09 | 40 | 1230 | MB/JJ/NB | rm frost | *
2016-02-10 | 41 | 1226 | SWD/JG/NB | clr | *
2016-02-11 | 42 | 1340 | MB/JJ | rm frost | *
2016-02-12 | 43 | 1625 | MB/JJ | rm frost | *
2016-02-13 | 44 | 1220 | MB/JJ | rm frost | *
2016-02-14 | 45 | 1700 | MB/JJ | rm frost | *
2016-02-15 | 46 | 1641 | JJ | rm frost | * | 10 | 5.5
2016-02-16 | 47 | 1321 | jj/mb | rm frost | *
2016-02-17 | 48 | 1415 | MB/JJ | rm frost | *
2016-02-18 | 49 | 1405 | MB/JJ | clr | *
2016-02-19 | 50 | 1349 | JJ/MB | rm frost | *
2016-02-20 | 51 | 1311 | JJ/MB | rm frost | *
2016-02-21 | 52 | 1510 | JJ/MB | rm frost | *
2016-02-22 | 53 | 1244 | JJ/MB | rm frost | * | 10 | 5.5
2016-02-23 | 54 | 1704 | MB | rm frost | *
2016-02-24 | 55 | 1314 | JJ/MB | clr | *
2016-02-25 | 56 | 1336 | jj/mb | clr | *
2016-02-26 | 57 | 1348 | JJ/MB | rm frost | *
2016-02-27 | 58 | 1305 | MB | rm frost | *
2016-02-28 | 59 | 1708 | JJ | rm frost | *
2016-02-29 | 60 | 1339 | jj/mb | rm frost | * | 9.9 | 5.5
2016-03-01 | 61 | 1248 | JJ/MB | clr | *
2016-03-02 | 62 | 1711 | MB/JJ | rm frost | *
2016-03-03 | 63 | 1338 | MB/JJ | rm frost | *
2016-03-04 | 64 | 1309 | jj/mb | rm frost | *
2016-03-05 | 65 | 1235 | jj/mb | rm frost | *
2016-03-06 | 66 | 1342 | MB | rm frost | *
2016-03-07 | 67 | 1321 | jj/mb | clr | * | 9.5 | 5.5
2016-03-08 | 68 | 1331 | jj/mb | clr | *
2016-03-09 | 69 | 1235 | JJ | rm frost | *
2016-03-10 | 70 | 0 | 0 |  |  |  |  | White out conditions (Condition 1) did not travel to TAWO this day.
2016-03-11 | 71 | 1352 | JJ | clr | *
2016-03-12 | 72 | 1437 | MB | rm frost | *
2016-03-13 | 73 | 1346 | JJ | rm frost | * |  |  | unexpected power outage today from 13:16 to 13:23 UTC
2016-03-14 | 74 | 1353 | jj/mb | clr | * | 9.6 | 5.5
2016-03-15 | 75 | 1533 | jj/mb | rm frost | *
2016-03-16 | 76 | 1319 | MB | rm frost | *
2016-03-17 | 77 | 1338 | jj/mb | rm frost | *
2016-03-18 | 78 | 1305 | MB/JJ | rm frost | *
2016-03-19 | 79 | 1854 | MB/JJ | rm frost | *
2016-03-20 | 80 | 1334 | MB | rm frost | *
2016-03-21 | 81 | 1307 | jj/mb | rm frost | * | 9.8 | 5
2016-03-22 | 82 | 1320 | MB | rm frost | *
2016-03-23 | 83 | 1744 | MB/JJ | clr | *
2016-03-24 | 84 | 1329 | jj/mb | rm frost | *
2016-03-25 | 85 | 1303 | JJ | rm frost | *
2016-03-26 | 86 | 1303 | MB | rm frost | *
2016-03-27 | 87 | 1715 | JJ | clr | *
2016-03-28 | 88 | 1327 | MB | clr | * | 9.5 | 5.5
2016-03-29 | 89 | 1151 | JJ | rm frost | *
2016-03-30 | 90 | 1134 | MB | clr | *
2016-03-31 | 91 | 1718 | jj/mb | clr | *
2016-04-01 | 92 | 1304 | JJ | clr | *
2016-04-02 | 93 | 1213 | jj/mb | clr | *
2016-04-03 | 94 | 1550 | MB | rm frost | *
2016-04-04 | 95 | 1613 | jj/mb | rm frost | * | 10 | 5.5
2016-04-05 | 96 | 1251 | MB/JJ | rm frost | *
2016-04-06 | 97 | 1332 | JJ | clr | * |  |  | Planned power outage today
2016-04-07 | 98 | 1224 | MB/JJ | clr | *
2016-04-08 | 99 | 1139 | JJ/MB | clr | *
2016-04-09 | 100 | 1254 | JJ/MB | clr | *
2016-04-10 | 101 | 1707 | JJ | clr | *
2016-04-11 | 102 | 1321 | JJ/MB | clr | * | 9.8 | 5.5
2016-04-12 | 103 | 1554 | JJ/MB | clr | *
2016-04-13 | 104 | 1317 | JJ/MB | clr | *
2016-04-14 | 105 | 1329 | JJ/MB | rm frost | *
2016-04-15 | 106 | 1711 | JJ/MB/SWD | clr | *
2016-04-16 | 107 | 1309 | JJ/MB | rm frost | *
2016-04-17 | 108 | 1628 | MB | clr | *
2016-04-18 | 109 | 1249 | MB | clr | * | 9.1 | 5.5
2016-04-19 | 110 | 1152 | JJ/MB | rm snow | *
2016-04-20 | 111 | 1143 | JJ/MB | rm snow | *
2016-04-21 | 112 | 1122 | JJ/MB/SWD | rm frost | *
2016-04-22 | 113 | 1838 | JJ/MB | clr | *
2016-04-23 | 114 | 1536 | 0 |  |  |  |  | White out conditions (Condition 1) did not travel to TAWO this day.
2016-04-24 | 115 | 1536 | JJ/MB | clr | *
2016-04-25 | 116 | 1207 | JJ/MB | clr | * | 9.3 | 5.5
2016-04-26 | 117 | 1130 | JJ | rm frost | *
2016-04-27 | 118 | 1549 | JJ/MB | rm frost | *
2016-04-28 | 119 | 1315 | JJ/MB | rm frost | *
2016-04-29 | 120 | 1300 | MB/JJ | clr | *
2016-04-30 | 121 | 1336 | MB/JJ | rm frost | *
2016-05-01 | 122 | 1624 | MB/JJ | clr | *
2016-05-02 | 123 | 1322 | JJ/MB | rm snow | * | 10.1 | 5.5 | Planned power outages starting from 1635 to 1700z
2016-05-03 | 124 | 1230 | MB/JJ | rm frost | *
2016-05-04 | 125 | 1451 | MB/JJ | clr | *
2016-05-05 | 126 | 1528 | MB | clr | *
2016-05-06 | 127 | 1845 | JJ/MB | clr | *
2016-05-07 | 128 | 1315 | JJ/MB | clr | *
2016-05-08 | 129 | 1558 | JJ | clr | *
2016-05-09 | 130 | 1635 | JJ/MB | clr | * | 9.8 | 5.5
2016-05-10 | 131 | 1648 | JJ/MB | rm snow | *
2016-05-11 | 132 | 1753 | MB/JJ | clr | * |  |  | O-ring broken and one spacer needs screw in sample inlet cap.
2016-05-12 | 133 | 1622 | MB/JJ | clr | * |  |  | Replaced spacer in sample inlet cap.
2016-05-13 | 134 | 1308 | MB/JJ | rm snow | *
2016-05-14 | 135 | 1149 | JJ/MB | clr | *
2016-05-15 | 136 | 1730 | MB | clr | *
2016-05-16 | 137 | 1609 | JJ/MB | clr | * | 10 | 5.5 | Tested spare O ring on inlet cap but it was too big. Inlet reinsatlled without complete O ring
2016-05-17 | 138 | 1722 | JJ/MB | clr | *
2016-05-18 | 139 | 1309 | JJ/MB | clr | *
2016-05-19 | 140 | 1322 | JJ | clr | *
2016-05-20 | 141 | 1850 | JJ/MB | clr | *
2016-05-21 | 142 | 1228 | JJ/MB | rm snow | *
2016-05-22 | 143 | 2327 | JJ | clr | *
2016-05-23 | 144 | 1656 | JJ | rm snow | * | 10.4 | 5.5
2016-05-24 | 145 | 1211 | JJ/MB | clr | *
2016-05-25 | 146 | 1726 | JJ | rm snow | *
2016-05-26 | 147 | 1232 | JJ/MB | rm snow | *
2016-05-27 | 148 | 1845 | JJ/mb | rm snow | *
2016-05-28 | 149 | 1221 | MB | clr | *
2016-05-29 | 150 | 1530 | MB | clr | *
2016-05-30 | 151 | 1226 | JJ | rm snow | * | 10 | 5.5 | Planned power outage today from 2130 to 2230 DRUM sampler remained running during outage.
2016-05-31 | 152 | 1731 | JJ | clr | *
2016-06-01 | 153 | 1840 | JJ | rm snow | *
2016-06-02 | 154 | 1831 | MB/JJ | clr | *
2016-06-03 | 155 | 1737 | MB/JJ/HJ | clr | *
2016-06-04 | 156 | 1156 | MB | clr | *
2016-06-05 | 157 | 1403 | JJ/TG/EB | clr | * |  |  | Replaced broken O-ring in inlet
2016-06-06 | 158 | 1613 | MB/JJ/TG/EB | clr | * | 10.5 | 6
2016-06-07 | 159 | 1536 | TG/EB/JJ | rm snow | * |  |  | Pump off for replacement of drums at 1839-1848z
2016-06-08 | 160 | 1309 | TG/EB/JJ/MB | clr | *
2016-06-09 | 161 | 1553 | TG/EB | clr | *
2016-06-10 | 162 | 1239 | TG/EB | clr | * |  |  | Work on ventilation duct from 1105 to 1330, higher room temps.
2016-06-11 | 163 | 1203 | TG/EB | clr | *
2016-06-12 | 164 | 1509 | TG/EB | clr | *
2016-06-13 | 165 | 1637 | TG/EB | clr | * | 10.5 | 5.5
2016-06-14 | 166 | 1153 | TG/EB | clr | *
2016-06-15 | 167 | 1202 | TG/EB | clr | *
2016-06-16 | 168 | 1156 | TG | clr | *
2016-06-17 | 169 | 1314 | TG | clr | *
2016-06-18 | 170 | 1241 | TG/EB | clr | *
2016-06-19 | 171 | 1808 | EB | clr | *
2016-06-20 | 172 | 1812 | TG/EB | clr | * | 10.6 | 5.5
2016-06-21 | 173 | 1204 | TG/EB | clr | *
2016-06-22 | 174 | 1332 | TG/EB | clr | *
2016-06-23 | 175 | 1313 | EB | clr | *
2016-06-24 | 176 | 1329 | TG | clr | *
2016-06-25 | 177 | 1717 | EB/TG | clr | *
2016-06-26 | 178 | 1828 | TG | clr | *
2016-06-27 | 179 | 1603 | EB/TG | clr | * | 11 | 5.5
2016-06-28 | 180 | 1452 | EB/TG | clr | *
2016-06-29 | 181 | 1135 | EB/TG | clr | *
2016-06-30 | 182 | 1143 | TG/EB | clr | *
2016-07-01 | 183 | 1202 | TG/EB | clr | * |  |  | 1227z pump and drum sampler turned off in preparation for carp work at TAWO
2016-07-02 | 184 | 1239 | TG/EB | clr | - |  |  | pump and sampler remained off for TAWO construction 22:02 turned pump and drum sampler back on. Vacuum 10.5 inHG and Flow 5.5
2016-07-03 | 185 | 1835 | EB
2016-07-04 | 186 | 1029 | TG | rm frost | * |  |  | turned pump and drum sampler off at 1039z for construction on TAWO tower instruments
2016-07-05 | 187 | 1207 | TG/EB | clr | * | 10.5 | 5.3 | turned pump and drum sampler on at 1203z
2016-07-06 | 188 | 0.47153 | TG | clr | *
2016-07-07 | 189 | 1214 | TG/EB | clr | *
2016-07-08 | 190 | 1604 | TG/EB | clr | *
2016-07-09 | 191 | 1342 | TG | clr | *
2016-07-10 | 192 | 1319 | TG | clr | *
2016-07-11 | 193 | 1323 | EB | clr | * | 11 | 5.4
2016-07-12 | 194 | 1228 | TG/EB | clr | *
2016-07-13 | 195 | 1137 | TG | clr | *
2016-07-14 | 196 | 1147 | TG/EB | clr | *
2016-07-15 | 197 | 1235 | TG/EB | clr | *
2016-07-16 | 198 | 1301 | TG/EB | clr | *
2016-07-17 | 199 | 1627 | EB | clr | *
2016-07-18 | 200 | 1207 | TG | clr | * | 11 | 5.5 | in and out of N winds, but construction has continued on station power outage 2048-54z
2016-07-19 | 201 | 1237 | TG | clr | *
2016-07-20 | 202 | 1609 | TG/EB | clr | *
2016-07-21 | 203 | 1224 | TG/EB | clr | *
2016-07-22 | 204 | 1141 | TG | rm frost | *
2016-07-23 | 205 | 1317 | TG | clr | *
2016-07-24 | 206 | 1747 | TG | clr | *
2016-07-25 | 207 | 1234 | TG/EB | clr | *
2016-07-26 | 208 | 1214 | TG/EB | clr | *
2016-07-27 | 209 | 0 | 0
2016-07-28 | 210 | 1313 | TG/EB | clr | *
2016-07-29 | 211 | 1611 | TG/EB | clr | *
2016-07-30 | 212 | 1236 | TG/EB | clr | *
2016-07-31 | 213 | 1521 | EB | clr | *
2016-08-01 | 214 | 1254 | EB | clr | * | 11 | 5.5
2016-08-02 | 215 | 1259 | EB | clr | *
2016-08-03 | 216 | 1121 | TG | clr | *
2016-08-04 | 217 | 1400 | EB | clr | *
2016-08-05 | 218 | 1353 | EB/TG | clr | *
2016-08-06 | 219 | 1717 | EB | clr | *
2016-08-07 | 220 | 1632 | TG | clr | *
2016-08-08 | 221 | 1541 | TG | clr | * | 11 | 5.5
2016-08-09 | 222 | 1832 | TG | clr | *
2016-08-10 | 223 | 1400 | EB | clr | *
2016-08-11 | 224 | 1243 | EB | clr | *
2016-08-12 | 225 | 1534 | TG | clr | * |  |  | Power outage 1609-1611z. 1642-1644z
2016-08-13 | 226 | 1755 | EB | clr | *
2016-08-14 | 227 | 1759 | EB | clr | *
2016-08-15 | 228 | 1907 | TG/EB | clr | *
2016-08-16 | 229 | 1810 | TG | clr | * | 11 | 5.5
2016-08-17 | 230 | 1652 | TG | clr | *
2016-08-18 | 231 | 1731 | TG | rm snow | *
2016-08-19 | 232 | 1612 | EB/TG | clr | *
2016-08-20 | 233 | 1222 | TG | clr | *
2016-08-21 | 234 | 1545 | EB | rm frost | *
2016-08-22 | 235 | 1634 | TG | rm frost | * | 10.8 | 5.5
2016-08-23 | 236 | 1255 | EB | rm frost | *
2016-08-24 | 237 | 1330 | TG | rm frost | *
2016-08-25 | 238 | 1223 | EB | clr | *
2016-08-26 | 239 | 1352 | TG | rm frost | *
2016-08-27 | 240 | 1744 | EB | clr | *
2016-08-28 | 241 | 1715 | TG | clr | *`
2016-08-29 | 242 | 1354 | EB | clr | * | 10.5 | 5.5
2016-08-30 | 243 | 1300 | TG | rm frost | *
2016-08-31 | 244 | 1306 | EB | clr | *
2016-09-01 | 245 | 1316 | TG | rm frost | *
2016-09-02 | 246 | 1700 | EB | clr | *
2016-09-03 | 247 | 1323 | TG | rm frost | *
2016-09-04 | 248 | 1419 | EB | clr | *
2016-09-05 | 249 | 1250 | TG | rm frost | * | 10 | 5.5
2016-09-06 | 250 | 1342 | EB | clr | *
2016-09-07 | 251 | 1243 | TG | rm | *
2016-09-08 | 252 | 1237 | EB | clr | *
2016-09-09 | 253 | 1249 | TG | rm frost | *
2016-09-10 | 254 | 1233 | EB | clr | *
2016-09-11 | 255 | 1742 | TG | rm frost | *
2016-09-12 | 256 | 1237 | EB | rm frost | * | 10.5 | 5.5
2016-09-13 | 257 | 1237 | TG | rm frost | *
2016-09-14 | 258 | 1246 | EB | clr | *
2016-09-15 | 259 | 1220 | TG | rm frost | *
2016-09-16 | 260 | 1340 | EB | rm frost | *
2016-09-17 | 261 | 1247 | TG | rm frost | *
2016-09-18 | 262 | 1457 | EB | rm frost | *
2016-09-19 | 263 | 1323 | TG | rm frost | * | 9.8 | 5.5
2016-09-20 | 264 | 1324 | EB | rm frost | *
2016-09-21 | 265 | 1226 | TG | rm frost | *
2016-09-22 | 266 | 1318 | EB | rm frost | *
2016-09-23 | 267 | 1216 | TG | rm frost | *
2016-09-24 | 268 | 1409 | EB | clr | *
2016-09-25 | 269 | 1337 | TG | rm frost | *
2016-09-26 | 270 | 1257 | EB | rm frost | * | 10.5 | 5.5
2016-09-27 | 271 | 1207 | TG | rm frost | *
2016-09-28 | 272 | 1325 | EB | rm frost | *
2016-09-29 | 273 | 1247 | TG | rm frost | *
2016-09-30 | 274 | 1402 | EB | clr | *
2016-10-01 | 275 | 1323 | TG | rm frost | *
2016-10-02 | 276 | 1403 | EB | clr | *
2016-10-03 | 277 | 1510 | TG | rm snow | * | 9.8 | 5.5
2016-10-04 | 278 | 1219 | EB | rm frost | *
2016-10-05 | 279 | 1408 | TG | rm hv frost | *
2016-10-06 | 280 | 1314 | EB | rm frost | *
2016-10-07 | 281 | 1303 | TG | rm frost | *
2016-10-08 | 282 | 0 | 0 |  |  |  |  | High winds/low visibility, didn't make it out to TAWO
2016-10-09 | 283 | 1459 | TG | rm frost | *
2016-10-10 | 284 | 1644 | EB | clr | * | 10 | 5.5
2016-10-11 | 285 | 1338 | TG | rm frost | *
2016-10-12 | 286 | 0 | 0
2016-10-13 | 287 | 1303 | TG | rm blwn snow | *
2016-10-14 | 288 | 1716 | EB | rm frost | *
2016-10-15 | 289 | 1300 | TG | rm blwn snow | *
2016-10-16 | 290 | 0 | 0
2016-10-17 | 291 | 1431 | TG | rm frost | * | 10 | 5.5
2016-10-18 | 292 | 1259 | EB | rm frost | *
2016-10-19 | 293 | 1257 | TG | rm frost | *
2016-10-20 | 294 | 0 | 0 |  |  |  |  | Con 1. Didn't go to TAWO.
2016-10-21 | 295 | 1619 | TG | rm blwn snow | *
2016-10-22 | 296 | 1213 | EB | rm frost | *
2016-10-23 | 297 | 1331 | TG | rm frost | *
2016-10-24 | 298 | 1544 | TG/MG/NB | rm frost | *
2016-10-25 | 299 | 1133 | MG/TG | rm lt frost | * | 10 | 5.5
2016-10-26 | 300 | 1213 | TG | rm lt frost | *
2016-10-27 | 301 | 1128 | MG/NB | clr | *
2016-10-28 | 302 | 1135 | MG/NB | rm frost | *
2016-10-29 | 303 | 1150 | MG/NB | rm lt frost | *
2016-10-30 | 304 | 1133 | NB | clr | *
2016-10-31 | 305 | 1318 | NB | rm frost | *
2016-11-01 | 306 | 1253 | MG | rm frost | *
2016-11-02 | 307 | 1251 | NB | clr | *
2016-11-03 | 308 | 1245 | MG | rm frost | * | 10 | 5.5
2016-11-04 | 309 | 1309 | MG | rm lt frost | *
2016-11-05 | 310 | 1230 | MG | rm lt frost | *
2016-11-06 | 311 | 1325 | NB | clr | *
2016-11-07 | 312 | 1230 | MG | rm lt frost | *
2016-11-08 | 313 | 1225 | MG | rm lt frost | *
2016-11-09 | 314 | 1216 | MG | rm lt frost | *
2016-11-10 | 315 | 1242 | MG | rm lt frost | *
2016-11-11 | 316 | [6 empty] | Did not go to TAWO condition 1
2016-11-12 | 317 | 1240 | MG/NB | rm lt frost | * | 9 | 5
2016-11-13 | 318 | 1250 | MG | rm lt frost | *
2016-11-14 | 319 | 1900 | NB | rm frost | *
2016-11-15 | 320 | 1400 | MG | clr | *
2016-11-16 | 321 | 1302 | MG | rm lt frost | *
2016-11-17 | 322 | 1218 | MG/NB | rm lt frost | *
2016-11-18 | 323 | 1323 | MG | rm lt frost | *
2016-11-19 | 324 | 1235 | NB | clr | *
2016-11-20 | 325 | 1312 | NB | rm lt frost | *
2016-11-21 | 326 | 1305 | NB | clr | *
2016-11-22 | 327 | 1301 | MG/NB | rm lt frost | * | 8.7 | 5.5
2016-11-23 | 328 | 1230 | NB | clr | *
2016-11-24 | 329 | 1250 | NB | clr | *
2016-11-25 | 330 | 1249 | MG | rm lt frost | *
2016-11-26 | 331 | 1335 | NB | clr | *
2016-11-27 | 332 | 1420 | MG | clr | *
2016-11-28 | 333 | 1235 | NB | clr | * | 8.3 | 5.5
2016-11-29 | 334 | 1245 | NB | rm frost | * | 7 | 5.5
2016-11-30 | 335 | 1325 | MG | rm lt frost | ~ |  |  | Pump vacuum dropping
2016-12-01 | 336 | 1250 | NB | rm lt frost | * | 11.5 | 5.2 | Swapped to spare vacuum pump 1325-1329z. Vacuum now reading 11.5, flow 5.2.
2016-12-02 | 337 | 1321 | MG | rm lt frost | *
2016-12-03 | 338 | 1235 | NB | rm lt frost | * | 12 | 5.2 | Vacuum and flow look good after pump change.
2016-12-04 | 339 | 1348 | NB | rm lt frost | *
2016-12-05 | 340 | 1242 | NB | clr | *
2016-12-06 | 341 | 1244 | MG | rm lt frost | *
2016-12-07 | 342 | 1240 | MG | rm lt frost | *
2016-12-08 | 343 | 1234 | MG/NB | clr | *
2016-12-09 | 344 | 1328 | MG | rm lt frost | *
2016-12-10 | 345 | 1920 | MG | rm lt frost | *
2016-12-11 | 346 | 1340 | MG | rm lt frost | *
2016-12-12 | 347 | 1230 | NB | rm lt frost | * | 12 | 5.1
2016-12-13 | 348 | 1225 | NB | rm lt frost | *
2016-12-14 | 349 | 1304 | MG | rm lt frost | *
2016-12-15 | 350 | 1415 | MG | rm lt frost | *
2016-12-16 | 351 | 1308 | MG | rm lt frost | *
2016-12-17 | 352 | 1252 | MG | rm lt frost | *
2016-12-18 | 353 | 1328 | NB | clr | *
2016-12-19 | 354 | 1220 | NB | clr | * | 11.3 | 5
2016-12-20 | 355 | 1506 | MG | clr | *
2016-12-21 | 356 | 1325 | MG | clr | *
2016-12-22 | 357 | 1350 | MG | clr | *
2016-12-23 | 358 | 1244 | NB | clr | * |  |  | Power outages for e-gen swap: 1330-1333z and 1645-1648z.
2016-12-24 | 359 | 1326 | NB | clr | *
2016-12-25 | 360 | 1456 | MG | rm lt frost | *
2016-12-26 | 361 | 1235 | NB | clr | * | 12 | 5
2016-12-27 | 362 | 1034 | MG | rm lt frost | *
2016-12-28 | 363 | 1358 | MG | rm lt frost | *
2016-12-29 | 364 | 1308 | MG/NB | clr | *
2016-12-30 | 365 | 1314 | MG/NB | rm frost | *
2016-12-31 | 366 | 1332 | NB | clr | *

## NOAA CATS GC
|  |  |  | Cylinder Pressures | [5 empty] | CATS Gas Chromatograph | [8 empty] | CATS Data Computer | [7 empty] | Chromatograms
Date | Julian Day | Time (UTC) | Observer's Initials | N2 as of 4/30/16 For Channel 4 (previously P5) (high psi) | N2 For Channels 1,2, and 3 (high/flow psi) | N2O (high/low psi) | Cal. 1 (high/flow psi) | Cal. 2 (high/flow psi) | CO2 (high/flow psi) | N2 Purifier Hot? | ECD1 Temp (330) | ECD2 TEMP (280) | ECD3 TEMP (280) | ECD4 TEMP (330) | Column 1 Temp (80C) | Column 2 Temp (85C) | Column 3 Temp (65C) | Column 4 Temp (75C) | Main Ch. 1 Flow (40±5 sccm) | Main Ch. 2 Flow (35±5 sccm) | Main Ch. 3 Flow (34±5 sccm) | Main Ch. 4 Flow (37±5 sccm) | B.F. Ch. 1 Flow (27±5 sccm) | B.F. Ch. 2 Flow ( 25±5 sccm) | B.F. Ch. 3 Flow (26±5 sccm) | B.F. Ch. 4 Flow (27±5 sccm) | Channel 1 Graph | Channel 2 Graph | Channel 3 Graph | Channel 4 Graph | Comments
2016-01-01 | 1 | 1510 | JG | 350 | 450/61 | 630/19 | 1200/24 | 1590/25 |  | * | * | * | 323.2 | * | unsteady | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Channel 3 response rate still high and chromatogram baseline is shifted up.
2016-01-02 | 2 | 1229 | NB | 310 | 380/61 | 630/19 | 1200/24 | 1590/25 |  | * | * | * | * | * | unsteady | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Channel 3 response rate still high and chromatogram baseline is shifted up. NB climbed tower and removed frost from inlets 1404-1450z
2016-01-03 | 3 | 1310 | JG | 270 | 290/61 | 630/19 | 1200/24 | 1590/24 |  | * | * | * | * | * | unsteady | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-01-04 | 4 | 1228 | NB | 210 | 190/61 | 630/19 | 1200/25 | 1590/25 |  | * | * | * | * | * | unsteady | * | * | * | * | * | * | * | * | * | * | * | no | no | no | no | Chromatograms and chroms screen not updating. Stuck on ssv position 1. Web server is down. Air 1 flow: 6.2, Air 2 flow: 6.2, Air 1 BP: 8.1, Air 2 BP: 12.0. JG climbed tower 1355-1417z and removed frost from inlets. PC rebooted at approx 1730z and chroms screen started updating again.
2016-01-05 | 5 | 1230 | JG | 160 | 2520/61 | 630/19 | 1200/25 | 1590/24 |  | * | * | * | unsteady | * | unsteady | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Changed N2 tank 1230-1236z; ending pres 100 psi; new tank starting pres 2520 psi.
2016-01-06 | 6 | 1307 | NB | 110 | 2400/61 | 630/19 | 1200/25 | 1580/27 |  | * | * | * | * | * | unsteady | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Changed P5 tank 1308-1321z; ending pres 110 psi; new tank starting pres 2180 psi.
2016-01-07 | 7 | 1235 | JG | 2100 | 2300/61 | 630/19 | 1190/25 | 1570/27 |  | * | * | * | * | * | unsteady | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-01-08 | 8 | 1347 | NB | 2070 | 2240/61 | 630/19 | 1190/26 | 1570/26 |  | * | * | * | * | * | unsteady | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-01-09 | 9 | 1240 | JG | 2020 | 2120/61 | 630/19 | 1190/25 | 1570/25 |  | * | * | * | * | * | unsteady | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-01-10 | 10 | 1504 | NB | 1950 | 2020/61 | 630/19 | 1190/25 | 1580/28 |  | * | * | * | unsteady | * | unsteady | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | JG climbed tower 1513-1530z and removed frost from inlets.
2016-01-11 | 11 | 1240 | JG | 1920 | 1910/61 | 630/19 | 1190/25 | 1580/28 |  | * | * | * | * | * | unsteady | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 flow: 6.2, Air 2 flow: 6.4, Air 1 BP: 8.1, Air 2 BP: 12.0.
2016-01-12 | 12 | 1232 | NB | 1870 | 1850/61 | 630/19 | 1190/24 | 1580/28 |  | * | * | * | * | * | unsteady | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | JG climbed tower 1243-1307z and removed frost from inlets. Adjusted air 2 flow to 50cc/min at 1707z.
2016-01-13 | 13 | 1255 | JG | 1820 | 1760/61 | 630/19 | 1190/25 | 1580/26 |  | * | * | * | unsteady | * | unsteady | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Planned power outage for e-gen test 1307-1309 and 1321-1324 UTC.
2016-01-14 | 14 | 1225 | NB | 1780 | 1700/61 | 630/19 | 1190/25 | 1580/26 |  | * | * | * | unsteady | * | unsteady | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-01-15 | 15 | 1255 | JG | 1740 | 1590/61 | 630/19 | 1190/25 | 1570/26 |  | * | * | * | * | * | unsteady | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | JG climbed tower 1327-1347 and cleaned frost from inlets.
2016-01-16 | 16 | 1304 | NB | 1690 | 1500/61 | 630/19 | 1190/25 | 1570/26 |  | * | * | * | unsteady | * | unsteady | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-01-17 | 17 | 1725 | JG | 1620 | 1390/61 | 630/19 | 1180/25 | 1560/26 |  | * | * | * | * | * | 85.2 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-01-18 | 18 | 1229 | NB | 1580 | 1340/61 | 630/19 | 1180/25 | 1560/25 |  | * | * | * | * | * | unsteady | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 flow: 6.2, Air 2 flow: 6.3, Air 1 BP: 8.2, Air 2 BP: 11.5.
2016-01-19 | 19 | 1310 | JG | 1540 | 1260/61 | 630/19 | 1180/25 | 1560/25 |  | * | * | * | unsteady | * | unsteady | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-01-20 | 20 | 1232 | NB | 1500 | 1190/61 | 630/19 | 1180/25 | 1560/26 |  | * | * | * | unsteady | * | unsteady | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted cal 2 flow to 50cc/min.
2016-01-21 | 21 | 1250 | JG | 1460 | 1100/61 | 630/19 | 1170/25 | 1560/26 |  | * | * | * | * | * | 85.2 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-01-22 | 22 | 1234 | NB | 1400 | 1000/61 | 630/19 | 1170/25 | 1550/25 |  | * | * | * | unsteady | * | unsteady | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-01-23 | 23 | 1210 | JG | 1360 | 900/61 | 630/19 | 1160/25 | 1550/25 |  | * | * | * | * | * | 85.2 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | NB climbed tower and removed frost from inlets 1217-1234z.
2016-01-24 | 24 | 1258 | NB | 1310 | 810/61 | 630/19 | 1160/25 | 1550/26 |  | * | * | * | unsteady | * | unsteady | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-01-25 | 25 | 1230 | JG | 1260 | 720/31 | 630/19 | 1160/24 | 1550/25 |  | * | * | * | * | * | unsteady | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Channel 3 baseline appears to have shifted up again. Air 1 flow: 6.0, Air 2 flow: 6.3, Air 1 BP: 7.9, Air 2 BP: 11.0.
2016-01-26 | 26 | 1254 | NB | 1210 | 650/61 | 630/19 | 1160/24 | 1550/26 |  | * | * | * | unsteady | * | unsteady | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Channel 3 baseline still high.
2016-01-27 | 27 | 1340 | JG | 1160 | 560/61 | 630/19 | 1160/25 | 1550/25 |  | * | * | * | unsteady | * | unsteady | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-01-28 | 28 | 1234 | NB | 1110 | 490/61 | 630/19 | 1160/25 | 1550/27 |  | * | * | * | unsteady | * | unsteady | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Swapped ECD2 and ECD3 "In" sample lines, and swapped the data acquisition cables on electrometer 2 and electrometer 3. During the swap the crimp on the coaxial cable that connects ECD3 to electrometer 3 contacted the adjacent thermal limit switch and produced an electrical arc that shut down all of the PIDs. A blown AC fuse was found and replaced. GC was restarted. The PIDs for Col1 and Col3 were found to have been damaged as well and were replaced with spares. Also the thermal limit switch on ECD3 was damaged (found in the open position and would not reset). ECD3 was taken offline (unplugged) and should remain offline until Geoff Dutton is able to service it during the summer season. NB
2016-01-29 | 29 | 1255 | JG | 1070 | 390/61 | 630/20 | 1160/25 | 1550/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-01-30 | 30 | 1225 | NB | 1010 | 300/60 | 630/19 | 1160/25 | 1550/28 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-01-31 | 31 | 1455 | JG | 960 | 210/60 | 630/19 | 1150/25 | 1550/25 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-02-01 | 32 | 1225 | NB | 910 | 150/61 | 630/19 | 1150/23 | 1540/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted cal1 flow to 50cc/min. Air 1 flow: 6.1, Air 2 flow: 6.3, Air 1 BP: 8.0, Air 2 BP: 11.2. JG climbed tower 1310-1327z and removed frost from inlets. Swapped N2 tank 1302-1306z, end tank pressure 150psi, start tank pressure 2410psi.
2016-02-02 | 33 | 1240 | JG | 910 | 2340/61 | 630/19 | 1150/24 | 1540/25 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-02-03 | 34 | 1248 | NB | 820 | 2250/61 | 630/19 | 1150/24 | 1530/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-02-04 | 35 | 1220 | JG | 780 | 2160/61 | 630/19 | 1150/25 | 1530/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | NB climbed tower 1224-1239z and removed frost from inlets.
2016-02-05 | 36 | 1333 | NB | 720 | 2050/61 | 630/19 | 1150/24 | 1530/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | JG climbed tower 1342-1359z and removed frost from inlets.
2016-02-06 | 37 | 1355 | JG | 690 | 1960/61 | 630/19 | 1140/24 | 1530/25 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | NB climbed tower 1407-1421z and cleaned frost from inlets.
2016-02-07 | 38 | 1455 | MB/JJ | 640 | 1880/60 | 630/19 | 1150/24 | 1540/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | at 1553 UTC JJ/MB adjusted Cal 2 flow to 50 cc/min.
2016-02-08 | 39 | 1240 | MB/JJ/SD | 600 | 1780/60 | 630/20 | 1140/23 | 1530/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 flow: 6.0, Air 2 flow: 6.2, Air 1 BP: 7.9, Air 2 BP: 11.1. Sample flows all 50cc/min
2016-02-09 | 40 | 1230 | MB/JJ/NB | 530 | 1680/61 | 630/19 | 1140/24 | 1520/28 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-02-10 | 41 | 1226 | SWD/JG/NB | 500 | 1600/61 | 630/20 | 1140/22 | 1520/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-02-11 | 42 | 1340 | MB/JJ | 420 | 1500/61 | 630/19 | 1140/23 | 1520/24 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Cal2 flow to 50cc/min. Tightened fittings on ECD4 OUT and P5 inlet relay 2. Snoop tested and saw no bubbles.
2016-02-12 | 43 | 1625 | MB/JJ | 380 | 1400/61 | 630/19 | 1140/25 | 1520/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Cal1 flow to 50cc/min. Adjusted Air1 flow to 50cc/min. Re-ran P5 exhaust through top of blue instrument housing. Placed computer CPU into rack.
2016-02-13 | 44 | 1220 | MB/JJ | 330 | 1310/61 | 630/19 | 1140/24 | 1520/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-02-14 | 45 | 1700 | MB/JJ | 290 | 1210/61 | 620/19 | 1140/23 | 1510/25 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-02-15 | 46 | 1641 | JJ | 230 | 1140/60 | 630/19 | 1130/22 | 1510/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted cal 1 flow to 50cc/min. Air 1 flow: 6.0, Air 2 flow: 6.0, Air 1 BP: 7.7, Air 2 BP: 11.1.
2016-02-16 | 47 | 1321 | jj/mb | 190 | 1090/60 | 640/19 | 1140/23 | 1520/25 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Preformed leak test on Channel4 manifold, found leaks in the 4-, 5-, and 10-o'clock positions, tightened fittings.
2016-02-17 | 48 | 1415 | MB/JJ | 120 | 990/60 | 640/19 | 1140/24 | 1520/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | JJ climbed tower 1428-1455z and removed frost from instruments. Changed P5 tank 1514-1515z; ending pres 120 psi; new tank starting pres 2140 psi.
2016-02-18 | 49 | 1405 | MB/JJ | 2110 | 980/61 | 630/19 | 1120/24 | 1510/24 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-02-19 | 50 | 1349 | JJ/MB | 2080 | 800/61 | 630/19 | 1120/24 | 1500/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-02-20 | 51 | 1311 | JJ/MB | 2020 | 710/61 | 630/19 | 1130/24 | 1510/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | JJ climbed tower 1332-1358z and 1457-1526z to troubleshoot 2m RTD cable.
2016-02-21 | 52 | 1510 | JJ/MB | 1980 | 620/61 | 630/18 | 1110/22 | 1500/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-02-22 | 53 | 1244 | JJ/MB | 1920 | 550/61 | 630/19 | 1110/24 | 1500/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 flow: 6.0, Air 2 flow: 6.5, Air 1 BP: 7.8, Air 2 BP: 11.2. Adjusted Cal1 and Cal2 to 50cc/min. All other sample flows all 50cc/min.
2016-02-23 | 54 | 1704 | MB | 1880 | 450/60 | 630/19 | 1120/24 | 1500/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-02-24 | 55 | 1314 | JJ/MB | 1820 | 390/60 | 630/19 | 1110/22 | 1490/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-02-25 | 56 | 1336 | jj/mb | 1800 | 300/60 | 630/19 | 1110/24 | 1500/25 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Changed N2 mol sieve trap to stored trap at 1353z. Began N2 trap bake at 1425z. Adjusted Cal2 to 50cc/min.
2016-02-26 | 57 | 1348 | JJ/MB | 1750 | 200/61 | 630/19 | 1110/23 | 1490/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-02-27 | 58 | 1305 | MB | 1700 | 100/61 | 630/20 | 1110/24 | 1500/28 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Changed N2 tank 1327-1331z; ending pres 100 psi; new tank starting pres 2410 psi. Finished baking spare N2 trap ar 1335z. SWD climbed tower to clean instruments and troubleshoot 2m-RTD cable from 1444-1542z.
2016-02-28 | 59 | 1707 | JJ | 1630 | 2380/61 | 630/19 | 1100/24 | 1490/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-02-29 | 60 | 1339 | jj/mb | 1600 | 2300/60 | 630/19 | 1100/23 | 1480/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 flow: 6.0, Air 2 flow: 6.5, Air 1 BP: 7.9, Air 2 BP: 11.2. Adjusted Air1, Air 2 and Cal2 to 50cc/min.
2016-03-01 | 61 | 1248 | JJ/MB | 1570 | 2210/60 | 630/19 | 1100/23 | 1490/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-03-02 | 62 | 1711 | MB/JJ | 1510 | 2110/60 | 630/19 | 1100/24 | 1490/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Cal1 flow to 50cc/min. JJ climbed tower to clean instruments from 1737-1833z.
2016-03-03 | 63 | 1338 | MB/JJ | 1480 | 2010/60 | 630/19 | 1100/23 | 1490/25 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-03-04 | 64 | 1309 | jj/mb | 1420 | 1970/60 | 630/19 | 1100/24 | 1490/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-03-05 | 65 | 1235 | jj/mb | 1400 | 1890/60 | 630/19 | 1100/24 | 1490/24 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-03-06 | 66 | 1342 | MB | 1330 | 1780/60 | 630/19 | 1100/23 | 1490/25 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-03-07 | 67 | 1321 | jj/mb | 1300 | 1690/60 | 630/19 | 1100/24 | 1490/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 flow: 6.0, Air 2 flow: 6.5, Air 1 BP: 7.8, Air 2 BP: 11.2. Adjusted Cal 1 and Cal 2 to 50cc/min.
2016-03-08 | 68 | 1331 | jj/mb | 1230 | 1590/60 | 630/19 | 1090/25 | 1480/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Cal1 to 50 cc/min. Adjusted Air1 to 50 cc/min.
2016-03-09 | 69 | 1235 | JJ | 1200 | 1500/60 | 630/19 | 1090/23 | 1480/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-03-10 | 70 | 0 | 0 | [27 empty] | White out conditions (Condition 1) did not travel to TAWO this day.
2016-03-11 | 71 | 1352 | JJ | 1100 | 1320/60 | 630/19 | 1080/23 | 1470/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-03-12 | 72 | 1437 | MB | 1060 | 1250/60 | 630/19 | 1080/24 | 1470/28 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-03-13 | 73 | 1346 | JJ | 1010 | 1180/60 | 630/19 | 1080/24 | 1470/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | unexpected power outage today from 13:16 to 13:23 UTC
2016-03-14 | 74 | 1353 | jj/mb | 970 | 1090/59 | 630/19 | 1070/24 | 1460/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 flow: 6.2, Air 2 flow: 6.2, Air 1 BP: 7.8, Air 2 BP: 11.2. Adjusted Cal1, Cal2, Air1 and Air2 to 50cc/min.
2016-03-15 | 75 | 1533 | jj/mb | 900 | 1000/59 | 630/19 | 1070/24 | 1460/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Cal2 to 50 cc/min.
2016-03-16 | 76 | 1319 | MB | 890 | 900/60 | 630/19 | 1080/25 | 1460/29 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-03-17 | 77 | 1338 | jj/mb | 820 | 820/60 | 630/19 | 1080/25 | 1450/29 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-03-18 | 78 | 1305 | MB/JJ | 780 | 750/60 | 630/19 | 1070/24 | 1450/31 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Software found frozen, rebooted computer at 1305z. Adjusted Cal1 to 50cc/min.
2016-03-19 | 79 | 1854 | MB/JJ | 720 | 620/60 | 630/19 | 1060/24 | 1450/30 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-03-20 | 80 | 1334 | MB | 690 | 590/60 | 630/19 | 1060/24 | 1450/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-03-21 | 81 | 1307 | jj/mb | 630 | 490/60 | 630/19 | 1060/24 | 1450/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 flow: 6.5, Air 2 flow: 6.2, Air 1 BP: 8.2, Air 2 BP: 11.6. Adjusted Cal2, Air1 to 50cc/min. All other flows 50cc/min.
2016-03-22 | 82 | 1320 | MB | 600 | 400/60 | 630/19 | 1060/24 | 1450/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-03-23 | 83 | 1744 | MB/JJ | 520 | 300/60 | 630/19 | 1060/24 | 1450/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Cal2 to 50 cc/min. adjusted Cal1 to 50cc/min.
2016-03-24 | 84 | 1329 | jj/mb | 480 | 220/60 | 630/19 | 1050/25 | 1450/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-03-25 | 85 | 1303 | JJ | 420 | 150/60 | 630/19 | 1050/24 | 1440/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Changed N2 tank 1441z ending pressure 150psi new tank starting pressure 2490 psi.
2016-03-26 | 86 | 1303 | MB | 390 | 2400/60 | 630/19 | 1050/24 | 1440/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-03-27 | 87 | 1715 | JJ | 320 | 2300/60 | 630/19 | 1050/24 | 1440/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-03-28 | 88 | 1327 | MB | 300 | 2220/60 | 630/19 | 1050/24 | 1440/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 flow: 6.0, Air 2 flow: 6.0, Air 1 BP: 8.1, Air 2 BP: 11.5. Adjusted Cal1 and Cal2 to 50cc/min. All other flows 50cc/min.
2016-03-29 | 89 | 1151 | JJ | 220 | 2120/60 | 630/19 | 1050/23 | 1440/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-03-30 | 90 | 1134 | MB | 190 | 2040/60 | 630/19 | 1050/23 | 1440/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-03-31 | 91 | 1718 | jj/mb | 120 | 1970/60 | 630/19 | 1050/23 | 1440/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Changed P5 tank 1722-1725z; ending pres 120 psi; new tank starting pres 2150 psi.
2016-04-01 | 92 | 1304 | JJ | 2150 | 1900/60 | 630/19 | 1050/24 | 1440/28 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-04-02 | 93 | 1213 | jj/mb | 2100 | 1800/60 | 630/19 | 1050/24 | 1440/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-04-03 | 94 | 1550 | MB | 2040 | 1690/60 | 630/19 | 1040/24 | 1430/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-04-04 | 95 | 1613 | jj/mb | 2000 | 1600/60 | 630/19 | 1040/24 | 1420/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 flow: 6.1, Air 2 flow: 6.2, Air 1 BP: 8.2, Air 2 BP: 11.6. Adjusted Cal2, Air 2 to 50cc/min. All other flows 50cc/min. Removed activated charcol carbon trap from the P5 line and started to bake. P5 line on GC was reconnected without this trap.
2016-04-05 | 96 | 1251 | MB/JJ | 1950 | 1510/60 | 630/19 | 1040/24 | 1430/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-04-06 | 97 | 1332 | JJ | 1900 | 1420/60 | 630/19 | 1030/23 | 1410/25 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Activated carbon charcol trap was reinstalled on the P5 line after it was reconditioned. JJ Planned power outage today
2016-04-07 | 98 | 1224 | MB/JJ | 1870 | 1350/60 | 630/19 | 1030/24 | 1420/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-04-08 | 99 | 1139 | JJ/MB | 1820 | 1280/60 | 630/19 | 1030/24 | 1420/25 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | SWD climbed the tower from 1715-1800z to measure 10m and 16m windbird alignment.
2016-04-09 | 100 | 1254 | JJ/MB | 1800 | 1200/60 | 630/19 | 1030/24 | 1410/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-04-10 | 101 | 1707 | JJ | 1720 | 1100/60 | 630/19 | 1020/25 | 1400/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Channel 1 baseline shiffted down.
2016-04-11 | 102 | 1321 | JJ/MB | 1700 | 1010/60 | 630/19 | 1030/25 | 1420/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 flow: 6.3, Air 2 flow: 6.2, Air 1 BP: 8.4, Air 2 BP: 11.6. Adjusted Cal1 and Air 1 to 50cc/min. All other flows 50cc/min.
2016-04-12 | 103 | 1554 | JJ/MB | 1680 | 920/60 | 630/19 | 1040/24 | 1430/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Cal2 to 50cc/min.
2016-04-13 | 104 | 1317 | JJ/MB | 1600 | 820/60 | 630/19 | 1030/25 | 1410/30 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Cal2 to 50cc/min.
2016-04-14 | 105 | 1329 | JJ/MB | 1550 | 780/60 | 630/19 | 1030/22 | 1410/28 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | JJ climbed tower from 1356-1422z to work on surface ozone inlet and clean instruments.
2016-04-15 | 106 | 1711 | JJ/MB/SWD | 1500 | 670/59 | 630/19 | 1025/23 | 1410/25 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted cal 2 flow to 50cc/min.
2016-04-16 | 107 | 1309 | JJ/MB | 1480 | 600/60 | 630/19 | 1020/23 | 1410/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Air2 and Cal1 to 50cc/min.
2016-04-17 | 108 | 1628 | MB | 1410 | 500/60 | 630/19 | 1010/24 | 1400/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-04-18 | 109 | 1249 | MB | 1380 | 410/60 | 630/19 | 1020/24 | 1400/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 flow: 6.2, Air 2 flow: 6.0, Air 1 BP: 8.1, Air 2 BP: 11.5. Adjusted Cal2 and Air1 to 50cc/min. All other flows 50cc/min.
2016-04-19 | 110 | 1152 | JJ/MB | 1320 | 320/60 | 630/19 | 1020/23 | 1410/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-04-20 | 111 | 1143 | JJ/MB | 1300 | 250/60 | 630/19 | 1010/22 | 1400/25 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-04-21 | 112 | 1122 | JJ/MB/SWD | 1230 | 150/60 | 630/19 | 1010/25 | 1400/28 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | 1354-1358z changed N2 tanks. Ending tank pressure 150 psi. Beginning tank pressure 2410 psi. 1426-1458z and 1513-1527z SWD climbed tower to clean instruments and realign 10m windbird.
2016-04-22 | 113 | 1838 | JJ/MB | 1200 | 2420/60 | 630/19 | 1010/25 | 1390/26 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted cal 2 flow to 50cc/min.
2016-04-23 | 114 | 1536 | 0 | [27 empty] | White out conditions (Condition 1) did not travel to TAWO this day.
2016-04-24 | 115 | 1536 | JJ/MB | 1100 | 2240/60 | 630/19 | 1000/22 | 1390/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | adjusted Cal 1 flow to 50cc/min
2016-04-25 | 116 | 1207 | JJ/MB | 1060 | 2180/60 | 630/19 | 1000/24 | 1380/25 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 flow: 6.2, Air 2 flow: 5.8, Air 1 BP: 8.5, Air 2 BP: 11.7. Adjusted Air 1 and 2 to 50cc/min. All other flows 50cc/min. Baseline dropped on channel 2.
2016-04-26 | 117 | 1130 | JJ | 1010 | 2050/60 | 630/19 | 1000/25 | 1380/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-04-27 | 118 | 1547 | JJ/MB | 970 | 1950/60 | 630/19 | 1000/24 | 1390/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | 1839z Began baking P5 carbon trap.
2016-04-28 | 119 | 1315 | JJ/MB | 910 | 1890/60 | 630/19 | 1000/24 | 1390/27 |  | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | 1255-1305z JJ climbed the tower to clean instruments. 2138z completed baking carbon trap. 2248z turned EDC 4 set point to 20 C in preparation for CO2-N2 switch from P5.
2016-04-29 | 120 | 1300 | MB/JJ | 890 | 1790/60 | 630/19 | 1000/25 | 1380/24 |  | * | * | * | offline | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Cal2 and Air2 to 50cc/min.
2016-04-30 | 121 | 1336 | MB/JJ | 2590 | 1700/60 | 630/19 | 1000/25 | 1380/26 | 810/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | 1235z removed P5 tank, ending pressure 850psi. Installed N2 (initial pressure 2590 psi) and CO2 (initial pressure 810 psi). ** Channel 4 graph is populating with first run on CO2-N2, first runs appear abnormal, will monitor.
2016-05-01 | 122 | 1624 | MB/JJ | 2550 | 1600/60 | 630/19 | 1000/23 | 1390/26 | 810/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | 1743-1815z JJ and MB climbed tower to replace filter on CUGC sample inlet.
2016-05-02 | 123 | 1322 | JJ/MB | 2500 | 1520/60 | 630/19 | 1000/23 | 1380/27 | 810/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 flow: 6.0, Air 2 flow: 5.5, Air 1 BP: 8.1, Air 2 BP: 11.5. Adjusted Cal1,Cal2, Air1, and Air2 to 50cc/min. Planned power outages starting from 1635 to 1700z
2016-05-03 | 124 | 1230 | MB/JJ | 2460 | 1420/60 | 630/19 | 990/24 | 1380/25 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | 1153-1217z JJ climbed tower to clean instruments and attach ties to CUGC intake hose. Reconditioning former P5 mole sieve to be N2 mole sieve only. Began baking at 1255z.
2016-05-04 | 125 | 1451 | MB/JJ | 2420 | 1350/60 | 630/19 | 990/24 | 1380/25 | 820/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-05-05 | 126 | 1528 | MB | 2350 | 1270/60 | 630/19 | 990/24 | 1380/27 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-05-06 | 127 | 1845 | JJ/MB | 2300 | 1180/60 | 630/19 | 990/24 | 1370/25 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-05-07 | 128 | 1315 | JJ/MB | 2270 | 1100/60 | 630/19 | 970/24 | 1370/25 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted cal 2 flow to 50cc/min.
2016-05-08 | 129 | 1558 | JJ | 2220 | 1110/60 | 630/19 | 970/24 | 1370/26 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-05-09 | 130 | 1635 | JJ/MB | 2200 | 980/60 | 630/19 | 1000/25 | 1370/26 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 flow: 6.2, Air 2 flow: 5.0, Air 1 BP: 8.0, Air 2 BP: 11.4. Adjusted Cal 1 and 2 and Air 2 to 50cc/min. Air 1 is 50cc/min.
2016-05-10 | 131 | 1648 | JJ/MB | 2130 | 850/60 | 630/19 | 980/25 | 1370/27 | 820/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Cal1 to 50cc/min.
2016-05-11 | 132 | 1753 | MB/JJ | 2100 | 780/60 | 630/19 | 975/25 | 1370/26 | 810/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Cal2 to 50cc/min.
2016-05-12 | 133 | 1622 | MB/JJ | 2040 | 690/60 | 630/19 | 970/25 | 1360/27 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Cal1 to 50cc/min.
2016-05-13 | 134 | 1308 | MB/JJ | 2000 | 600/60 | 630/19 | 960/23 | 1360/25 | 810/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-05-14 | 135 | 1149 | JJ/MB | 1980 | 510/60 | 630/19 | 970/23 | 1360/26 | 810/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-05-15 | 136 | 1730 | MB | 1910 | 410/60 | 630/19 | 970/23 | 1350/26 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Air1 to 50cc/min.
2016-05-16 | 137 | 1609 | JJ/MB | 1880 | 320/60 | 630/19 | 970/23 | 1350/26 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 flow: 6.2, Air 2 flow: 5.0, Air 1 BP: 8.4, Air 2 BP: 11.8. Adjusted Cal 1 and 2 to 50cc/min. All other flows 50cc/min.
2016-05-17 | 138 | 1722 | JJ/MB | 1820 | 240/60 | 630/19 | 960/23 | 1350/24 | 810/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-05-18 | 139 | 1309 | JJ/MB | 1800 | 170/60 | 630/19 | 960/23 | 1350/24 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Cal 1 to 50cc/min
2016-05-19 | 140 | 1322 | JJ | 1770 | 100/60 | 630/19 | 950/23 | 1350/27 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Changed N2 tank for channed 1,2,3
2016-05-20 | 141 | 1850 | JJ/MB | 1710 | 2390/60 | 630/19 | 950/23 | 1350/27 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Air2 to 50cc/min.
2016-05-21 | 142 | 1228 | JJ/MB | 1690 | 2310/60 | 630/19 | 960/23 | 1350/27 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-05-22 | 143 | 2327 | JJ | 1610 | 2200/60 | 630/19 | 950/23 | 1350/27 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-05-23 | 144 | 1656 | JJ | 1550 | 2110/60 | 630/19 | 950/24 | 1340/27 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 flow: 6.2, Air 2 flow: 4.5, Air 1 BP: 8.3, Air 2 BP: 11.7. All flows 50cc/min.
2016-05-24 | 145 | 1211 | JJ/MB | 1560 | 2050/60 | 630/19 | 950/24 | 1350/27 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Cal2 and Air2 to 50cc/min.
2016-05-25 | 146 | 1726 | JJ | 1500 | 1910/60 | 630/19 | 950/21 | 1340/25 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-05-26 | 147 | 1232 | JJ/MB | 1490 | 1860/60 | 630/19 | 950/23 | 1340/25 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-05-27 | 148 | 1845 | JJ/mb | 1420 | 1780/60 | 630/19 | 950/23 | 1340/27 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-05-28 | 149 | 1221 | MB | 1400 | 1700/60 | 630/19 | 950/23 | 1340/26 | 820/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Air1, Cal2, and Air2 to 50cc/min.
2016-05-29 | 150 | 1530 | MB | 1360 | 1600/60 | 630/19 | 950/23 | 1340/28 | 820/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Cal1 to 50cc/min.
2016-05-30 | 151 | 1226 | JJ | 1320 | 1510/60 | 630/19 | 940/24 | 1340/27 | 820/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 flow: 6.0, Air 2 flow: 5.2, Air 1 BP: 8.5, Air 2 BP: 11.8. Adjusted Cal 2 to 50cc/min All other flows 50cc/min. Planned power outage today from 2130 to 2230
2016-05-31 | 152 | 1731 | JJ | 1280 | 1410/60 | 630/19 | 940/25 | 1320/28 | 810/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Baseline dropped on channel 2.
2016-06-01 | 153 | 1840 | JJ | 1220 | 1320/60 | 630/19 | 940/25 | 1330/27 | 820/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-06-02 | 154 | 1831 | MB/JJ | 1200 | 1260/60 | 630/19 | 940/24 | 1330/27 | 820/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-06-03 | 155 | 1737 | MB/JJ/HJ | 1190 | 1200/59 | 630/19 | 960/25 | 1360/27 | 920/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-06-04 | 156 | 1156 | MB | 1120 | 1100/60 | 630/20 | 930/24 | 1310/31 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Cal1 to 50cc/min.
2016-06-05 | 157 | 1403 | JJ/TG/EB | 1080 | 1000/60 | 630/19 | 930/24 | 1310/29 | 810/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-06-06 | 158 | 1613 | MB/JJ/TG/EB | 1030 | 900/60 | 620/19 | 940/25 | 1310/27 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 flow: 6.0, Air 2 flow: 5.5, Air 1 BP: 8.4, Air 2 BP: 11.75. Adjusted Air 1 to 50cc/min, Cal 2 to 50cc/min, Air 2 to 50 cc/min, Cal 1 to 50 cc/min.
2016-06-07 | 159 | 1536 | TG/EB/JJ | 1000 | 800/59 | 630/19 | 925/24 | 1310/25 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Cal2 to 50cc/min. Air 1 valve left open after HATS sampling; resealed 1350 06/08/16.
2016-06-08 | 160 | 1309 | TG/EB/JJ/MB | 980 | 720/60 | 640/19 | 940/25 | 1320/28 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Cal2 to 50cc/min, adjusted Cal1 to 50cc/min.
2016-06-09 | 161 | 1553 | TG/EB | 910 | 620/60 | 620/19 | 910/24 | 1310/25 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-06-10 | 162 | 1239 | TG/EB | 900 | 600/59 | 630/19 | 940/24 | 1340/26 | 850/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Cal2 to 50cc/min, adjusted Air2 to 50cc/min. Work on ventilation duct from 1105 to 1330, higher room temps.
2016-06-11 | 163 | 1203 | TG/EB | 840 | 490/60 | 620/19 | 910/25 | 1310/26 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-06-12 | 164 | 1509 | TG/EB | 800 | 400/60 | 630/20 | 920/25 | 1300/30 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-06-13 | 165 | 1637 | TG/EB | 780 | 300/60 | 620/19 | 910/25 | 1310/25 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | BP1: 8.2, Air1: 6.0. BP2: 11.5, Air1: 5.5. Adjusted Air 2 to 50cc/min, adjusted Cal1 to 50 cc/min, adjusted Air 1 to 50 cc/min, adjusted Cal2 to 50cc/min.
2016-06-14 | 166 | 1153 | TG/EB | 740 | 210/60 | 620/20 | 910/22 | 1305/29 | 810/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Cal2 to 50cc/min.
2016-06-15 | 167 | 1202 | TG/EB | 700 | 2600/61 | 610/20 | 900/25 | 1290/26 | 780/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Changed Ch1,2,3 N2 Cylinder: old cylinder final psi=100, new cylinder initial psi=2600
2016-06-16 | 168 | 1156 | TG | 650 | 2390/61 | 610/20 | 900/23 | 1290/28 | 720/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Cal2 to 50cc/min
2016-06-17 | 169 | 1314 | TG | 620 | 2340/60 | 620/19 | 900/25 | 1300/25 | 790/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-06-18 | 170 | 1241 | TG/EB | 600 | 2240/60 | 620/20 | 900/24 | 1300/28 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-06-19 | 171 | 1808 | EB | 600 | 2120/60 | 610/20 | 910/22 | 1300/30 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-06-20 | 172 | 1812 | TG/EB | 520 | 2020/60 | 620/19 | 900/24 | 1300/29 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Cal1 to 50cc/min, Air 1 to 50cc/min, Cal2 to 50cc/min, Air 2 to 50cc/min. Adjustments were made but flows and air were not recorded
2016-06-21 | 173 | 1204 | TG/EB | 480 | 1980/60 | 620/19 | 900/25 | 1290/23 | 810/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-06-22 | 174 | 1332 | TG/EB | 450 | 1870/60 | 620/19 | 900/25 | 1290/27 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-06-23 | 175 | 1313 | EB | 410 | 1800/60 | 625/20 | 900/25 | 1290/27 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-06-24 | 176 | 1329 | TG | 350 | 1700/60 | 620/19 | 900/25 | 1290/27 | 810/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-06-25 | 177 | 1717 | EB/TG | 300 | 1610/60 | 620/19 | 900/25 | 1290/26 | 810/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-06-26 | 178 | 1828 | TG | 270 | 1500/60 | 620/19 | 900/24 | 1270/25 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-06-27 | 179 | 1603 | EB/TG | 220 | 1420/60 | 620/19 | 900/25 | 1280/27 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | BP1: 8.5, Air1: 6.0, BP2: 11.8, Air2: 5.0. Air2 was already at 50cc/min, Adjusted Cal1 to 50cc/min, Adjusted Air1 to 50cc/min, Adjusted Cal2 to 50cc/min
2016-06-28 | 180 | 1452 | EB/TG | 200 | 1325/60 | 640/20 | 890/24 | 1260/25 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-06-29 | 181 | 1135 | EB/TG | 150 | 1300/60 | 620/19 | 900/24 | 1270/24 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Cal2 to 50cc/min
2016-06-30 | 182 | 1143 | TG/EB | 100 | 1200/60 | 620/20 | 890/24 | 1270/27 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | changed N2 channel 4, 1217z, now at 2720psi
2016-07-01 | 183 | 1202 | TG/EB | 2410 | 1100/59 | 620/19 | 880/24 | 1270/27 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-07-02 | 184 | 1239 | TG/EB | 2250 | 1000/59 | 620/19 | 900/22 | 1270/27 | 810/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-07-03 | 185 | 1835 | EB | 2050 | 900/60 | 620/20 | 890/25 | 1260/28 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-07-04 | 186 | 1029 | TG | 1960 | 840/60 | 620/19 | 890/25 | 1260/30 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air1: 6. Air2: 4.6. BP1: 8.6. BP2: 11.9. Adusted Cal1 to 50cc/min, Adjusted Air1 to 50cc/min, adjusted Cal2 to 50cc/min, adjusted Air2 to 50cc/min
2016-07-05 | 187 | 1207 | TG/EB | 1820 | 780/60 | 620/19 | 890/25 | 1260/30 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-07-06 | 188 | 0.47153 | TG | 1610 | 700/60 | 620/19 | 890/24 | 1260/27 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-07-07 | 189 | 1214 | TG/EB | 1600 | 600/59 | 620/19 | 890/25 | 1260/27 | 800/21 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-07-08 | 190 | 1604 | TG/EB | 1500 | 500/60 | 620/19 | 880/25 | 1250/28 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-07-09 | 191 | 1342 | TG | 1420 | 400/61 | 620/20 | 880/25 | 1260/29 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-07-10 | 192 | 1319 | TG | 1350 | 320/59 | 620/19 | 870/25 | 1250/28 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-07-11 | 193 | 1323 | EB | 1310 | 250/60 | 620/19 | 860/25 | 1250/28 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Line flow Air1: 6.0, Line flow Air2: 5.0, Back pressure Air1: 8.8, Back pressure Air2: 11.9; Air2 at 50cc/min, Cal1 at 50cc/min, Air1 at 50cc/min, Cal2 adjusted to 50cc/min.
2016-07-12 | 194 | 1228 | TG/EB | 1260 | 150/60 | 620/19 | 860/25 | 1250/26 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-07-13 | 195 | 1137 | TG | 1210 | 2550/60 | 620/19 | 860/24 | 1250/26 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | ? | ? | * | * | changed N2 channels 123 1203-1219z, now at 2450psi adjusted cal2 to 50cc/min
2016-07-14 | 196 | 1147 | TG/EB | 1200 | 2440/60 | 620/19 | 860/25 | 1250/28 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | ? | ? | * | * | graphs are a little off, will see if adjusting flow rates helps; Adjusted Cal1 to 50cc/min, Air2 to 50cc/min, Cal2 to 50cc/min
2016-07-15 | 197 | 1235 | TG/EB | 1150 | 2340/60 | 620/19 | 860/24 | 1250/25 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-07-16 | 198 | 1301 | TG/EB | 1000 | 2230/60 | 610/20 | 850/25 | 1240/26 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | graph issue was not with the instrument but the scale of the graphs
2016-07-17 | 199 | 1627 | EB | 1050 | 2110/60 | 610/20 | 860/24 | 1240/26 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-07-18 | 200 | 1207 | TG | 1010 | 2090/60 | 620/19 | 850/24 | 1240/26 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air1: 6, Air2: 5, BP1: 8.6, BP2: 11.8 in and out of N winds, but construction has continued on station Adjusted Cal1, Air1, and Air2 to 50cc/mini power outage 2048-54z
2016-07-19 | 201 | 1237 | TG | 970 | 1950/60 | 620/19 | 850/24 | 1240/25 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Main and B.F. numbers were in an okay range, but seemed to be fluctuating with greater variance than normal
2016-07-20 | 202 | 1609 | TG/EB | 920 | 1820/60 | 620/19 | 850/25 | 1240/26 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-07-21 | 203 | 1224 | TG/EB | 900 | 1770/60 | 600/19 | 850/24 | 1240/26 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-07-22 | 204 | 1141 | TG | 860 | 1700/59 | 620/19 | 850/25 | 1240/27 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | contacted Geoff regarding Main and B.F. number flucutations
2016-07-23 | 205 | 1317 | TG | 810 | 1590/60 | 620/19 | 850/24 | 1230/25 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Main and B.F. numbers are sometimes momentarily out of range and still fluctuating greatly. In contact with Geoff
2016-07-24 | 206 | 1747 | TG | 750 | 1490/59 | 620/19 | 840/25 | 1220/27 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Main and B.F. numbers are sometimes momentarily out of range and still fluctuating greatly. In contact with Geoff
2016-07-25 | 207 | 1234 | TG/EB | 710 | 1400/59 | 620/19 | 840/23 | 1220/25 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Main and B.F. numbers are sometimes momentarily out of range and still fluctuating greatly. In contact with Geoff. Adjusted Cal1 to 50cc/min, adjusted Air1 to 50cc/min, adjusted Cal2 to 50cc/min, Air2 at 50cc/min. Air1 Flow: 5.8, Air2 Flow: 12.0, Air1 BP: 8.8, Air2 BP: 4.8.
2016-07-26 | 208 | 1214 | TG/EB | 695 | 1300/60 | 620/19 | 840/25 | 1220/28 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Main and B.F. numbers are sometimes momentarily out of range and still fluctuating greatly. In contact with Geoff.
2016-07-27 | 209 | 0 | 0
2016-07-28 | 210 | 1313 | TG/EB | 600 | 1110/60 | 620/19 | 840/24 | 1220/26 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Main and B.F. numbers are sometimes momentarily out of range and still fluctuating greatly. In contact with Geoff.
2016-07-29 | 211 | 1611 | TG/EB | 560 | 1000/60 | 620/19 | 840/25 | 1210/27 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Main and B.F. numbers are sometimes momentarily out of range and still fluctuating greatly. In contact with Geoff.
2016-07-30 | 212 | 1236 | TG/EB | 500 | 920/59 | 620/19 | 840/24 | 1220/27 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Main and B.F. numbers are sometimes momentarily out of range and still fluctuating greatly. In contact with Geoff.
2016-07-31 | 213 | 1521 | EB | 470 | 800/59 | 630/19 | 840/24 | 1220/26 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Main and B.F. numbers are sometimes momentarily out of range and still fluctuating greatly. Ch 3 BF very out of range: 1.1-2.7. In contact with Geoff. Checked but didn't need to adjust Air1, Cal2, Air2.
2016-08-01 | 214 | 1254 | EB | 420 | 750/60 | 620/19 | 830/24 | 1210/26 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Main and B.F. numbers are further out of range and still fluctuating greatly. Ch 3 BF very out of range (1.1-2.7), and Ch 3 graph slightly offset. In contact with Geoff. Checked Air1,Cal1, Air2, Cal2. LF Air1=5.0, LF Air2=12.0, BP Air1=8.8, BP Air2=4.7.
2016-08-02 | 215 | 1259 | EB | 400 | 700/60 | 620/19 | 840/24 | 1210/26 | 800/20 | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Rebooted computer, GC turned off from 1844 to 1900 and 2010 to 2011 UTC
2016-08-03 | 216 | 1121 | TG | 360 | 700/59 | 620/19 | 830/24 | 1210/25 | 800/20 | * | * | * | offline | * | * | * | * | * | x | x | x | x | x | x | x | x | x | x | x | x | Turned off GC 1255-1257z as per Geoff's instructions (different than manual). Quit GC software at 1255z waited ~30 mins. Rebooted computer at 1327z
2016-08-04 | 217 | 1400 | EB | 360 | 600/59 | 610/19 | 830/25 | 1210/28 | 800/20 | * | * | * | offline | * | * | * | * | * | x | x | x | x | x | x | x | x | x | x | x | x | Stopped software (GC was not measuring anyway) at 1330z to check voltages/flows. 2025z ECDs lowered to 40C.
2016-08-05 | 218 | 1353 | EB/TG | 350 | 700/60 | 620/20 | 830/23 | 1210/24 | 800/20 | * | x | x | offline | x | * | * | * | * | x | x | x | x | x | x | x | x | x | x | x | x | GC measurements stopped, ECDs turned down to 40C. Completely shut down GC at 1528z.
2016-08-06 | 219 | 1717 | EB | 350 | 700/60 | 600/21 | 820/28 | 1140/30 | 800/20 | * | x | x | offline | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | Instrument off; carrier gas, CO2, and N2O cylinders left open.
2016-08-07 | 220 | 1632 | TG | 350 | 690/59 | 600/21 | 820/28 | 1210/34 | 810/20 | * | x | x | offline | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | Instrument off; carrier gas, CO2, and N2O cylinders left open.
2016-08-08 | 221 | 1541 | TG | 350 | 690/60 | 600/22 | 820/27 | 1200/32 | 800/20 | * | x | x | offline | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | Instrument off; carrier gas, CO2, and N2O cylinders left open. Air1: 4.7, Air2, 4.5, BP1: 9.1, BP2: 12
2016-08-09 | 222 | 1832 | TG | 350 | 690/60 | 610/21 | 810/27 | 1200/34 | 800/20 | * | x | x | offline | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | Instrument off; carrier gas, CO2, and N2O cylinders left open.
2016-08-10 | 223 | 1400 | EB | 360 | 700/60 | 610/20 | 820/28 | 1140/34 | 800/20 | * | x | x | offline | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | Instrument off; carrier gas, CO2, and N2O cylinders left open.
2016-08-11 | 224 | 1243 | EB | 350 | 670/60 | 600/22 | 820/28 | 1110/34 | 800/20 | * | x | x | offline | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | Instrument off; carrier gas, CO2, and N2O cylinders left open.
2016-08-12 | 225 | 1534 | TG | 350 | 650/60 | 600/21 | 810/27 | 1100/33 | 800/20 | * | x | x | offline | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | Instrument off; carrier gas, CO2, and N2O cylinders left open. Power outage 1609-1611z. 1642-1644z
2016-08-13 | 226 | 1755 | EB | [6 empty] | * | x | x | offline | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | Instrument off; carrier gas, CO2, and N2O cylinders left open.
2016-08-14 | 227 | 1759 | EB | [6 empty] | * | x | x | offline | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | Instrument off; carrier gas, CO2, and N2O cylinders left open.
2016-08-15 | 228 | 1907 | TG/EB | - | - | - | - | - | - | * | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | Geoff here and currently working on instrument
2016-08-16 | 229 | 1810 | TG | 310 | 400/60 | 610/20 | 810/20 | 1190/22 | 740/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air1: 5.0, Air2: 4.8, BP1: 6.2, BP2: 8.6 Geoff Adjusted and changed all the Main and BF flows as well as some of the ECD and Column Temps. CO2 was low, so I informed Geoff who said it shouldn't drop. He is currently looking into it and checking for leaks
2016-08-17 | 230 | 1652 | TG | 290 | 410/59 | 620/19 | 810/17 | 1200/22 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-08-18 | 231 | 1731 | TG | 270 | 350/59 | 620/19 | 820/17 | 1210/22 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Cal1 to 50cc/min. Channel 1 lines were offset, so I contacted Geoff. He'll take a look when he gets back to Boulder
2016-08-19 | 232 | 1612 | EB/TG | 250 | 200/59 | 620/19 | 820/19 | 1200/20 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Channel 1 graph offset between red and blue. Changed N2 Ch4 at 1702z, changed N2 Ch1,2,3 at 1730z.
2016-08-20 | 233 | 1222 | TG | 2390 | 2270/60 | 620/19 | 810/19 | 1210/20 | 800/20 | * | * | * | * | * | * | offline | * | * | * | ? | * | * | * | * | * | * | * | * | * | * | Channel 1 graph offset between red and blue. Channel 4 Main has been hovering around 10 (+/-2) instead of 37, emailed Geoff
2016-08-21 | 234 | 1545 | EB | 2380 | 2200/60 | 600/20 | 810/17 | 1200/22 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | ? | * | * | * | * | * | * | * | * | Ch4 graph baseline off. Geoff working remotely, recheck tomorrow.
2016-08-22 | 235 | 1634 | TG | 2350 | 2100/60 | 620/19 | 810/19 | 1200/20 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | ? | * | * | * | * | * | * | * | * | Air1: 5.2, Air2: 5.5, BP1: 6.5, BP2: 9.0 Channel 1 graph still offset. Waiting to hear back from Geoff still. Channel 4 Main flow is at 17 (+/-3) instead of 37
2016-08-23 | 236 | 1255 | EB | 2310 | 2040/61 | 610/20 | 800/20 | 1200/21 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | ? | * | * | * | * | * | * | * | ? | Main Ch. 4 low (~17), Ch. 4 graph baseline off. 2030z Ch4 N2 valve pressure increased to 90psi.
2016-08-24 | 237 | 1330 | TG | 2300 | 1980/60 | 620/19 | 810/19 | 1200/21 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | ? | ? | ? | ? | Channel 1 graph still offset, all 4 graphs have a strange and strong negative spike downwards about 1/2 way along the x-axis. Contacted Geoff
2016-08-25 | 238 | 1223 | EB | 2300 | 1900/60 | 600/20 | 810/19 | 1200/20 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-08-26 | 239 | 1352 | TG | 2270 | 1830/59 | 620/19 | 810/19 | 1200/21 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in channel 1 graph
2016-08-27 | 240 | 1744 | EB | 2230 | 1810/60 | 610/20 | 810/20 | 1200/21 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in channel 1 graph
2016-08-28 | 241 | 1715 | TG | 2210 | 1700/60 | 610/20 | 810/20 | 1190/19 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-08-29 | 242 | 1354 | EB | 2200 | 1640/60 | 600/20 | 810/20 | 1200/21 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Slight offsets in ch. 1 and 3 graphs. Air1: 5.0, Air2: 4.5, BP1: 6.3, BP2: 8.5
2016-08-30 | 243 | 1300 | TG | 2180 | 1590/59 | 610/20 | 810/19 | 1190/19 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in channel 1 graph
2016-08-31 | 244 | 1306 | EB | 2140 | 1540/60 | 600/20 | 805/20 | 1195/20 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in channel 1 graph
2016-09-01 | 245 | 1316 | TG | 2140 | 1420/60 | 610/19 | 810/19 | 1195/20 | 810/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in channel 1 graph
2016-09-02 | 246 | 1700 | EB | 2100 | 1400/60 | 600/19 | 800/19 | 1190/20 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | switched out ECD 2 and 3 lines 1358z-1409z
2016-09-03 | 247 | 1323 | TG | 2100 | 1300/60 | 610/19 | 800/18 | 1190/19 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in channel 1 graph
2016-09-04 | 248 | 1419 | EB | 2080 | 1240/60 | 610/20 | 800/16 | 1190/19 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in channel 1 graph
2016-09-05 | 249 | 1250 | TG | 2050 | 1190/59 | 610/19 | 800/18 | 1195/19 | 810/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in channel 1 graph. Air1: 5.2, Air2: 4.0, BP1: 6.2, BP2: 8.5
2016-09-06 | 250 | 1342 | EB | 2020 | 1110/60 | 610/19 | 800/19 | 1190/16 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-09-07 | 251 | 1243 | TG | 2020 | 1090/59 | 610/19 | 795/18 | 1190/20 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-09-08 | 252 | 1237 | EB | 2010 | 1000/60 | 600/19 | 800/18 | 1180/20 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-09-09 | 253 | 1249 | TG | 1980 | 900/59 | 610/19 | 800/18 | 1180/20 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-09-10 | 254 | 1233 | EB | 1950 | 900/60 | 600/19 | 800/18 | 1190/20 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-09-11 | 255 | 1742 | TG | 1920 | 790/60 | 620/19 | 800/18 | 1190/21 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-09-12 | 256 | 1237 | EB | 1900 | 720/60 | 620/19 | 800/16 | 1170/22 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Still slight offset in ch1 graph. Air1: 5.0, Air2: 5.2, BP1: 6.3, BP2: 8.5
2016-09-13 | 257 | 1237 | TG | 1890 | 680/59 | 620/19 | 800/17 | 1180/20 | 810/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-09-14 | 258 | 1246 | EB | 1880 | 630/60 | 620/19 | 790/16 | 1180/22 | 820/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-09-15 | 259 | 1220 | TG | 1850 | 550/60 | 620/19 | 790/19 | 1170/21 | 810/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-09-16 | 260 | 1340 | EB | 1820 | 570/60 | 630/19 | 800/20 | 1190/21 | 810/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-09-17 | 261 | 1247 | TG | 1810 | 400/60 | 620/19 | 790/15 | 1170/21 | 810/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-09-18 | 262 | 1457 | EB | 1800 | 400/60 | 620/19 | 790/18 | 1170/20 | 820/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-09-19 | 263 | 1323 | TG | 1770 | 290/60 | 620/19 | 790/18 | 1170/20 | 810/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph Air 1: 5.0, Air2: 5.0, BP1: 6.2, BP2: 8.8
2016-09-20 | 264 | 1324 | EB | 1740 | 200/60 | 620/19 | 790/20 | 1170/21 | 810/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-09-21 | 265 | 1226 | TG | 1720 | 100/60 | 620/19 | 790/18 | 1160/20 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Adjusted Cal2 to 50cc/min. Changed N2 Channels 1,2,3 at 1235-1244z new pressure 2500psi
2016-09-22 | 266 | 1318 | EB | 1710 | 2420/60 | 630/19 | 790/19 | 1160/19 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-09-23 | 267 | 1216 | TG | 1700 | 2390/61 | 620/19 | 790/19 | 1160/19 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph COL #4 is ranging from 74.2-75.5
2016-09-24 | 268 | 1409 | EB | 1690 | 2300/60 | 630/19 | 780/19 | 1160/19 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-09-25 | 269 | 1337 | TG | 1650 | 2220/60 | 620/19 | 770/18 | 1160/19 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-09-26 | 270 | 1257 | EB | 1630 | 2200/60 | 620/19 | 780/20 | 1160/20 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph. Air1: 5.0, Air2: 5.0, BP1: 6.5, BP2: 8.5
2016-09-27 | 271 | 1207 | TG | 1600 | 2100/60 | 620/19 | 770/19 | 1160/20 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-09-28 | 272 | 1325 | EB | 1590 | 2010/61 | 640/19 | 770/19 | 1170/19 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-09-29 | 273 | 1247 | TG | 1560 | 1960/60 | 630/19 | 770/19 | 1160/20 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-09-30 | 274 | 1402 | EB | 1560 | 1880/60 | 630/19 | 760/19 | 1150/19 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-10-01 | 275 | 1323 | TG | 1500 | 1790/60 | 630/19 | 770/17 | 1160/17 | 810/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-10-02 | 276 | 1403 | EB | 1510 | 1730/60 | 620/19 | 770/19 | 1150/19 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-10-03 | 277 | 1510 | TG | 1490 | 1640/60 | 630/19 | 770/19 | 1160/19 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph Air1: 5.2, Air2: 5.0, BP1: 6.3, BP2: 8.7
2016-10-04 | 278 | 1219 | EB | 1490 | 1610/60 | 640/19 | 760/20 | 1160/20 | 810/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-10-05 | 279 | 1408 | TG | 1440 | 1510/60 | 620/19 | 760/20 | 1150/20 | 790/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-10-06 | 280 | 1314 | EB | 1430 | 1600/60 | 630/19 | 760/20 | 1160/20 | 810/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-10-07 | 281 | 1303 | TG | 1400 | 1400/60 | 630/19 | 760/19 | 1160/20 | 810/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-10-08 | 282 | 0 | 0 | [27 empty] | High winds/low visibility, didn't make it out to TAWO
2016-10-09 | 283 | 1459 | TG | 1350 | 1280/60 | 620/19 | 760/19 | 1160/20 | 820/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-10-10 | 284 | 1644 | EB | 1350 | 1200/60 | 620/19 | 760/19 | 1160/20 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air1 flow: 5.0, Air1 BP: 6.5, Air2 flow: 4.0, Air2 BP: 8.5; Cal2: 51.7cc/min, Cal1:53.4cc/min, Air1: 50.2cc/min
2016-10-11 | 285 | 1338 | TG | 1300 | 1120/62 | 620/19 | 760/18 | 1150/18 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-10-12 | 286 | 0 | 0
2016-10-13 | 287 | 1303 | TG | 1290 | 1000/61 | 620/19 | 750/19 | 1140/20 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph computer screen was black upon checks. Apparently a software crash. Reboot of the computer fixed it
2016-10-14 | 288 | 1716 | EB | 1270 | 980/60 | 620/19 | 760/20 | 1150/20 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-10-15 | 289 | 1300 | TG | 1230 | 880/60 | 620/19 | 760/19 | 1150/19 | 820/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph
2016-10-16 | 290 | 0 | 0
2016-10-17 | 291 | 1431 | TG | 1200 | 750 | 620/19 | 750/19 | 1150/20 | 810/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph. Adjust Cal1 to 50cc/min (1652z) Air1 flow: 5.2, Ai1 BP: 6.3, Air2 flow: 4.2, Air2 BP: 8.7
2016-10-18 | 292 | 1259 | EB | 1200 | 700/60 | 630/20 | 740/19 | 1140/19 | 810/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-10-19 | 293 | 1257 | TG | 1190 | 610/60 | 640/19 | 740/17 | 1140/17 | 810/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | still slight offset in ch1 graph. Slight offset in channel 3.
2016-10-20 | 294 | 0 | 0 | [27 empty] | Con 1. Didn't go to TAWO.
2016-10-21 | 295 | 1619 | TG | 1120 | 500/61 | 620/19 | 750/19 | 1140/21 | 810/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air1 flow: < 1, Air1 BP: 1.5 slight offset in channel 3, channel 1 looks good
2016-10-22 | 296 | 1213 | EB | 1110 | 400/60 | 610/19 | 740/19 | 1140/19 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * |  |  |  |  | Inlet pumps turned off from ~1215-1400z for diaphragm change, Air 1 pump left off (fan broken).
2016-10-23 | 297 | 1331 | TG | 1090 | 340/60 | 620/19 | 750/19 | 1140/21 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * |  |  |  |  | Air 1 pump left off (fan broken)
2016-10-24 | 298 | 1544 | TG/MG/NB | 1025 | 250/60 | 620/19 | 750/21 | 1155/20 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * |  |  |  |  | Air 1 pump left off (fan broken)
2016-10-25 | 299 | 1133 | MG/TG | 1025 | 250/60 | 620/18 | 750/19 | 1125/15 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * |  |  |  |  | Air 1 pump left off (fan broken), Cal2 50cc/min 1208z Air1 flow: 0, Ai1 BP: 1.5, Air2 flow: 4.2, Air2 BP: 8.5, Changed N2 1,2,3, cylindar: 1334-1345
2016-10-26 | 300 | 1213 | TG | 1010 | 2430/60 | 620/19 | 740/18 | 1125/20 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * |  |  |  |  | Air 1 pump left off (fan broken)
2016-10-27 | 301 | 1128 | MG/NB | 1000 | 2360/60 | 610/19 | 740/18 | 1130/19 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | The GC is making chomatagrams, previous techs left out the "*" above because of unstable chomatagrams. (Red and Blue lines do not match up, due to pump issue)
2016-10-28 | 302 | 1135 | MG/NB | 990 | 2200/60 | 620/19 | 750/18 | 1120/20 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Removed frost from sample inlets at 1133z. Air 1 pump offline.
2016-10-29 | 303 | 1150 | MG/NB | 990 | 2200/60 | 620/19 | 750/19 | 1120/20 | 800/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Cal 2 flow 52. Air 1 pump offline.
2016-10-30 | 304 | 1133 | NB | 950 | 2120/60 | 610/19 | 730/18 | 1120/21 | 790/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 pump offline.
2016-10-31 | 305 | 1318 | NB | 920 | 2060/60 | 610/19 | 730/17 | 1120/20 | 790/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 pump offline. Air 2 flow: 5.5 lpm, Air 2 BP: 8.7; Cal 1 flow: 51cc/min, Cal 2 flow: 52cc/min. Adjusted air 2 to 50cc/min at 1930z.
2016-11-01 | 306 | 1253 | MG | 900 | 2000/60 | 610/19 | 730/20 | 1120/22 | 790/20 | * | * | * | * | * | * | offine | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 2 reads 50.1 at 1252z, Cal 1 is 49.8 at 1320z
2016-11-02 | 307 | 1251 | NB | 890 | 1960/60 | 610/19 | 730/20 | 1120/21 | 780/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 pump still offline.
2016-11-03 | 308 | 1245 | MG | 870 | 1890/60 | 610/19 | 730/19 | 1120/21 | 780/20 | * | * | * | * | * | * | offline | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Installed Air 1 sample line on spare port on Air 2 pump. Air 1 flow is high (~75cc/min) and can not be adjusted independently of Air 2. NB
2016-11-04 | 309 | 1309 | MG | 850 | 1790/60 | 610/19 | 730/18 | 1120/20 | 780/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 pump still offline.
2016-11-05 | 310 | 1230 | MG | 830 | 1720/60 | 610/19 | 730/20 | 1120/20 | 780/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * |  | * | * | * | * | * | * | Air 2 at 50.6 lmp, Channel 2 Bf reading 0.0
2016-11-06 | 311 | 1325 | NB | 810 | 1670/60 | 610/19 | 730/17 | 1120/19 | 780/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * |  | * | * | * | * | * | * | Channel 2 BF still reading 0.0. Air 1 flow 74.0cc/min
2016-11-07 | 312 | 1230 | MG | 790 | 1620/60 | 610/20 | 730/19 | 1120/20 | 780/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * |  | * | * | * | * | * | * | Channel 2 BF still reading 0.0.
2016-11-08 | 313 | 1225 | MG | 770 | 1540/60 | 610/19 | 720/20 | 1110/20 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Cal 2 flow 50.4cc/min Air 2 flow 48-52 cc/min
2016-11-09 | 314 | 1216 | MG | 750 | 1500/60 | 610/19 | 720/20 | 1110/21 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | CAL 2 flow 50.1cc/min
2016-11-10 | 315 | 1242 | MG | 720 | 1420/61 | 610/19 | 720/20 | 1100/20 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-11-11 | 316 | [29 empty] | Did not go to TAWO condition 1
2016-11-12 | 317 | 1240 | MG/NB | 660 | 1250/61 | 590/20 | 690/20 | 1050/21 | 590/21 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Cold in TAWO today do to strong east wind (42F). Tank pressures low as a result.
2016-11-13 | 318 | 1250 | MG | 680 | 1230/61 | 610/19 | 710/20 | 1100/20 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-11-14 | 319 | 1900 | NB | 650 | 1150/60 | 610/19 | 710/19 | 1100/16 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-11-15 | 320 | 1400 | MG | 630 | 1100/60 | 610/19 | 710/19 | 1100/15 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-11-16 | 321 | 1302 | MG | 610 | 1020/60 | 610/19 | 710/19 | 1100/15 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-11-17 | 322 | 1218 | MG/NB | 600 | 1000/60 | 610/19 | 690/18 | 1110/21 | 730/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Turned up Cal 2 to ~51cc/min, Air 1 flow ~72.8 cc/min
2016-11-18 | 323 | 1323 | MG | 590 | 900/60 | 610/19 | 690/19 | 1100/22 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Cal 1 flow ~51.8 cc/min
2016-11-19 | 324 | 1235 | NB | 580 | 840/60 | 610/19 | 700/19 | 1110/16 | 780/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Cal 2 flow ~50cc/min
2016-11-20 | 325 | 1312 | NB | 550 | 780/60 | 610/19 | 710/20 | 1100/21 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-11-21 | 326 | 1305 | NB | 520 | 710/60 | 610/19 | 710/19 | 1100/21 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 2 flow: 4.2lpm, Air 2 BP: 9.0; Air 1 pump offline; Cal 1 flow: 51cc/min, Cal 2 flow 50cc/min. Adjusted Air 2 flow to 50cc/min.
2016-11-22 | 327 | 1301 | MG/NB | 500 | 650/60 | 610/19 | 710/19 | 1100/21 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-11-23 | 328 | 1230 | NB | 480 | 500/60 | 610/19 | 700/20 | 1080/25 | 740/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-11-24 | 329 | 1250 | NB | 460 | 400/60 | 610/19 | 700/20 | 1090/20 | 740/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Cal 2 flow 51cc/min
2016-11-25 | 330 | 1249 | MG | 450 | 450/60 | 610/19 | 700/20 | 1100/20 | 740/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-11-26 | 331 | 1335 | NB | 420 | 380/60 | 610/19 | 700/20 | 1100/21 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 2: 52cc/min
2016-11-27 | 332 | 1420 | MG | 400 | 210/60 | 610/19 | 700/20 | 1100/20 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1: 80cc/min
2016-11-28 | 333 | 1235 | NB | 390 | 250/60 | 610/19 | 700/20 | 1100/20 | 780/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 pump offline; Air 2 line flow: 3.9lpm, Air 2 line BP: 9.0psi; Cal 1 flow: 51cc/min; Cal 2 flow 51cc/min
2016-11-29 | 334 | 1245 | NB | 360 | 200/60 | 610/19 | 700/20 | 1090/20 | 790/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Swapped N2 for channel 1,2, 3 at 1351 UTC
2016-11-30 | 335 | 1325 | MG | 350 | 2380/60 | 610/19 | 700/20 | 1090/20 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-12-01 | 336 | 1250 | NB | 320 | 2340/60 | 610/19 | 690/20 | 1090/20 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-12-02 | 337 | 1321 | MG | 310 | 2250/60 | 610/19 | 690/20 | 1090/20 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-12-03 | 338 | 1235 | NB | 290 | 2200/60 | 610/19 | 690/19 | 1090/20 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Cal 1 flow = 52cc/min
2016-12-04 | 339 | 1348 | NB | 280 | 2130/60 | 610/19 | 690/20 | 1090/21 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-12-05 | 340 | 1242 | NB | 270 | 2060/60 | 610/19 | 690/19 | 1090/20 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 flow 80cc/min, Air 2 flow 50cc/min, Cal 1 flow 52cc/min, Cal 2 flow 51cc/min; Air 1 pump offline, Air 2 pump flow 3.9lpm, Air 2 pump BP 9.0.
2016-12-06 | 341 | 1244 | MG | 250 | 2000/60 | 610/19 | 690/19 | 1090/20 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-12-07 | 342 | 1240 | MG | 220 | 2020/60 | 610/19 | 690/19 | 1090/20 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 flow ~81cc/min
2016-12-08 | 343 | 1234 | MG/NB | 200 | 1880/60 | 610/19 | 690/18 | 1090/20 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-12-09 | 344 | 1328 | MG | 180 | 1770/60 | 610/19 | 690/20 | 1090/20 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-12-10 | 345 | 1920 | MG | 160 | 1700/60 | 610/19 | 690/20 | 1090/20 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-12-11 | 346 | 1340 | MG | 140 | 1660/60 | 610/19 | 690/20 | 1090/20 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-12-12 | 347 | 1230 | NB | 130 | 1600/60 | 610/19 | 690/18 | 1090/19 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Air 1 pump offline; Air 2 line flow: 3.9lpm, Air 2 line BP: 8.9psi; Cal 1 flow: 51cc/min; Cal 2 flow 51cc/min
2016-12-13 | 348 | 1225 | NB | 110 | 1520/60 | 610/19 | 690/17 | 1080/21 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Changed Chan 4 N2 tank at 1242 - 1248z.
2016-12-14 | 349 | 1304 | MG | 2540 | 1450/60 | 610/19 | 690/19 | 1070/20 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-12-15 | 350 | 1415 | MG | 2520 | 1400/60 | 610/19 | 690/20 | 1090/20 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-12-16 | 351 | 1308 | MG | 2500 | 1310/60 | 610/20 | 690/20 | 1080/20 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-12-17 | 352 | 1252 | MG | 2480 | 1280/60 | 610/19 | 690/18 | 1070/19 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-12-18 | 353 | 1328 | NB | 2450 | 1200/60 | 610/19 | 680/19 | 1070/20 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
2016-12-19 | 354 | 1220 | NB | 2400 | 1120/60 | 610/19 | 680/18 | 1070/19 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | ~ | * | Air 1 pump offline. Air 2 pump flow: 3.6lpm. Air 2 pump BP: 8.8. Air 1 flow: 78cc/min. Air 2 flow: 52cc/min. Cal 1 flow: 54cc/min. Cal 2 flow: 51cc/min. Channel 3 graph appears to have an unusually high baseline and peaks that differ significantly from their usual arrangement.
2016-12-20 | 355 | 1506 | MG | 2390 | 1040/60 | 610/19 | 670/20 | 1060/20 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | ~ | * | Channel 3 base line even higher than yesterday
2016-12-21 | 356 | 1325 | MG | 2360 | 1000/60 | 610/19 | 680/20 | 1040/19 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | ~ | * | Channel 3 base line even higher than yesterday
2016-12-22 | 357 | 1350 | MG | 2340 | 920/60 | 610/19 | 670/20 | 1040/19 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * |  | * | CAL 2 flow ~51 cc/min. Changed ECD 3 temp setpoint from 280C to 250C at 1421z per request from Geoff Dutton.
2016-12-23 | 358 | 1244 | NB | 2310 | 860/60 | 610/19 | 670/20 | 1060/21 | 740/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | ~ | * | Air 2 flow: 52cc/min. Channel 3 baseline is still high.
2016-12-24 | 359 | 1326 | NB | 2300 | 800/60 | 610/19 | 670/19 | 1060/21 | 740/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | ~ | * | Channel 3 baseline is still high, but starting to come down a bit.
2016-12-25 | 360 | 1456 | MG | 2280 | 710/60 | 610/19 | 660/20 | 1050/21 | 740/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | ~ | * | Channel 3 looks almost back to normal
2016-12-26 | 361 | 1235 | NB | 2260 | 680/60 | 610/19 | 660/20 | 1060/20 | 740/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | ~ | * | Channel 3 baseline is still high, but getting better. Air 1 pump offline. Air 2 pump flow: 3.6 lpm. Air 2 pump BP: 8.9. Air 1 flow: 78cc/min. Air 2 flow: 53cc/min. Cal 1 flow: 54cc/min. Cal 2 flow: 51cc/min.
2016-12-27 | 362 | 1034 | MG | 2150 | 600/60 | 560/21 | 640/20 | 1010/20 | 590/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | ~ | * | Cylindars reading low because TAWO is cold, 38 degrees
2016-12-28 | 363 | 1358 | MG | 2210 | 530/60 | 610/20 | 660/20 | 1060/25 | 750/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * |  | * | channel 3 base line worse, almost competely maxed out
2016-12-29 | 364 | 1308 | MG/NB | 2200 | 490/60 | 610/19 | 660/19 | 1060/22 | 740/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * |  | * | Channel 3 baseline about halfway up plot
2016-12-30 | 365 | 1314 | MG/NB | 2190 | 400/60 | 610/19 | 660/20 | 1050/22 | 740/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * |  | * | Changes ECD lines between channel 2 and channel 3 at 1524 UTC
2016-12-31 | 366 | 1332 | NB | 2150 | 350/60 | 610/19 | 660/20 | 1060/21 | 740/20 | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | * | Chromatograms seem to be looking better at ECD line swap.

## NOAA Solar
Date | Julian Day | Time (UTC) | Observer's Initials | SPN1 | CM11 Pyranometer | Kipp & Zonen ventilator inlet | Did I launch the NOAA missile if I cleaned a radiometer? | comments
2016-08-17 | 230 | 1652 | TG | clr | rm frost | clr
2016-08-18 | 231 | 1731 | TG | rm ltfrost | r ltfrost | clr
2016-08-19 | 232 | 1612 | EB/TG | clr | clr | clr
2016-08-20 | 233 | 1222 | TG | clr | clr | clr
2016-08-21 | 234 | 1545 | EB | lt frost | lt frost | clr
2016-08-22 | 235 | 1634 | TG | clr | lt frost | clr
2016-08-23 | 236 | 1255 | EB | rm ltfrost | rm frost | clr
2016-08-24 | 237 | 1330 | TG | rm ltfrost | rm frost | clr |  | New batter installed for CM11 ventilator heater
2016-08-25 | 238 | 1223 | EB | clr | clr | clr
2016-08-26 | 239 | 1352 | TG | clr | rm lt frost | clr
2016-08-27 | 240 | 1744 | EB | clr | clr | rm frost
2016-08-28 | 241 | 1715 | TG | clr | clr | clr
2016-08-29 | 242 | 1354 | EB | clr | clr | clr
2016-08-30 | 243 | 1300 | TG | clr | rm frost | rm frost |  | The pyranometer dome was covered in frost as was the housing. It appears as though initially the housing had melted the frost but then it froze as the frost on the housing wasn't light and fluffy but was icy, like the frost you have to scrape off a car windshield
2016-08-31 | 244 | 1306 | EB | clr | clr | clr
2016-09-01 | 245 | 1316 | TG | clr | rm frost | rm frost |  | The pyranometer dome was covered in frost as was the housing. It appears as though initially the housing had melted the frost but then it froze as the frost on the housing wasn't light and fluffy but was icy, like the frost you have to scrape off a car windshield
2016-09-02 | 246 | 1700 | EB |  |  |  |  | Switched ventilator heater from PS12 battery to 12VDC power supply
2016-09-03 | 247 | 1323 | TG | clr | rm hv frost | hv frost
2016-09-04 | 248 | 1419 | EB | clr | rm lt frost | clr
2016-09-05 | 249 | 1250 | TG | clr | rm frost | rm lt frost
2016-09-06 | 250 | 1342 | EB | clr | clr | clr
2016-09-07 | 251 | 1243 | TG | clr | dome clear, rm frost from housing | rm frost | y | dome on CM11 clear, but entire housing frosted
2016-09-08 | 252 | 1237 | EB | clr | rm frost | clr | y
2016-09-09 | 253 | 1249 | TG | clr | rm frost | clr | y* | inlet clear, rm frost from tube and mesh wire *pressed button ~30mins after cleaning
2016-09-10 | 254 | 1233 | EB | clr | clr | clr
2016-09-11 | 255 | 1742 | TG | clr | rm frost | clr | y | rm frost from mesh wire
2016-09-12 | 256 | 1237 | EB | clr | clr | clr
2016-09-13 | 257 | 1237 | TG | rm ltfrost | rm frost | clr | y | rm frost from mesh wire
2016-09-14 | 258 | 1246 | EB | clr | clr | rm lt frost
2016-09-15 | 259 | 1220 | TG | clr | dome clear, rm frost from housing | clr | y | rm frost from mesh wire
2016-09-16 | 260 | 1340 | EB | clr | clr | rm lt frost
2016-09-17 | 261 | 1247 | TG | clr | rm frost | clr | y | rm frost from mesh wire
2016-09-18 | 262 | 1457 | EB | clr | rm frost | clr | y | about a half hour late
2016-09-19 | 263 | 1323 | TG | clr | rm frost | clr | Y | rm frost from mesh wire
2016-09-20 | 264 | 1324 | EB | clr | clr | rm frost
2016-09-21 | 265 | 1226 | TG | rm ltfrost | rm frost | clr | Y | rm heavy frost from mesh wire. Frost on SPN1 was fluffy and came off easily with a simple brush stroke. Frost on CM11 is icy underneath and always requires OH
2016-09-22 | 266 | 1318 | EB | clr | rm frost | rm frost | y
2016-09-23 | 267 | 1216 | TG | clr | clr | clr
2016-09-24 | 268 | 1409 | EB | clr | clr | clr
2016-09-25 | 269 | 1337 | TG | clr | clr | clr
2016-09-26 | 270 | 1257 | EB | rm ltfrost | rm frost | rm frost | y
2016-09-27 | 271 | 1207 | TG
2016-09-28 | 272 | 1325 | EB | clr | rm frost | rm frost | y
2016-09-29 | 273 | 1247 | TG | clr | rm lt frost | clr | y | rm frost from mesh wire
2016-09-30 | 274 | 1402 | EB | clr | clr | clr
2016-10-01 | 275 | 1323 | TG | clr | clr | clr
2016-10-02 | 276 | 1403 | EB | clr | clr | clr
2016-10-03 | 277 | 1510 | TG | clr | rm lt frost | clr | y
2016-10-04 | 278 | 1219 | EB | clr | rm lt frost | clr | y
2016-10-05 | 279 | 1408 | TG | rm ltfrost | rm hv frost | rm frost | y | rm heavy frost from mesh wire
2016-10-06 | 280 | 1314 | EB | clr | clr | clr
2016-10-07 | 281 | 1303 | TG | clr | rm frost | rm blown snow | y
2016-10-08 | 282 | 0 | 0 |  |  |  |  | High winds/low visibility, didn't make it out to TAWO
2016-10-09 | 283 | 1459 | TG | clr | rm lt frost | clr | y
2016-10-10 | 284 | 1644 | EB | clr | clr | clr
2016-10-11 | 285 | 1338 | TG | rm frost | rm frost | clr | y | rm heavy frost from mesh wire
2016-10-12 | 286 | 0 | 0
2016-10-13 | 287 | 1303 | TG | clr | clr | clr
2016-10-14 | 288 | 1716 | EB | clr | clr | clr
2016-10-15 | 289 | 1300 | TG | clr | rm lt frost | rm blown snow | y
2016-10-16 | 290 | 0 | 0
2016-10-17 | 291 | 1431 | TG | rm ltfrost | rm frost | clr | y | rm heavy frost from mesh wire
2016-10-18 | 292 | 1259 | EB | clr | rm lt frost | rm lt frost | y
2016-10-19 | 293 | 1257 | TG | clr | rm lt frost | clr | y
2016-10-20 | 294 | 0 | 0 |  |  |  |  | Con 1. Didn't go to TAWO.
2016-10-21 | 295 | 1619 | TG | clr | rm lt frost | clr | y
2016-10-22 | 296 | 1213 | EB | clr | clr | clr
2016-10-23 | 297 | 1331 | TG | clr | clr | clr
2016-10-24 | 298 | 1544 | TG/MG/NB | clr | rm frost | clr | y
2016-10-25 | 299 | 1133 | MG/TG | rm ltfrost | rm frost | clr | y
2016-10-26 | 300 | 1213 | TG | clr | rm frost | clr | y
2016-10-27 | 301 | 1128 | MG/NB | clr | rm frost | clr | y
2016-10-28 | 302 | 1135 | MG/NB | rm ltfrost | rm frost | clr | y
2016-10-29 | 303 | 1150 | MG/NB | clr | rm lt frost | clr | Y
2016-10-30 | 304 | 1133 | NB | clr | clr | clr
2016-10-31 | 305 | 1318 | NB | rm frost | rm frost | clr | Y
2016-11-01 | 306 | 1253 | MG | clr | rm lt frost | clr | Y
2016-11-02 | 307 | 1251 | NB | clr | rm lt frost | clr | Y
2016-11-03 | 308 | 1245 | MG | clr | rm lt frost | rm lt frost | Y | The inside diameter at the top of the ventilaor inlet was starting to skrink, knocks the inlet tube a few times and the frost fell out.
2016-11-04 | 309 | 1309 | MG | rm lt frost | rm frost | clr | Y
2016-11-05 | 310 | 1230 | MG | clr | clr | clr | Y
2016-11-06 | 311 | 1325 | NB | clr | clr | clr
2016-11-07 | 312 | 1230 | MG | clr | rm lt frost | clr | Y
2016-11-08 | 313 | 1225 | MG | clr | rm lt frost | clr | Y
2016-11-09 | 314 | 1216 | MG | clr | rm lt frost | clr | Y
2016-11-10 | 315 | 1242 | MG | clr | rm lt frost | clr | Y
2016-11-11 | 316 | [6 empty] | Con 1. Didn't go to TAWO.
2016-11-12 | 317 | 1240 | MG/NB |  |  |  |  | did not go on the roof due to 35 knot winds
2016-11-13 | 318 | 1250 | MG | clr | rm frost | clr | Y
2016-11-14 | 319 | 1900 | NB | rm frost | rm frost |  | Y
2016-11-15 | 320 | 1400 | MG | clr | rm frost | clr | Y
2016-11-16 | 321 | 1302 | MG | clr | rm frost | clr | Y
2016-11-17 | 322 | 1218 | MG/NB | clr | rm frost | clr | Y
2016-11-18 | 323 | 1323 | MG | clr | rm frost | clr | Y
2016-11-19 | 324 | 1235 | NB | clr | rm lt frost | clr | Y
2016-11-20 | 325 | 1312 | NB | rm lt frost | rm lt frost | clr | Y
2016-11-21 | 326 | 1305 | NB | rm lt frost | rm lt frost | rm lt frost | Y | The inside diameter at the top of the ventilaor inlet was starting to shrink, knocks the inlet tube a few times and the frost fell out.
2016-11-22 | 327 | 1301 | MG/NB | rm lt frost | rm lt frost | rm lt frost | Y
2016-11-23 | 328 | 1230 | NB | clr | rm lt frost | clr | Y
2016-11-24 | 329 | 1250 | NB | clr | rm lt frost | clr | Y
2016-11-25 | 330 | 1249 | MG | clr | rm lt frost | clr | Y
2016-11-26 | 331 | 1335 | NB | clr | clr | clr
2016-11-27 | 332 | 1420 | MG | clr | rm lt frost | clr | Y
2016-11-28 | 333 | 1235 | NB | rm lt frost | rm lt frost | clr | Y
2016-11-29 | 334 | 1245 | NB | rm lt frost | rm lt frost | rm lt frost | Y | The inside diameter at the top of the ventilaor inlet was starting to skrink, knocks the inlet tube a few times and the frost fell out.
2016-11-30 | 335 | 1325 | MG | rm lt frost | rm lt frost | clr | Y
2016-12-01 | 336 | 1250 | NB | rm lt frost | rm lt frost | rm lt frost | Y
2016-12-02 | 337 | 1321 | MG | rm forst | rm frost | clr | Y
2016-12-03 | 338 | 1235 | NB | clr | rm frost | clr | Y
2016-12-04 | 339 | 1348 | NB | rm lt frost | rm lt frost | rm lt frost | Y
2016-12-05 | 340 | 1242 | NB | clr | rm lt frost | clr | N | Forgot to launch missle. Cleaned CMP @ 1247z.
2016-12-06 | 341 | 1244 | MG | rm frost | rm lt frost | rm lt frost | Y
2016-12-07 | 342 | 1240 | MG | clr | clr | clr | Y
2016-12-08 | 343 | 1234 | MG/NB | clr | clr | clr | Y
2016-12-09 | 344 | 1328 | MG | rm frost | rm frost | clr | Y | Both domes had a hard ice glaze on the windward side
2016-12-10 | 345 | 1920 | MG | clr | rm lt frost | clr | Y
2016-12-11 | 346 | 1340 | MG | rm lt frost | rm lt frost | clr | Y
2016-12-12 | 347 | 1230 | NB | rm lt frost | rm lt frost | clr | Y
2016-12-13 | 348 | 1225 | NB | clr | rm lt frost | clr | Y
2016-12-14 | 349 | 1304 | MG | clr | rm lt frost | clr | Y
2016-12-15 | 350 | 1415 | MG | clr | rm lt frost | clr | Y
2016-12-16 | 351 | 1308 | MG | clr | rm lt frost | clr | Y
2016-12-17 | 352 | 1252 | MG | clr | rm lt frost | clr | Y
2016-12-18 | 353 | 1328 | NB | clr | rm lt frost | clr | Y
2016-12-19 | 354 | 1220 | NB | clr | rm lt frost | clr | N | Forgot to launch missle. Cleaned CMP @ 1237z.
2016-12-20 | 355 | 1506 | MG | clr | rm lt frost | clr | Y
2016-12-21 | 356 | 1325 | MG | clr | rm lt frost | clr | Y
2016-12-22 | 357 | 1350 | MG
2016-12-23 | 358 | 1244 | NB | clr | rm lt frost | clr | N | Forgot to launch missle. Cleaned CMP @ 1242z.
2016-12-24 | 359 | 1326 | NB | clr | rm lt frost | clr | Y
2016-12-25 | 360 | 1456 | MG | clr | rm lt frost | clr | Y
2016-12-26 | 361 | 1235 | NB | rm lt frost | rm lt frost | clr | Y
2016-12-27 | 362 | 1034 | MG | - | - | - | - | Did not go on the roof due to 32 knot winds
2016-12-28 | 363 | 1358 | MG | clr | rm lt frost | clr | Y
2016-12-29 | 364 | 1308 | MG/NB | clr | rm lt frost | clr | Y
2016-12-30 | 365 | 1314 | MG/NB | clr | rm lt frost | clr | Y
2016-12-31 | 366 | 1332 | NB | clr | rm lt frost | clr | Y

## 50-m Met Suite
Date | Julian Day | Time (UTC) | Observer's Initials | 2m RH | 2m RTD | 2m Sonic | 10m Windbird | 10m RTD | 10m Sonic | Comments (Climbing, removing instruments, etc) | clr= elements clear, no cleaning necessary lf= lite frost hf= heavy frost rf= removed frost (NO RIME! All frost)
2016-07-05 | 187 | 12:00 | HJ | clr | clr | clr | clr | clr | clr
2016-07-06 | 188 | 11:50 | HJ | clr | clr | clr | clr | clr | clr
2016-07-07 | 189 | 11:50 | HJ | clr | clr | clr | clr | clr | clr
2016-07-08 | 190 |  | HJ | clr | clr
2016-07-09 | 191 |  | HJ
2016-07-10 | 192 | 14:01 | HJ | clr | clr | clr | clr | clr | clr
2016-07-11 | 193 | 14:00 | HJ | clr | clr | clr | clr | clr | clr
2016-07-12 | 194 | 12:00 | HJ | clr | clr | clr | clr | clr | clr
2016-07-13 | 195 | 12:20 | HJ | clr | clr | clr | clr | clr | clr
2016-07-14 | 196 | 11:36 | HJ | clr | clr | clr | clr | clr | clr
2016-07-15 | 197 | 11:45 | HJ | rf | rf | clr | lf | lf | clr | Very little frost on the windbird
2016-07-16 | 198 | 12:00 | HJ | clr | clr | clr | clr | clr | clr | HJ climbed tower and moved in 2m sonic for training 12-1215utc.
2016-07-17 | 199 |  | HJ
2016-07-18 | 200 | 18:20 | HJ | clr | clr | clr | clr | clr | clr
2016-07-19 | 201 | 16:00 | HJ | clr | clr | clr | clr | clr | clr
2016-07-20 | 202 | 13:10 | HJ | lf | lf | clr | lf | lf | clr
2016-07-21 | 203 | 13:40 | HJ | clr | clr | clr | clr | clr | clr
2016-07-22 | 204 | 15:15 | HJ | clr | clr | clr | clr | clr | clr | HJ climbed tower 15:15-15:40z for weekly inspection of 10m instruments
2016-07-23 | 205 | 12:03 | HJ | lf | lf | clr | clr | lf | clr
2016-07-24 | 206 | 13:17 | HJ | clr | clr | clr | clr | clr | clr
2016-07-25 | 207 | 12:05 | HJ | hf | hf | clr | clr | hf | clr
2016-07-26 | 208 |  | HJ
2016-07-27 | 209 | 12:30 | HJ | clr | clr | clr | clr | clr | clr
2016-07-28 | 210 | 11:50 | HJ | clr | clr | clr | clr | clr | clr
2016-07-29 | 211 | 12:40 | HJ | lf | lf | clr | clr | lf | clr
2016-07-30 | 212 | 12:00 | HJ | hf | hf | clr | lf | hf | clr
2016-07-31 | 213 | 12:45 | HJ | hf | hf | clr | clr | hf | clr
2016-08-01 | 214 | 12:00 | TG/HJ | hf | hf | clr | clr | hf | clr
2016-08-02 | 215 | 12:35 | TG/HJ | clr | clr | clr | clr | clr | clr
2016-08-03 | 216 | 12:15 | HJ | clr | clr | clr | clr | clr | clr
2016-08-04 | 217 | 11:48 | HJ | lf | hf | clr | clr | hf | clr
2016-08-05 | 218 | 12:00 | HJ | hf | hf | clr | clr | hf | clr
2016-08-06 | 219 | 11:55 | HJ | lf | lf | clr | clr | lf | clr
2016-08-07 | 220
2016-08-08 | 221 | 12:05 | HJ/EB | clr | clr | clr | clr | clr | clr
2016-08-09 | 222 | 12:20 | HJ/EB/TG | lf | lf | clr | clr | lf | lf
2016-08-10 | 223 | 12:10 | HJ | clr | lf | clr | clr | clr | clr
2016-08-11 | 224
2016-08-12 | 225
2016-08-13 | 226 | 11:20 | EB/TG | lf | lf | clr | clr | lf | clr
2016-08-14 | 227
2016-08-15 | 228 | 13:30 | HJ | lf | lf | clr | clr | hf | clr | Hannah and Grey on tower from 1330-1445z
2016-08-16 | 229 | 11:45 | TG | lf | lf | clr | clr | lf | clr
2016-08-17 | 230 | 17:00 | EB | lf | lf | clr | clr | clr | clr
2016-08-18 | 231 | 14:40 | EB | lf | lf | clr | clr | clr | clr
2016-08-19 | 232 | 11:51 | TG | lf | lf | clr | clr | lf | clr
2016-08-20 | 233 | 12:25 | EB | lf | lf | clr | clr | clr | clr
2016-08-21 | 234 | 18:01 | TG | hf | hf | clr | clr | hf | clr
2016-08-22 | 235 | 13:02 | EB/TG | lf | lf | clr | clr | lf | clr
2016-08-23 | 236 | 11:15 | TG | hf | hf | clr | clr | lf | clr
2016-08-24 | 237 | 12:45 | EB | lf | lf | clr | clr | lf | clr | Protective cap had fallen off 2m RTD, frost collected on sensor. Cap replaced 13:00z.
2016-08-25 | 238 | 12:00 | TG | clr | lf | clr | clr | lf | clr
2016-08-26 | 239 | 15:15 | EB | lf | lf | clr | clr | clr | clr
2016-08-27 | 240 | 13:46 | TG | hf | hf | clr | lf | hf | clr
2016-08-28 | 241 | 14:35 | EB | lf | lf | clr | clr | lf | clr
2016-08-29 | 242 | 12:37 | TG | lf | lf | clr | clr | lf | clr
2016-08-30 | 243 | 12:45 | EB | lf | lf | clr | clr | lf | clr
2016-08-31 | 244 | 12:01 | TG | lf | lf | clr | clr | lf | clr
2016-09-01 | 245 | 12:03 | EB | lf | lf | clr | clr | lf | clr
2016-09-02 | 246 | 11:54 | TG | lf | clr | clr | clr | clr | clr
2016-09-03 | 247
2016-09-04 | 248 | 12:24 | TG | lf | lf | clr | clr | fr | clr
2016-09-05 | 249 | 12:15 | EB | lf | lf | clr | clr | clr | clr
2016-09-06 | 250 | 11:54 | TG | lf | lf | clr | clr | fr | clr
2016-09-07 | 251 | 11:55 | EB | lf | lf | clr | clr | fr | clr
2016-09-08 | 252 | 11:51 | TG | lf | lf | clr | clr | fr | clr
2016-09-09 | 253 | 12:40 | EB | lf | lf | clr | clr | lf | clr
2016-09-10 | 254 | 11:51 | TG | lf | lf | clr | clr | lf | clr
2016-09-11 | 255 | 19:30 | EB | lf | lf | clr | clr | lf | clr
2016-09-12 | 256 | 13:21 | TG | fr | fr | clr | clr | fr | clr
2016-09-13 | 257 | 12:38 | EB | lf | lf | clr | lf | lf | clr
2016-09-14 | 258 | 11:52 | TG | fr | fr | clr | clr | fr | clr
2016-09-15 | 259 | 12:00 | EB | lf | lf | clr | lf | lf | clr
2016-09-16 | 260
2016-09-17 | 261 | 14:10 | EB | lf | lf | clr | lf | lf | clr
2016-09-18 | 262 | 13:02 | TG | fr | fr | clr | lf | fr | clr
2016-09-19 | 263 | 12:35 | EB | fr | fr | clr | lt fr | fr | clr
2016-09-20 | 264 | 12:03 | TG | fr | fr | clr | lt fr | fr | clr
2016-09-21 | 265 | 12:30 | EB | fr | fr | clr | lt fr | fr | clr
2016-09-22 | 266 | 12:05 | TG | fr | fr | clr | clr | fr | clr
2016-09-23 | 267 | 12:15 | EB | fr | fr | clr | clr | fr | clr
2016-09-24 | 268 | 14:36 | TG | lf | lf | clr | clr | lf | clr
2016-09-25 | 269 | 14:15 | EB | lf | lf | clr | clr | fr | clr
2016-09-26 | 270 | 13:23 | TG | fr | fr | clr | fr | fr | clr | Tguy on tower from ~1730-1810z
2016-09-27 | 271 | 19:00 | EB | lf | lf | clr | clr | fr | clr
2016-09-28 | 272 | 13:32 | TG | fr | fr | clr | lf | fr | clr
2016-09-29 | 273 | 16:00 | EB | fr | fr | clr | lf | fr | clr
2016-09-30 | 274 | 12:02 | TG | lf | lf | clr | clr | lf | clr
2016-10-01 | 275 | 13:00 | EB | clr | clr | clr | clr | lf | clr
2016-10-02 | 276 | 13:57 | TG | clr | clr | clr | clr | clr | clr
2016-10-03 | 277 | 13:10 | EB | clr | clr | clr | clr | lf | clr
2016-10-04 | 278 | 11:59 | TG | clr | clr | clr | clr | clr | clr
2016-10-05 | 279 | 17:00 | EB | clr | clr | clr | clr | lf | clr
2016-10-06 | 280 | 15:50 | TG | clr | lf | clr | clr | lf | clr
2016-10-07 | 281 | 12:15 | EB | clr | clr | clr | clr | lf | clr
2016-10-08 | 282 | 12:29 | TG | fr | fr | clr | lf | fr | clr
2016-10-09 | 283
2016-10-10 | 284 | 13:54 | TG | fr | fr | clr | fr | fr | clr
2016-10-11 | 285 | 16:43 | EB | lt fr | lt fr | clr | clr | lt fr | clr
2016-10-12 | 286 | 12:18 | TG | clr | clr | clr | clr | lt fr | clr
2016-10-13 | 287 | 14:10 | EB | clr | clr | clr | clr | clr | clr
2016-10-14 | 288 | 12:49 | TG | clr | clr | clr | clr | clr | clr
2016-10-15 | 289 | 13:50 | EB | clr | clr | clr | clr | clr | clr
2016-10-16 | 290 | [8 empty] | did not make it to tower due to all-day ICESAT traverse
2016-10-17 | 291 | 16:00 | EB | lf | lf | clr | lf | fr | clr
2016-10-18 | 292 | 12:09 | TG | lf | lf | clr | lf | fr | clr
2016-10-19 | 293 | 13:10 | EB | lf | lf | clr | lf | lf | clr | climbed the 50 m tower, RTDs still good, diagonal support loose at about 9m
2016-10-20 | 294 | 15:10 | TG | clr | clr | clr | clr | clr | clr | got report from Any and Christy that weather station was down. At 1513z I cycled power and this got the lights blinking. Will see if this fixed the issue. Heard back in the late afternoon that the station was up and running. Suspected static issue from high winds
2016-10-21 | 295 | 15:30 | EB | clr | clr | clr | clr | clr | clr
2016-10-22 | 296 | 14:19 | TG | clr | clr | clr | clr | clr | clr
2016-10-23 | 297 | 15:30 | EB | clr | clr | clr | clr | clr | clr
2016-10-24 | 298 | 15:30 | EB | lt fr | lt fr | clr | clr | lt fr | clr
2016-10-25 | 299 | 11:59 | SWD | lf | lf | clr | lf | lf | clr
2016-10-26 | 300 | 12:04 | SWD | rf | rf | clr | hf | hf | clr | Climbed at ~18z, RTD shield snowpacked, removed lower shield to clean.
2016-10-27 | 301 | 12:05 | SWD | clr | clr | clr | hf | clr | clr
2016-10-28 | 302 | 12:06 | SWD | clr | clr | clr | hf | lf | clr
2016-10-29 | 303 | 12:12 | SWD | lf | lf | clr | hf | lf | clr
2016-10-30 | 304 | 12:08 | SWD | clr | clr | clr | clr | clr | clr
2016-10-31 | 305 | 13:11 | SWD | lf | lf | clr | hf | lf | clr
2016-11-01 | 306 | 12:19 | SWD | clr | clr | clr | clr | clr | clr
2016-11-02 | 307 | 12:38 | SWD | lf | lf | clr | lf | lf | clr
2016-11-03 | 308 | 12:37 | SWD | lf | lf | clr | lf | lf | clr
2016-11-04 | 309 | 12:28 | SWD | lf | lf | clr | lf | lf | clr
2016-11-05 | 310 | 11:41 | SWD | lf | lf | clr | lf | lf | clr
2016-11-06 | 311 | 12:20 | SWD | hf | hf | clr | lf | lf | clr
2016-11-07 | 312 | 12:25 | SWD | hf | hf | clr | lf | lf | clr
2016-11-08 | 313 | 12:17 | SWD | lf | lf | clr | lf | lf | clr
2016-11-09 | 314 | 12:35 | SWD | lf | lf | clr | lf | lf | clr
2016-11-10 | 315 | 12:36 | SWD | lf | lf | clr | lf | lf | clr
2016-11-11 | 316 | - | SWD | [6 empty] | Not visited due to high winds.
2016-11-12 | 317 | - | SWD | [6 empty] | Not visited due to high winds.
2016-11-13 | 318 | 12:19 | SWD | clr | clr | clr | clr | clr | clr
2016-11-14 | 319 | 12:06 | SWD | lf | lf | clr | lf | lf | clr
2016-11-15 | 320 | 13:57 | SWD | lf | lf | clr | lf | lf | clr
2016-11-16 | 321 | 12:28 | SWD | hf | hf | clr | hf | hf | clr
2016-11-17 | 322 | 12:04 | SWD | hf | hf | clr | hf | hf | clr | Climbed at 14:20z to confirm 10 m RTD fan operation.
2016-11-18 | 323 | 12:32 | SWD | hf | hf | clr | hf | hf | clr
2016-11-19 | 324 | 12:31 | SWD | hf | hf | clr | hf | hf | clr
2016-11-20 | 325 | 11:35 | SWD | hf | hf | clr | hf | hf | clr
2016-11-21 | 326 | 12:45 | SWD | hf | hf | clr | hf | hf | clr
2016-11-22 | 327 | 12:34 | SWD | hf | hf | clr | hf | hf | clr
2016-11-23 | 328 | 12:25 | SWD | hf | hf | clr | hf | hf | clr
2016-11-24 | 329 | 12:33 | SWD | hf | hf | clr | hf | hf | clr
2016-11-25 | 330 | 12:12 | SWD | hf | hf | clr | hf | hf | clr
2016-11-26 | 331 | 11:36 | SWD | hf | hf | clr | hf | hf | clr
2016-11-27 | 332 | 11:55 | SWD | lf | lf | clr | lf | lf | clr
2016-11-28 | 333 | 12:06 | SWD | lf | lf | clr | lf | lf | clr
2016-11-29 | 334 | 12:33 | SWD | lf | lf | clr | lf | lf | clr
2016-11-30 | 335 | 12:27 | SWD | lf | lf | clr | lf | lf | clr
2016-12-01 | 336 | 12:16 | SWD | lf | lf | clr | lf | lf | clr
2016-12-02 | 337 | 12:35 | SWD | lf | lf | clr | hf | lf | clr | 10 m vane anemometer fully rimed.
2016-12-03 | 338 | 12:25 | SWD | lf | lf | clr | hf | hf | clr | 10 m vane anemometer fully rimed.
2016-12-04 | 339 | 11:35 | SWD | lf | lf | clr | hf | hf | clr | 10 m vane anemometer fully rimed.
2016-12-05 | 340 | 11:56 | SWD | lf | lf | clr | hf | hf | clr | 10 m vane anemometer fully rimed.
2016-12-06 | 341 | 12:25 | SWD | lf | lf | clr | hf | lf | clr | Tower climbed 14:35z to check 10 m RTD fan. No fan noise was detected, but difficult to confirm given the sound of the wind. The lower element of the radiation shield was removed and cleaned out. Quiet but distinct fan noise was detected. The 10 m sonic anemometer appears level in both axes. The 10 m vane anemometer is fully rimed.
2016-12-07 | 342 | 12:22 | SWD | lf | lf | clr | hf | clr | clr | 10 m vane anemometer fully rimed.
2016-12-08 | 343 | 12:12 | SWD | lf | lf | clr | hf | clr | clr
2016-12-09 | 344 | 12:23 | SWD | lf | lf | clr | hf | clr | clr
2016-12-10 | 345 | 11:55 | SWD | lf | lf | clr | hf | clr | clr
2016-12-11 | 346 | 11:28 | SWD | lf | lf | clr | hf | clr | clr
2016-12-12 | 347 | 12:21 | SWD | lf | lf | clr | hf | clr | clr
2016-12-13 | 348 | 12:33 | SWD | lf | lf | clr | hf | clr | clr
2016-12-14 | 349 | 12:20 | SWD | lf | hf | clr | hf | clr | clr
2016-12-15 | 350 | 12:16 | SWD | lf | lf | clr | hf | clr | clr
2016-12-16 | 351 | 12:09 | SWD | lf | lf | clr | lf | lf | clr
2016-12-17 | 352 | 12:13 | SWD | lf | lf | clr | lf | lf | clr
2016-12-18 | 353 | 12:03 | SWD | lf | lf | clr | lf | lf | clr
2016-12-19 | 354 | 12:05 | SWD | lf | lf | clr | lf | lf | clr
2016-12-20 | 355 | 12:14 | SWD | lf | lf | clr | lf | lf | clr
2016-12-21 | 356 | 12:11 | SWD | lf | lf | clr | lf | lf | clr
2016-12-22 | 357 | 12:15 | SWD | lf | lf | clr | lf | lf | clr
2016-12-23 | 358 | 12:09 | SWD | lf | lf | clr | lf | lf | clr
2016-12-24 | 359 | 11:42 | SWD | lf | lf | clr | lf | lf | clr
2016-12-25 | 360 | 12:25 | SWD | lf | lf | clr | lf | lf | clr
2016-12-26 | 361 | 12:11 | SWD | lf | lf | clr | lf | lf | clr
2016-12-27 | 362 | 12:23 | SWD | lf | lf | clr | lf | lf | clr
2016-12-28 | 363 | 12:20 | SWD | - | - | - | - | - | - | Tower damaged during winds 12/27.
2016-12-29 | 364 | [8 empty] | Daily checks suspended.
2016-12-30 | 365
2016-12-31 | 366
2017-10-12
2017-10-13
2017-10-14
2017-10-15
2017-10-16
2017-10-17
2017-10-18
2017-10-19
2017-10-20
2017-10-21
2017-10-22
2017-10-23
2017-10-24
2017-10-25
2017-10-26
2017-10-27
2017-10-28
2017-10-29
2017-10-30
2017-10-31
2017-11-01
2017-11-02
2017-11-03
2017-11-04
2017-11-05
2017-11-06
2017-11-07
2017-11-08
2017-11-09
2017-11-10
2017-11-11
2017-11-12
2017-11-13
2017-11-14
2017-11-15
2017-11-16
2017-11-17
2017-11-18
2017-11-19
2017-11-20
2017-11-21
2017-11-22
2017-11-23
2017-11-24
2017-11-25
2017-11-26
2017-11-27
2017-11-28
2017-11-29
2017-11-30
2017-12-01
2017-12-02
2017-12-03
2017-12-04
2017-12-05
2017-12-06
2017-12-07
2017-12-08
2017-12-09
2017-12-10
2017-12-11
2017-12-12
2017-12-13
2017-12-14
2017-12-15
2017-12-16
2017-12-17
2017-12-18
2017-12-19
2017-12-20
2017-12-21
2017-12-22
2017-12-23
2017-12-24
2017-12-25
2017-12-26
2017-12-27
2017-12-28
2017-12-29
2017-12-30
2017-12-31