**PFR**

**Weekly data download:**

* Before undocking the Toughbook in the GH, check that the time is correct.
* AT the solar tracker connect the Toughbook to the CR10 (in the S-Facing enclosure) using the SC32.
  + Open S-Facing enclosure and disconnect the DB-9 connector from the CR-10 datalogger.
  + Connect the laptop to the datalogger with the flat ribbon cable and the SC32 inline.
  + Start PC208W if it isn’t already running.
  + Click the‘Connect’ button and select ‘CR10’ from the list of instruments.
  + Click ‘connect’ on the bottom right.
  + Check the datalogger time and correct if needed. Record the offset in time for the ETH spreadsheet.
  + Disconnect the cable to the laptop and reconnect the Storage Module cable to the CR10.
  + Swap storage modules. There is one wrapped in white and another in black. Take the one you removed to TAWO.